## Entity Summary
Field | Extracted Data | Remarks
Establishment Name | ANODYNE THERAPY, LLC
Registration No | 1055581
Registration Year | 2023
Total No Of Devices | 2
Class 3 Devices | 0
No of Inspections | 7
No Of Recalls | -
Warning Letters | -
1.Response Letter | N/A
2.Closeout Letter | N/A
Adverse Events | 3
Is This an adverse Report? | 3
Establishment | ANODYNE THERAPY, LLC 301B Mears Blvd Oldsmar, FL 34677 Registration Number: 1055581 FEI Number*: 1000113231 Status: Active Date of Registration Status: 2023
Owner/Operator | ANODYNE THERAPY, LLC 301B Mears Blvd Oldsmar, FL US 34677 Owner/Operator Number: 9004867
Official Correspondent | jonathan ward AJW Technology Consultants, Inc Suite #503 11705 Boyette Road Riverview, FL 33569 Phone: 813-6452855
510(k) Device Name | Applicant | 510(k)Number | Decision Date
510K Table Not Found

## Recalls
Recall Link | Trade Name | Recall Class | Posted internet Date | Termination Date | FDA Determine Cause | Quantity in Commerce | Manufacturer Reason For Recall | Action | Distribution
Result Not Found

## MAUDE
Report Number | Event Date | FDA Received Date | Report Lag | Event Type | Report Source | Manufacturer-D | Manufacturer-G | Brand Name | Type Of Device | Manufacturer Review | Event Text | Manufacturer Narrative | Was Device Available For Evaluation ? | Is This an Adverse Event Report ? | MAUDE Link
1055581-2015-00001 | 05/07/2015 | 06/03/2015 | 27 | Injury | Manufacturer | ANODYNE THERAPY, LLC 14105 mccormick dr tampa FL 33636 | - | ANODYNE THERAPY | INFRARED DEVICE | - | Pt advises that she received a topical burn following treatment with anodyne therapy administered by a healthcare professional. This was the 11th treatment received, with no prior problems. Pt reports she went to the er for treatment 2 days following this anodyne treatment. Pt provided the er notes which confirm absence of treatment consistent with a topical burn, ie no topical treatment and no wound covering. Er notes do not mention a thermal injury, rather that pt was treated for cellulites with antibiotics and pain medication. Pt also provided visual evidence that does not show a topical burn, rather multiple skin red marks covering both legs in places where anodyne pads were not placed based upon both clinical protocols and soap notes from the treating facility / physician. | Anodyne therapy contacted treating facility to discuss the reports made by the pt. Facility provided soap notes and treatment protocols which dod not support a topical burn or any other injury. Facility was requested to identify the anodyne system used during treatment, and return it anodyne for eval, but has failed to do so. Follow up calls were made to schedule add'l training for facility users as is our normal practice. All calls were not returned, and facility stated we were not to contact staff members. In summation, no evidence has been provided to confirm a topical burn has occurred, or that anodyne therapy contributed to the reported event. We do not anticipate receiving any further info to continue an investigation, however if we do, we will report as appropriate. Pt reports she is currently 90-95% healed. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=4827570&pc=ILY
1055581-2013-00004 | 10/18/2013 | 11/25/2013 | 38 | Injury | Manufacturer | ANODYNE THERAPY, LLC tampa FL 33626 | - | ANODYNE THERAPY | INFRARED LAMP | - | Treating facility reports a patient developed a burn with blisters approximately 1/2" x 1/2" and a reddened area 5cm x 2cm on the calf following an anodyne therapy treatment. Reported patient injury has completely healed. | The system involved in this complaint was returned for evaluation after patient reportedly developed a 2nd degree burn with blisters approximately 1/2 x 1/2" and a reddened area approximately 5cm x 2cm on the le calf following treatment with an anodyne therapy professional system. Patient self administered silvadene, and did receive medical intervention of a topical antibiotic. Reported treatment time and intensity was within the guidelines provided to the user in the ifu manual. Evaluation identified a broken wire in one therapy pad, and a bad connection in one jack, neither of which would be expected to cause high temperatures. Additionally, during simulated use, the therapy pad temperatures were found to be well within levels that would not be expected to cause a burn. The facility reported that they used this same anodyne unit on various other patients around the same time as the reported injury without incident. Accordingly, anodyne is not able to confirm that this anodyne unit caused the reported injury. The number of incidents of this type remain within the expected rate for this product based upon the number of incidents reported and the number of units in distribution. Company continues to monitor and trend events of this type. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=3503984&pc=ILY
1055581-2013-00003 | 07/16/2013 | 08/20/2013 | 35 | Injury | Manufacturer | ANODYNE THERAPY, LLC tampa FL 33626 | - | ANODYNE THERAPY | INFRARED LAMP | - | Treating facility reports a pt developed a burn the size of a quarter on the calf following an anodyne therapy treatment. | The system involved in this complaint was returned for eval after pt reportedly developed a 2nd degree burn the size of a quarter on the rt calf following treatment with an anodyne therapy professional system. Pt did received medical intervention - a topical antibiotic. Reported treatment time and intensity was within the guidelines provided to the user in the ifu manual. Upon eval, a defective transistor was identified; however, therapy pad temperatures during simulated use were found to be will within levels that would not be expected to cause a burn. Additionally, the facility reported they used this same anodyne unit on other pts without incident around the same time as the reported injury. Accordingly, anodyne is not able to confirm that this anodyne unit caused the reported injury. The number of incidents of this type remain within the expected rate for this product based upon the number of incidents reported and the number of units in distribution. Company continues to monitor and trend events of this type. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=3313383&pc=ILY

## Inspections
City | State | Zip Code | Inspection End Date | Project Area | Classification
Oldsmar | Florida | 34677 | 01/23/2020 | Compliance: Devices | Voluntary Action Indicated (VAI)
Oldsmar | Florida | 34677 | 01/23/2020 | Postmarket Assurance: Devices | No Action Indicated (NAI)
Oldsmar | Florida | 34677 | 01/23/2020 | Radiation Control and Health Safety Act | Voluntary Action Indicated (VAI)
Oldsmar | Florida | 34677 | 01/09/2013 | Compliance: Devices | No Action Indicated (NAI)
Oldsmar | Florida | 34677 | 01/09/2013 | Postmarket Assurance: Devices | No Action Indicated (NAI)
Oldsmar | Florida | 34677 | 04/17/2009 | Compliance: Devices | No Action Indicated (NAI)
Oldsmar | Florida | 34677 | 04/17/2009 | Postmarket Assurance: Devices | No Action Indicated (NAI)

## Devices
Device Summary | Device Detail | TPLC Information | # Reported Events
Device Name | SPECTROPAD SYSTEM | TPLC Device Problem | MDR | Event In MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem32 | 5 | 5
Product Code | ILY | Overheating of Device33 | 3 | 3
Regulation No | 890.55 | Patient-Device Incompatibility34 | 3 | 3
Premarket Submission No | K931261 | Improper or Incorrect Procedure or Method35 | 2 | 2
Establishment Operations | Manufacturer | Defective Device36 | 1 | 1
Recalls | -
Device Name | PAD, HEATING, POWERED | TPLC Device Problem | MDR | Event In MDR
Device Class | 2 | Improper or Incorrect Procedure or Method29 | 190 | 190
Product Code | IRT | Adverse Event Without Identified Device or Use Problem30 | 154 | 154
Regulation No | 890.57 | Fire31 | 137 | 137
Premarket Submission No | - | Sparking32 | 115 | 115
Establishment Operations | Manufacturer | Excessive Heating33 | 109 | 109
Recalls | -

## Info
Establishment Name | Error | Timestamp
ANODYNE THERAPY, LLC | 510K Table Not Found | 2023-07-28 18:14
ANODYNE THERAPY, LLC | Recalls Not Found | 2023-07-28 18:14
ANODYNE THERAPY, LLC | Warning Result Not Found | 2023-07-28 18:18