## Entity Summary
Field | Extracted Data | Remarks
Establishment Name | BIOMET FRANCE
Registration No | 3006946279
Registration Year | 2023
Total No Of Devices | 92
Class 3 Devices | -
No of Inspections | 2
No Of Recalls | 2
Warning Letters | -
1.Response Letter | N/A
2.Closeout Letter | N/A
Adverse Events | 500
Is This an adverse Report? | 221
Establishment | BIOMET FRANCE Plateau de Lautagne Valence Drome, FR 26000 Registration Number: 3006946279 FEI Number*: 3006946279 Status: Active Date of Registration Status: 2023
Owner/Operator | Biomet, Inc. 56 East Bell Dr. PO Box 587 Warsaw, IN US 46582 Owner/Operator Number: 1825034
Official Correspondent | yannick bossert Plateau de Lautagne Valence, Drome FR 26000 Phone: 33-675-227628 US Agent: Holly Seppanen Zimmer Biomet 1800 West Center Street Warsaw, IN US 46580 Phone: 952 8470494 ext Email: holly.seppanen@zimmerbiomet.com
510(k) Device Name | Applicant | 510(k)Number | Decision Date
510K Table Not Found

## Recalls
Recall Link | Trade Name | Recall Class | Posted internet Date | Termination Date | FDA Determine Cause | Quantity in Commerce | Manufacturer Reason For Recall | Action | Distribution
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=143158 | Device Recall EndobonXenograft Granules | Class 2 | February 09, 2016 | May 31, 2017 | Under Investigation by firm | 145,116 units | Cytotoxicity tests were not passed at 36 months. Use of Endobon-Xenograft Granules past 18 months of its specified shelf life may result in patient illness/injury. | Zimmer BIOMET France SARL has initiated a voluntary recall via certified mail on January 2, /2016 for the Endobon¿ Xenograft Granules, based on limited (in vitro/bench) results which indicated cytotoxicity tests were not passed at 36 months (real-time) aging. Use of the Endobon¿ Xenograft Granules may result in patient may experiencing soft tissue irritation, up to and including infection. Customers area asked to do the following: 1. Please review this notice. 2. Identify, remove from service and segregate all affected Endobon ROX* product in your inventory. 3. Complete the attached Business Reply Form (Attachment 1) and: a. Fax it to 877-650-0368 or email it to Biomet5351@stericycle.com. 4. If you have product to return, contact BIOMET 3i at the number on the Business Reply Form for an RMA number and return shipping label. a. Return a copy of the Business Reply Form along with any affected product to the address on the Business Reply Form. Replacement product will immediately ship based upon availability. For assistance or other questions that customers may have relative to this notice, please contact BIOMET 3i at 1-800-443-8166 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. | Worldwide Distribution - US including AL, AK, AR, AK, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MN, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to Canada, Japan, Europe.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=137454 | Device Recall Optigun Ratchet | Class 2 | June 10, 2015 | October 27, 2015 | Nonconforming Material/Component | 3 | The pin which maintains the knob button, may disconnect and become lost. This may result in a delay of surgery to obtain a new Optigun and detect/retrieve the pin, and may necessitate revision surgery if the pin falls into the surgical site and is not detected. | BIOMET sent an URGENT ADVISORY NOTICE dated May 14, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to check the product to identify to determine if the pin is defective, if defective they should return it to Biomet Warsaw, following the directions on the Response Form and confirm receipt of the notice by sending back the response form. Questions related to this notice should be directed to (574) 372- 1570 Monday through Friday 8 a.m. to 5 p.m. | Nationwide distribution to MD only

## MAUDE
Report Number | Event Date | FDA Received Date | Report Lag | Event Type | Report Source | Manufacturer-D | Manufacturer-G | Brand Name | Type Of Device | Manufacturer Review | Event Text | Manufacturer Narrative | Was Device Available For Evaluation ? | Is This an Adverse Event Report ? | MAUDE Link
3006946279-2023-00055 | Result Not Found | 06/27/2023 | 97 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOB BONE CMT R-3 | BONE CEMENT | - | It was reported that during surgery the cement paste set too fast. The cement was taken from the refrigerator about (10) minutes before the start of the surgery. The surgeon began using the cement approximately (4) minutes after the start of mixing. Insertion of the implant was about (5) minutes from the start of mixing. The cement then hardened in about (7) minutes. No impact on a patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ "japan". No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A retain sample of batch has been tested in the laboratory under standardized conditions. No unusual behavior during mixing, handling or setting. The reported behavior of the cement cannot be reproduced. A review of the device manufacturing records confirmed no abnormalities or deviations that could be related to the reported event. Device used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17209551&pc=MBB
3006946279-2023-00054 | 11/26/2022 | 06/20/2023 | 206 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOB BONE CMT R-3 | BONE CEMENT | - | It was reported that the patient was admitted to the hospital after a fall in the nursing home resulting in a traumatic brain injury and subsequent aspiration pneumonia. While repositioning the patient, pain was noted in the right hip. X-rays revealed medial femoral neck fracture. Surgery as delayed 2 days due to chronic anticoagulant therapy. Subsequently a right total hip arthroplasty was performed. Due to difficulty obtaining stabilization after the femoral stem and head were implanted, it was decided to place an acetabular cup. Once excess cement from the cup was removed, the surgeon waited for the cement to set before continuing. During this time, the patient¿s oxygen saturation and circulation collapsed. The surgical team initiated resuscitation measures, but after 30 minutes without success, measures were terminated, and the patient expired. All final implants were placed, and the patient was closed. The diagnosis provided was fulminant pulmonary embolism with right heart failure. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical product: optipac 60 refob bone cmt r-3; item# 4711500396-3; lot# ax14bc1701, cement restrictor; item# 237014; lot# 22c010201, avantage cemented cup s52; item# p0463052; lot# 0001551688, avantage e1 inlay s52 / 28; item# p0561e52; lot# 0001556471, protasul; item# 30. 28. 05; lot# 3114546, muller stem; item# 35. 00. 39. 125; lot# 3125258. Report source foreign: germany. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00027, 3006946279 - 2023 - 00054. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. An iso-setting test has been performed. A retain sample of batches has been tested in the laboratory under standardized conditions. No unusual cement paste temperature and hardening time has been noticed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. The review of the raw material certificate shows no nonconformity or deviation. Device used for treatment. Medical records were provided and reviewed from a healthcare professional. It was reported that the patient had multiple cardiac comorbidities and compromise prior to the procedure from a fall that resulted in a traumatic brain injury and subsequent aspiration pneumonia. The surgery was delayed to prevent excessive blood loss due to chronic anti-coagulation medications taken for the comorbid heart issues. After a fracture, it is common to develop clots and hematomas around the bone as the body attempts to repair. These clots can break off and travel through the vascular and respiratory systems, increasing the risk for pulmonary embolism. As the cement cured, the patient's respiratory and circulatory systems collapsed. It has been noticed that the cement was expired at the time of the surgery. According to the r&d department, the fact that the optipac has been used expired can cause bone cement implantation syndrome (clinical performance: intraoperative mortality or morbidity, pressure increase, potential driving fat into the circulation). The allergic risk also causing anaphylactic shock. From the information provided, it appears that the cause of death is a culmination of the patient's comorbidities, recent trauma, and intraoperative events. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17160290&pc=LOD
3006946279-2023-00060 | - | 06/19/2023 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | ENDOBON XENOGRAFT GRANULATES LARGE 5 ML | BONE GRAFTING MATERIAL, SYNTHETIC | - | It was reported that the implant failed to heal after sinus lift procedure. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 foreign austria. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17150576&pc=LYC
3006946279-2023-00059 | - | 06/19/2023 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | ENDOBON XENOGRAFT GRANULATES LARGE 5 ML | BONE GRAFTING MATERIAL, SYNTHETIC | - | It was reported that the implant failed to heal after sinus lift procedure. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ austria. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17150530&pc=LYC
3006946279-2023-00058 | 12/25/2022 | 06/16/2023 | 173 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that patient underwent left hip revision due to dislocation after initial surgery. No further event information available at the time of this report. | (b)(4). No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Unable to confirm complaint, no product information or medical records provided. Part and lot identification are necessary for review of device history records, neither were provided. Device is used for treatment. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. H3 other text : insufficient information available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17142770&pc=MBB
3006946279-2023-00057 | 01/19/2022 | 06/15/2023 | 512 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FMRL STEM SIZE +4 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that the patient underwent a revision surgery due to fracture of femoral component, caused by fall trauma. Revision was performed eight (8) years eight (8) months from initial hip arthroplasty. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: regenerex ringloc item#pt104056, lot#814810. Ringloc e1 item#ep043254, lot#2754272. G2 - foreign: the netherlands. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17133522&pc=OQI
3006946279-2023-00056 | 05/11/2023 | 06/09/2023 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIP 80 REFOB PLUS BONE CMT-3 | BONE CEMENT | - | It was reported that the cement paste took short time for setting. This event is related to a malfunction that could potentially lead to a sterility issue. Due diligence is in progress for this complaint; to date, whatever additional information received has been included in this report. | (b)(4). G2-foreign-south africa the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17097928&pc=LOD
3006946279-2023-00053 | 01/30/2023 | 05/22/2023 | 112 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVAN UHMWPE INSERT 50/28 | HIP PROSTHESIS | - | It was reported that patient underwent revision of hip prosthesis due to head disassociation from the liner after trauma. Disassociation led also to peri-prosthetic metallosis. Attempts have been made and no further information has been provided. | (b)(4). G2: foreign- france. Associated products: item number:14. 28. 05-20, item name: cocr head 28/-4 's' 12/14, item lot:3024199. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records confirmed no abnormalities or deviations. Device used for treatment. Radiograph was provided and reviewed from an health care professional/radiologist : right total hip arthroplasty is noted in which the femoral component is somewhat eccentrically positioned superolateral with respect of the central acetabular cup. Difficult to exclude additional metallic density material more inferomedial to the femoral head component (or whether this is simply heads from the threaded screws). Multiple metallic plates and screws noted spanning through the right iliac bone, acetabulum, and right pubic bone. Slight eccentric positioning of the femoral head component with respect to the acetabular cup. Findings could be related to polyethylene disassembly or wear. Otherwise fit alignment appears normal. No signs of loosening, wear, radiolucency or other contributing factors. Disassociation can not be confirmed, although there are secondary findings which can suggest this association of the polyethylene from the acetabular cup. No contributing factors are detected. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16973614&pc=JDI
3006946279-2023-00052 | 04/11/2023 | 05/22/2023 | 41 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CMNT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the bone cement hardened faster than normal; within 5 minutes after mixing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16972335&pc=LOD
3006946279-2023-00049 | 03/27/2023 | 05/17/2023 | 51 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CMK REVSTM CM MD12/14 301 L200 | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the left hip approximately three (3) years post implantation due to pain, increasing functional impotence and femoral implant fracture in its cement sheath. Pseudarthrosis of the greater trochanter was diagnosed. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2-foreign-france. G2-other-ansm r2309708 investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16946403&pc=LZO
3006946279-2023-00051 | 04/14/2022 | 05/17/2023 | 398 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported the study patient developed a gradual onset of persistent pain and instability approximately one year post implantation. Radiographic images showed radiolucency surrounding the tibial implant and a triple phase bone scan displayed aseptic loosening. The patient underwent a revision where the femoral and tibial components were found to have minimal ingrowth. All components, except patella, were revised. No intra-op complications reported. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: femur cemented standard left size 7; item# 42-5046-062-01; lot# 64881021. Femoral distal augment cemented size 7, 7+ 10 mm thickness; item# 42-5566-062-10; lot# 64882685. Femoral distal augment cemented size 7, 7+ 10 mm thickness; item# 42-5566-062-10; lot# 64866990. Femoral posterior augment cemented size 7, 7+ 5 mm thickness; item# 42-5568-062-05; lot# 64859470. Stem extension tapered cemented 14 mm diameter +75 mm length; item# 42-5600-075-14; lot# 64911499. Rticular surface fixed bearing constrained posterior stabilized (cps) left 14 mm height use with tibia sizes e-f / ps femur sizes 6-9; item# 42-5126-007-14; lot# 64629122 tibia fixed cemented left size e stem extension use required; item# 42-5420-071-01; lot# 64817993. Tibial augment cemented half block left lateral size ef 10 mm thickness; item# 42-5558-051-10; lot# 64847601. Tibial augment cemented half block left medial size ef 10 mm thickness; item# 42-5558-053-10; lot# 64783331. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16944081&pc=MBB
3006946279-2023-00050 | 04/14/2023 | 05/15/2023 | 31 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | BONE CEMENT, ANTIBIOTIC | - | It was reported that bone cement hardened faster than normal. It took about 5 minutes from the start of mixing. Patient outcome: delayed surgery. | (b)(4). G2 ¿ foreign ¿ japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16929473&pc=MBB
3006946279-2023-00048 | 03/29/2023 | 05/02/2023 | 34 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED CUP S46 | HIP PROSTHESIS | - | It was reported that an initial right total hip arthroplasty was performed on an unknown date with unknown product. The patient underwent first revision surgery after suffering at home that resulted in a periprosthetic acetabular fracture with loosening of the cup. The initial stem was retained, the acetabular components were revised without complication. A plate and screws were placed to stabilize the fracture. The patient then underwent a second acetabular revision due to loosening of the cup and disassociation of the head and dual mobility components. A smaller cup was cemented into place with a dual mobility head and liner. The initial stem remained intact. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2-foreign- austria. G2-other-austrian federal office for safety in health care(basg ref. : (b)(4)). D10. Item#:unknown ;lot#:unknown ;item name: dual mobility liner ; item#:unknown ;lot#:unknown ;item name: metal head 28mm+10. 5 ; item#:unknown ;lot#:unknown ;item name: unknown cement; item#:unknown ;lot#:unknown ;item name: unknown initial stem. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16852869&pc=KWA
3006946279-2023-00046 | 03/08/2023 | 04/20/2023 | 43 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery the bone cement hardened faster than normal, within 5 minutes after mixing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign - japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16774983&pc=LOD
3006946279-2023-00045 | 06/12/2020 | 04/20/2023 | 1042 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VG ROCC TIB TRAY COCR HA/PC 79 | KNEE PROSTHESIS | - | It was reported that a patient underwent a revision surgery due to pain, stiffness and allergic reaction to metal approximately two years post initial total left knee replacement. During the revision noted osteointegration defect and adhesions. The patella was repositioned above the native location with unknown screws. All components were exchanged with unknown titanium components without difficulty. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products:¿ vanguard rocc por fem s67. 5 l; item p09vmg06 , lot 0001043713, vg rocc tib brg uhmwpe 67. 5x10. 5x10mm; item p11v0610, lot 0000938659. Foreign france. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00044. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16774981&pc=JWH
3006946279-2023-00044 | 06/12/2020 | 04/20/2023 | 1042 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VANGUARD ROCC POR FEM S67.5 L | KNEE PROSTHESIS | - | It was reported that a patient underwent a revision surgery due to pain, stiffness and allergic reaction to metal approximately two years post initial total left knee replacement. During the revision noted osteointegration defect and adhesions. The patella was repositioned above the native location with unknown screws. All components were exchanged with unknown titanium components without difficulty. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products:¿ vg rocc tib tray cocr ha/pc 79; item p10vh006, lot 0001249199, vg rocc tib brg uhmwpe 67. 5x10. 5x10mm; item p11v0610, lot 0000938659. Foreign france. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00045. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16774877&pc=JWH
3006946279-2023-00043 | 02/21/2023 | 04/17/2023 | 55 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery the bone cement hardened faster than normal, within 5 minutes after mixing. Attempts have been made and no further information has been provided. | (b)(4). G2: foreign - japan. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A retain sample of batch x20aaf2301 has been tested in the laboratory under standardized conditions (23°c; rel. Humid > = 40 %). No unusual behavior during mixing, handling or setting. The reported behavior of the cement cannot be reproduced. A review of the raw material certificate confirmed no abnormalities or deviations. Device used for treatment. Reported event is not related to medical condition. Review of medical record is not applicable. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16749492&pc=LOD
3006946279-2023-00042 | 03/06/2023 | 04/13/2023 | 38 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | PROSTHESIS, KNEE | - | It was reported that the patient underwent a left knee arthroplasty revision approximately four (4) years post-operatively due to pain, instability, femoral loosening, recurrent patellar dislocations and patellar fragmentation. Competitor devices were used to complete the procedure. Revision operative records indicate that the knee joint was grossly unstable to varus-valgus stress and anterior-posterior stress. The femoral component was loose when tapped and the patellar component had uncovered eburnated bone. | (b)(4). Concomitant medical products: lps flex femoral component-option for cemented use only nitrogen hardened size e left; item#00596801551; lot#64017665. All poly patella standard size 32 mm diameter 8. 5 mm thickness for cemented use only; item#00597206632; lot#64178939. Stemmed tibial component precoat size 3 for cemented use only use of this tibial component with lcck articulating surfaces requires using a stem exten; item#00598003701; lot#64158604. Articular surface use with lps/lps-flex 51 or 52 suffix femorals size ef 12 mm height; item#00596203212; lot#63658236. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16730083&pc=MBB
3006946279-2023-00041 | - | 04/12/2023 | - | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT | - | It was reported that the patient underwent hip prosthesis insertion. Bodycardia, oxygen saturation reduction and circulatory instability has happened during operation. Patient has died. Further information in regards to the bone cement causing or contributing to the patient's death are unknown. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16721824&pc=MBB
3006946279-2023-00027 | 11/26/2022 | 04/04/2023 | 129 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT | - | It was reported that the patient died due to cardiac arrest. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16673763&pc=LOD
3006946279-2023-00038 | 02/28/2023 | 04/04/2023 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN GTS FEMORAL STEM STANDARD | HIP PROTHESIS | - | It was reported by the medical writing team that according to the german registry a femoral stem is noted to have a high re-revision rate. Attempts have been made and no further information has been provided. | (b)(4). Report source¿ foreign ¿ ¿germany¿ the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16672214&pc=KWA
3006946279-2023-00038 | 02/28/2023 | 04/04/2023 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN GTS FEMORAL STEM STANDARD | HIP PROTHESIS | - | It was reported by the medical writing team that according to the german registry a femoral stem is noted to have a high re-revision rate. Attempts have been made and no further information has been provided. | (b)(4). Report source¿ foreign ¿ ¿germany¿ the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16672214&pc=OQG
0009613350-2023-00152 | 02/28/2023 | 04/04/2023 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN GTS FEMORAL STEM STANDARD | HIP PROTHESIS | - | It was reported by the medical writing team that according to the german registry a femoral stem is noted to have a high septic re-revision rate. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign¿germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Product not returned. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16672213&pc=KWA
0009613350-2023-00152 | 02/28/2023 | 04/04/2023 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN GTS FEMORAL STEM STANDARD | HIP PROTHESIS | - | It was reported by the medical writing team that according to the german registry a femoral stem is noted to have a high septic re-revision rate. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign¿germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Product not returned. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16672213&pc=OQG
3006946279-2023-00040 | 02/07/2023 | 04/04/2023 | 56 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN REVISION 40 -3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the cement ratchet gun malfunctioned half way through a primary hip implantation causing the cement to begin to set. The surgeon removed the cement before it set, however the patient started to deteriorate causing low oxygen saturation and ecg changes. Fat embolus were also evident. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign- uk. Concomitant medical products: item#: unknown, lot#: unknown, item name: unknown optigun. Item#: 3011630001-3, lot#: x20cab0303, item name: refobacin revision 40 -3. Item#: 4160, lot#: 0001688806, item name: optivac m. Multiple mdr reports were filed for this event, please see associated reports: 3006946279-2023-00040. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16671676&pc=MBB
3006946279-2023-00039 | 02/07/2023 | 03/31/2023 | 52 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN REVISION 40 -3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the cement ratchet gun malfunctioned half way through a primary hip implantation causing the cement to begin to set. The surgeon removed the cement before it set, however the patient started to deteriorate causing low oxygen saturation and ecg changes. Fat embolus were also evident. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign- uk. Concomitant medical products: associated products: item#: unknown, lot#: unknown, item name: unknown optigun. Item#: 3011630001-3, lot#: x20cab0303, item name: refobacin revision 40 -3. Item#: 4160, lot#: 0001688806, item name: optivac m. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16654491&pc=MBB
3006946279-2023-00037 | 02/08/2022 | 03/27/2023 | 412 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN REFOBACIN BIOMET BONE CEMENT | BONE CEMENT, ANTIBIOTIC | - | It was reported that patient underwent revision surgery due to the pain, limited rom, stiffness, scar tissue and adhesions about two (2) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Femur cemented cruciate retaining (cr) narrow right size 7, item# 42502006202, lot# 64672793; articular surface medial congruent (mc) right 13 mm height use with tibia sizes cd/ cr femur sizes 6-7, item# 42522100413, lot# 64461580; all-poly patella cemented 29 mm diameter, item# 42540200029, lot# 64775977; the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Device remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16618395&pc=MBB
3006946279-2023-00036 | 02/21/2023 | 03/27/2023 | 34 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COCR FM HD D28/ 12-14/ 0/ M | HIP PROSTHESIS | - | It was reported that the patient underwent an initial right hip arthroplasty and subsequently during consultation the surgeon discovered that the implants had disassembled. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ france. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16617230&pc=LPH
3006946279-2023-00035 | - | 03/23/2023 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET® BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that patient suffered of burning pain in the foot, pain and the swelling in the knee with loss of range of motion after initial total knee surgery. Additionally, the nerve damage and tissue fusing to implant were discovered after doctor appointment in unknown timeframe after initial surgery. No medical or surgical intervention has been reported. Due diligence is completed for this complaint; no additional information is available and product has not been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16596336&pc=LOD
3006946279-2023-00033 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583986&pc=MBB
3006946279-2023-00030 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583560&pc=MBB
3006946279-2023-00034 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source : foreign ¿ sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583493&pc=MBB
3006946279-2023-00032 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign ¿ sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583447&pc=MBB
3006946279-2023-00031 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583408&pc=MBB
3006946279-2023-00028 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583402&pc=MBB
3006946279-2023-00029 | 02/01/2023 | 03/21/2023 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement has long setting time, approximately 14-15 minutes after vacuum mixing. The cement has a longer sticky and dough phases and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583401&pc=MBB
3006946279-2023-00025 | 02/08/2023 | 03/10/2023 | 30 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the liner would not seat properly. A delay occurred that required the surgeon to cement in the patient and spend an extra 20 (twenty) minutes. The patient coded when cementing the femur and expired due to pe (pulmonary embolism). Due diligence is in progress for this complaint; to date no additional information or product has been received. | Cmp-(b)(4). Medical devices: biomet bc r 1x40 us; item# 110035368; lot# ax21ab1101. Biomet bc r 1x40 us; item# 110035368; lot# ax27be0212. Liner constrained neutral 36 mm i. D. Size h for use with acetabular system only; item# 30203608; lot# 65606435. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00024; 3006946279 - 2023 - 00025; 3006946279 - 2023 - 00026. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16520166&pc=LOD
3006946279-2023-00024 | 02/08/2023 | 03/10/2023 | 30 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the liner would not seat properly. A delay occurred that required the surgeon to cement in the patient and spend an extra 20 (twenty) minutes. The patient coded when cementing the femur and expired due to pe (pulmonary embolism). Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: biomet bc r 1x40 us; item# 110035368; lot# ax05ba1003; biomet bc r 1x40 us; item# 110035368; lot# ax27be0212; liner constrained neutral 36 mm i. D. Size h for use with g7 acetabular system only; item# 30203608; lot# 65606435. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00025 and 3006946279 - 2023 - 00026. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16520165&pc=LOD
3006946279-2023-00026 | 02/08/2023 | 03/10/2023 | 30 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the liner would not seat properly. A delay occurred that required the surgeon to cement in the patient and spend an extra 20 (twenty) minutes. The patient coded when cementing the femur and expired due to pe (pulmonary embolism). Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: biomet bc r 1x40 us; item# 110035368; lot# ax21ab1101; biomet bc r 1x40 us; item# 110035368; lot# ax05ba1003; liner constrained neutral 36 mm i. D. Size h for use with g7 acetabular system only; item# 30203608; lot# 65606435. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2023 - 00024 and 3006946279 - 2023 - 00025. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16520162&pc=LOD
3006946279-2023-00022 | 01/26/2023 | 03/06/2023 | 39 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the bone cement hardened within 5 minute after mixing. Attempts have been made and no further information has been provided. | (b)(4). Foreign - japan the device will not be returned for analysis:however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16484829&pc=LOD
3006946279-2023-00021 | 02/15/2023 | 03/03/2023 | 16 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent revision surgery on left knee after 15-month post-implantation due to pain and possible loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: articular surface medial congruent (mc) left 12 mm height use with tibia sizes e-f/cr femur sizes 6-7, item# 42512100712, lot# 64767244, natural tibia trabecular metal two-peg porous fixed bearing left size e, item# 42530007101, lot# 65068961, femur trabecular metal cruciate retaining (cr) narrow porous, item# 42502206001, lot# 65100496, nexgenâ® complete knee solution, trabecular metal standard primary patella item# 00587806532, lot# 65081838. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16477710&pc=LOD
3006946279-2023-00020 | 02/01/2023 | 02/27/2023 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the bone cement has too long a setting time, approximately 14-15 minutes after vacuum mixing. The cement has longer sticky and dough phases, and it feels very soft during application. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source¿ foreign ¿ sweden. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16441847&pc=MBB
3006946279-2023-00008 | 01/12/2023 | 02/23/2023 | 42 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BIOMET BONE CEMENT | - | It was reported that when the surgeon mixed the bone cement, the monomer and polymer did not mix. The consistency remained powdery after waiting more than 3 minutes. The product could not be used. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source¿ foreign ¿ japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16426187&pc=LOD
3006946279-2023-00008 | 01/12/2023 | 02/23/2023 | 42 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BIOMET BONE CEMENT | - | It was reported that when the surgeon mixed the bone cement, the monomer and polymer did not mix. The consistency remained powdery after waiting more than 3 minutes. The product could not be used. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source¿ foreign ¿ japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16426187&pc=MBB
3006946279-2023-00019 | 01/08/2023 | 02/20/2023 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that approximately four months after total knee arthroplasty, the patient underwent revision due to infection. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products associated devices: item name:vngd ant stblzd brg 14x79, item number:189104, lot number:321720. Item name:vanguard cr ilok fem-rt 70, item number:141215, lot number:j7181424. Item name:series a pat std 34 3 peg, item number: 184766, lot number:152410. Item name:unk tibial stem ext, item number:unk tibial stem ext, lot number:240000. Item name:refobacin bc r 1x40 us, item number:110034355, lot number:y46baj2103. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16403477&pc=MBB
3006946279-2023-00018 | 01/08/2023 | 02/20/2023 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that approximately four months after total knee arthroplasty, the patient underwent revision due to infection. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: item name:vngd ant stblzd brg 14x79, item number:189104, lot number:321720. Item name:vanguard cr ilok fem-rt 70, item number:141215, lot number:j7181424. Item name:series a pat std 34 3 peg, item number: 184766, lot number:152410. Item name:unk tibial stem ext, item number:unk tibial stem ext, lot number:240000. Item name:refobacin bc r 1x40 us, item number:110034355, lot number:ay25ci2901. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16403323&pc=MBB
3006946279-2023-00017 | 01/23/2023 | 02/20/2023 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CMNT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the bone cement hardened within 5 minutes after mixing. Attempts have been made and no further information has been provided. | (b)(4). Foreign source: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16402372&pc=LOD
3006946279-2023-00017 | 01/23/2023 | 02/20/2023 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CMNT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the bone cement hardened within 5 minutes after mixing. Attempts have been made and no further information has been provided. | (b)(4). Foreign source: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16402372&pc=MBB
3006946279-2023-00015 | 02/01/2023 | 02/17/2023 | 16 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the patient went under revision surgery due to pain in the knee. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: femur trabecular metal cruciate retaining (cr) narrow porous , item# 42502205802 lot# 64894501, natural tibia trabecular metal two-peg porous fixed bearing right size d item# 42530006702 lot# 64696346, nexgenâ® complete knee solution, trabecular metalâ standard primary patella item# 00587806532 lot# 65019535, articular surface medial congruent (mc) right 11 mm height use with tibia sizes c-d/cr femur sizes 4-5 item# 42522100311 lot# 64821566, biomet bc r 1x40 us item# 110035368 lot# az50ba2102. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16395978&pc=LOD
3006946279-2023-00016 | 02/01/2023 | 02/17/2023 | 16 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent the revision surgery due to pain in the knee. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: femur trabecular metal cruciate retaining (cr) narrow porous item# 42502205802 lot# 64894501. Natural tibia trabecular metal two-peg porous fixed bearing right size d item# 42530006702. Lot# 64696346. Articular surface medial congruent (mc) right 11 mm height use with tibia sizes c-d/cr femur sizes 4-5 item# 42522100311 lot# 64821566. Nexgenâ® complete knee solution, trabecular metalâ¿¢ standard primary patella item# 00587806532. Lot# 65019535. Biomet bc r 1x40 us item 110035368 lot# y11baa2101. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16395977&pc=LOD
3006946279-2023-00012 | 12/16/2022 | 02/16/2023 | 62 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | . plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that patient blood pressure suddenly dropped during surgery in which bone cement has been used. Patient was given continuous noradrenaline injection. On the next day there was another blood pressure drop and patient died. According to hospital staff the bone cement is considered to be a major possibility regarding the cause of death, but the cement cannot be defined as the sole cause of death, as the patient's physical condition and anesthesia management may also have been factors. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign- japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16384671&pc=LOD
3006946279-2023-00014 | 09/18/2020 | 02/16/2023 | 881 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the patient experienced pain and diagnosed with synovitis three months post-surgery. A steroid injection was administered, and no further complications were reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products:¿ item#:42504606611 ; lot#:64463036 ; item name: femur cemented plus left size 9+; item#:42512800710 ; lot#: 64316885 ; item name: articular surface fixed bearing constrained; condylar knee (cck) left 10 mm height use with tibia sizes e, f / revision femur sizes 7, 7+, 9, 9+ with l ; item#:unknown femoral stem extension 14x30 ; lot#: unknown; item name: unknown femoral stem extension 14x30; item#:unknown tibial component size f ; lot#:unknown; item name:unknown tibial component size f ; item#:unknown tibial stem extension 14x75; lot#: unknown ; item name: unknown tibial stem extension 14x75; item#:unknown patella 32mm ; lot#:unknown ; item name: unknown patella 32mm; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16384255&pc=LOD
3006946279-2023-00013 | 01/16/2023 | 02/15/2023 | 30 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that cemented stem was dislodged from cement mantle during closed reduction of dislocation approximately 1-month post-implantation. Patient was revised. Due diligence is in progress for this complaint; to date no additional information. | (b)(4). Concomitant medical products - associated products; femoral stem low demand/fracture cemented 12/14 neck taper rasp 14 or 15 size 14 135 mm stem length , item# 00783301400 , lot# 64874473. Endo femoral head 12 / 14 taper , item# 00781805100 , lot# 65499504. Distal centralizer 12 mm o. D. , item# 00785901200, lot# 65509002. Endo femoral head adaptor 12 / 14 taper plus 7. 0 mm neck length item# 00781809900 , lot# 65366691. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16377962&pc=MBB
3006946279-2023-00011 | 01/05/2023 | 02/03/2023 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | OPTIPAC (E) : BBCR | - | It was reported that bone cement hardened faster than normal. It took about 5 minutes from the start of the mixing. Fifteen minute delay to replace the product. No health consequences to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16290058&pc=MBB
3006946279-2023-00010 | 01/09/2023 | 02/02/2023 | 24 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | . plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S56 / 28 | HIP PROSTHESIS | - | It was reported that after approximately three months post revision surgery, x-rays confirmed that another dislocation had occurred. Patient outcome: revision. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: femoral head -3. 5x26mm dia item#00-8018-026-01 lot#66341383. Foreign source: south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16283480&pc=KWA
3006946279-2023-00010 | 01/09/2023 | 02/02/2023 | 24 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | . plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S56 / 28 | HIP PROSTHESIS | - | It was reported that after approximately three months post revision surgery, x-rays confirmed that another dislocation had occurred. Patient outcome: revision. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: femoral head -3. 5x26mm dia item#00-8018-026-01 lot#66341383. Foreign source: south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16283480&pc=KWY
3006946279-2023-00009 | 12/29/2022 | 02/01/2023 | 34 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VARIZED FEMORAL STEM EXCEPTION -CEMENTLESS- 5 42- 12/14 - LEFT S5 | COATED HIP FEMUR PROSTHESIS | - | It was reported that the patient underwent revision surgery due to stem fracture approximately 8 years post-implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- france. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16276495&pc=LPH
3006946279-2023-00007 | 01/05/2023 | 01/31/2023 | 26 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that after cementing the hip stem, patient was transferred to recovery room and post-op films were taken. Films revealed that the stem had come out of the cement mantle. Patient was brought back to operation room and the stem was recemented. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16270383&pc=LOD
3006946279-2023-00006 | 01/12/2023 | 01/26/2023 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that the cement could not be processed, because it solidified too quickly, within 6-7 minutes. No health consequences to patient were reported. Attempts have been made and no further information has been provided. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16245328&pc=MBB
3006946279-2023-00005 | 01/12/2023 | 01/25/2023 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that the cement could not be processed, because it solidified too quickly, within 6-7 minutes. No health consequences to patient were reported. Attempts have been made and no further information has been provided. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16239982&pc=MBB
3006946279-2023-00003 | - | 01/16/2023 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | . plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent diagnostic intervention (biopsy for metal allergy) greater than six months from initial surgery. The patient also experiences pain and swelling. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Medical devices: oxf anat brg lt md size 4 pma; item# 159548; lot# unknown. Oxf twin-peg cmntd fem md pma; item# 161469; lot# unknown. Oxf uni tib tray sz d lm pma; item# 154724; lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16172718&pc=MBB
3006946279-2023-00001 | 12/18/2022 | 01/12/2023 | 25 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent revision due to infection. Due diligence is in progress for this complaint; to date no additional information or product has been received. | Cmp-0856181 concomitant medical products: item#:189103 ; lot#:019330 ; item name: vngd ant stab brg 13x79; item#:141235 : lot#:j6973615 ; item name: biomet cc cruciate tray 79mm; item#:183032 ; lot#:j7138334 ; item name: vanguard cr ilok fem-lt 70; item#:184788 ; lot#:356390 ; item name: series a pat thin 37x8. 6 3 peg; the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16145610&pc=LOD
3006946279-2023-00004 | 12/19/2022 | 01/11/2023 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 2X20-3 | CEMENT PACK (E) : RBCR | - | It was reported that consistency of the cement was not suitable for application. No further consequence or patient harm reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: austria. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16136317&pc=LOD
3006946279-2023-00002 | 08/22/2022 | 01/10/2023 | 141 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the patient experienced pain and implant loosening approximately four (4) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: lps fem comp sz e-r, item#: 00-5996-015-02, lot#: 63931769; st prc tib plt size 4, item#: 00-5980-037-02, lot#: 63987026; lps-flex fxd mld ef/3-4 12mm, item#: 00-5964-032-12, lot#: 63908007. Report source: foreign: germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16129584&pc=MBB
3006946279-2022-00129 | 12/15/2022 | 12/28/2022 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BC R | BONE CEMENT | - | It was reported that the bone cement hardened faster than normal and there was a delay of a few minutes to replace the product. No patient harm reported. Attempts have been made and no further information has been provided yet. | (b)(4). Report source: foreign: japan investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16058388&pc=LOD
3006946279-2022-00128 | 11/25/2022 | 12/22/2022 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the cement hardened in 2~3 minutes during knee surgery (uka). The surgeon needed to prepare another cement to complete surgery. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source - foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16033385&pc=MBB
3006946279-2022-00127 | 05/03/2017 | 12/22/2022 | 2059 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION VAR STEM LEFT SIZE 4 | HIP PROSTHESIS | - | It was reported that the patient underwent revision surgery due to pain and subluxation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Concomitant medical products: item (b)(4); lot#2016080307; item name: delta cer fm hd 036/+4mm 12/14. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16031792&pc=JDI
3006946279-2022-00127 | 05/03/2017 | 12/22/2022 | 2059 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION VAR STEM LEFT SIZE 4 | HIP PROSTHESIS | - | It was reported that the patient underwent revision surgery due to pain and subluxation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Concomitant medical products: item (b)(4); lot#2016080307; item name: delta cer fm hd 036/+4mm 12/14. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16031792&pc=LZO
3006946279-2022-00126 | 06/13/2022 | 12/20/2022 | 190 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION VAR STEM LEFT SIZE 5 | HIP PROSTHESIS | - | It was reported that the patient was revised due to stem neck fracture. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: france. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16014095&pc=LZO
3006946279-2022-00125 | - | 12/16/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that a patient underwent an initial right total knee arthroplasty. The patient has subsequently experienced pain for over one year post-implantation. Infection has been ruled out and the medical team cannot determine what is causing the pain. The current outcome of the patient is unknown and further planned medical intervention has not been reported. Due diligence is in process for this event; to date no additional information has been reported. | (b)(4). Medical devices: articular surface medial congruent (mc) right 12 mm height use with tibia sizes g-h/cr femur sizes 8- 11; item# 42522100912; lot# 64403887; femur cemented cruciate retaining (cr) standard right size 10; item# 42502606802; lot# 64824213; all-poly patella cemented 38 mm diameter; item# 42540200038; lot# 65029458; tibia cemented 5 degree stemmed right size g; item# 42532007902; lot# 64950764. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15993648&pc=MBB
3006946279-2022-00125 | - | 12/16/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN REFOBACIN BONE CEMENT | BONE CEMENT | - | It was reported that a patient underwent an initial right total knee arthroplasty. The patient has subsequently experienced pain for over one year post-implantation. Infection has been ruled out and the medical team cannot determine what is causing the pain. The current outcome of the patient is unknown and further planned medical intervention has not been reported. Due diligence is in process for this event; to date no additional information has been reported. | (b)(4). Medical devices: articular surface medial congruent (mc) right 12 mm height use with tibia sizes g-h/cr femur sizes 8- 11; item# 42522100912; lot# 64403887; femur cemented cruciate retaining (cr) standard right size 10; item# 42502606802; lot# 64824213; all-poly patella cemented 38 mm diameter; item# 42540200038; lot# 65029458; tibia cemented 5 degree stemmed right size g; item# 42532007902; lot# 64950764. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15993648&pc=LOD
3006946279-2022-00124 | 11/30/2022 | 12/13/2022 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R | OPTIPAC REFOBACIN R | - | It was reported that the product remained powdery after the mixing process. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: austria. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15968488&pc=MBB
3006946279-2022-00121 | 11/17/2022 | 12/12/2022 | 25 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN INLAY | UNKNOWN INLAY | - | It was reported that during the revision, the liner did not fit in the socket. The procedure had to be completed with the explanted item causing a 10 minutes delay. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958387&pc=KWA
3006946279-2022-00122 | 11/17/2022 | 12/12/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S58 / 28 | HIP PROSTHESIS | - | It was reported that during the revision the liner did not fit in the socket. The procedure had to be completed using the explant, causing a 10 minutes delay. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: avantage e1 inlay s58 / 28 item #p0561e58, lot#0001567491. Report source: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958377&pc=KWA
3006946279-2022-00123 | 11/16/2022 | 12/12/2022 | 26 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN BIOMET BC R BONE CEMENT | BONE CEMENT | - | It was reported that the patient underwent an initial left hemiarthroplasty. Approximately 2 months post implantation, the patient was revised due to pain, difficulty ambulating, and loosening. During the revision, periprosthetic fracture to the tibia was noted. All components were replaced with conversion to total knee replacement without complications. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products:¿ partial tibial cemented size d left medial; item# 42538000401; lot# 65066289; partial femur cemented size 4 left media; item# 42558000401; lot# 65286540; partial articular surface left medial size d 12 mm thickness; item# 42518200412; lot# 64730371. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15957647&pc=LOD
3006946279-2022-00120 | 11/08/2022 | 12/07/2022 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the cement could not be processed, because it solidified too quickly, within 5-6 minutes. No health consequences to patient were reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15931218&pc=MBB
3006946279-2022-00112 | 10/18/2022 | 11/28/2022 | 41 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 BIOMET BC R | ORTHOPEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that during surgery, the cement hardened within 6-7 minutes. No health consequences or impact to patient were reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15868493&pc=MBB
3006946279-2022-00117 | 10/19/2022 | 11/24/2022 | 36 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that the patient's blood pressure dropped while implanting the cementing system during a total hip arthroplasty. The patent expired a few days after the operation. Further information in regards to the bone cement causing or contributing to the patient's death are unknown. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: cem fem stem sz 11 std x 120, item# (b)(4), lot#64039395; zb 12/14 cocr hd 28mm x -3. 5, item# (b)(4), lot#2985404; distal centralizer 9mm, item# (b)(4), lot#63712104; ringloc bipolar vit e 28x41mm, item# (b)(4), lot#64900472. Report source: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15859923&pc=LOD
3006946279-2022-00119 | 10/28/2022 | 11/22/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the cement hardened within 6-7 minutes. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15843435&pc=LOD
3006946279-2022-00118 | 10/28/2022 | 11/22/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that during surgery, the cement hardened within 6-7 minutes. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign - japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Discarded. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15842657&pc=LOD
3006946279-2022-00116 | 10/24/2022 | 11/18/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that while mixing a batch of refobacin bone cement, the consistency became a white chalky texture. No known adverse event was reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: refobacin bc r 1x40 us, item#: 110034355, lot#: y46aal0702. Vanguard cr por fem-rt 60 item#: 183044, lot#: 025040. Biomet cc cruciate tray 67mm item#: 141232, lot#: j7300990. Vngd ant stblzd brg 10x67 item#: 189040, lot#: 165990. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15823730&pc=MBB
3006946279-2022-00115 | 09/26/2022 | 11/10/2022 | 45 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that during surgery, the cement hardened within 6-7 minutes. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15769997&pc=MBB
3006946279-2022-00114 | 07/22/2022 | 11/10/2022 | 111 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that after mixing the cement it set quicker than usually expected. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15767720&pc=LOD
3006946279-2022-00113 | 10/03/2022 | 11/07/2022 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It was reported that patient underwent a left total knee arthroplasty due to aseptic loosening, approximately four (4) years after initial surgery. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: femur cemented posterior stabilized (ps) standard left size 7 item#42500606201 lot#63946270; tibia cemented 5 degree stemmed left size d item#42532006701 lot#64019774; all poly patella cemented 32 mm diameter item#42540000032 lot#64098548; articular surface fixed bearing posterior stabilized (ps) left 10 mm height item#42512400510 lot#63756951. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15743133&pc=LOD
3006946279-2022-00111 | 10/24/2022 | 11/02/2022 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | . plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that: while mixing a batch of refobacin bone cement and it turned to white chalky texture. No known adverse event was reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: refobacin bc r 1x40 us, item 110034355, lot ay40ca0310. Vanguard cr por fem-rt 60, item183044, lot 025040. Biomet cc cruciate tray 67mm, item141232, lot j7300990. Vngd ant stblzd brg 10x67, item189040, lot 165990. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713497&pc=MBB
3006946279-2022-00110 | 05/21/2021 | 11/02/2022 | 530 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED ACETABULAR CUP / CEMENTED / SIZE 50X44 | HIP PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the right hip implants due to mobilization of the cup and loosening of the screws, approximately five (5) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: bone srew self-tapping item# 62506520 lot# 63387194; bone srew self-tapping item# 62506525 lot# 63347326; bone srew self-tapping item# 62506535 lot# 63333303; bone srew self-tapping item# 62506525 lot# 63373259; bone srew self-tapping item# 62506525 lot# 63347325; bone srew self-tapping item# 62506535 lot# 63387195. Report source foreign - italy. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713214&pc=LZO
3006946279-2022-00109 | 09/28/2022 | 10/27/2022 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent initial knee surgery. Subsequently, the patient underwent a two-stage revision operation due to infection. The surgeon removed the cr femur and cemented poly and replaced it with full knee revision components. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: e1 vngd as tib brg 12x67, item#: ep-189042, lot#013920. Vg 360 univ pst fm aug 67. 5x5, item#: 185345, lot#793090. Vg 360 dst fm ag 67. 5x5 ll/rm, item#: 185325, lot#901770. Vanguard cr ilok fem-rt 67. 5, item#: 183010, lot#j7125799. Vg 360 dst fm ag 67. 5x5 rl/lm, item#: 185305, lot#364420. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15678441&pc=LOD
3006946279-2022-00107 | 10/06/2022 | 10/27/2022 | 21 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOB BONE CMT R-3 | REFOBACIN PLUS BONE CEMENT | - | It was reported that: when changing the trial head, the stem slipped out of the cement mantle and the cup did not hold. Surgeon asks to check the cementing system reset sample; the socket slipped out of the cement mantle. Over 60 min delay. The event occurred during surgery. Patient experiences pain. Attempts have been made and no further information has been provided. | (b)(4). (b)(6). Product was not returned or pictures not provided. Device evaluation could not be performed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. This device is used for treatment. Reported event is not related to medical condition. Review of medical record is not applicable. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15678287&pc=LOD
3006946279-2022-00108 | 09/06/2022 | 10/27/2022 | 51 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOB BONE CMT R-3 | REFOBACIN PLUS BONE CEMENT | - | It was reported that: when changing the trial head, the stem slipped out of the cement mantle and the cup did not hold. Surgeon asks to check the cementing system reset sample; the socket slipped out of the cement mantle. Over 60 min delay. The event occurred during surgery. Patient experiences pain. Attempts have been made and no further information has been provided. | (b)(4). Report source ¿ foreign ¿ germany. Product was not returned or pictures not provided. Device evaluation could not be performed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. This device is used for treatment. Reported event is not related to medical condition. Review of medical record is not applicable. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15678275&pc=LOD
3006946279-2022-00106 | 09/22/2022 | 10/24/2022 | 32 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S46 / 22,2 | HIP PROSTHESIS | - | It was reported that the patient underwent revision surgery due to dislocation of the head and liner. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item# 802202202; lot# unknown; item name:femoral head 12/14 taper 22. 2 mm diameter +0 mm neck length. Foreign source: united kingdom. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15657110&pc=LPH
3006946279-2022-00103 | 10/05/2022 | 10/20/2022 | 15 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S56 / 28 | HIP PROSTHESIS | - | It was reported that four months post primary hip replacement, the patient had a fall and heard a clicking sound in his hip. Upon x-rays it was found that the left hip had dislocated and the patient underwent a revision surgery. Attempts have been made and no further information has been provided yet. | (b)(4). Concomitant medical devices: versys femoral head 28mm/-3. 5. Item# 8018-28-01, lot# 65270511. Foreign source: south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15638061&pc=LPH
3006946279-2022-00102 | 02/16/2022 | 10/20/2022 | 246 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent an initial right total knee arthroplasty. Subsequently, the patient had an arthroscopy with lysis of adhesions approximately 10 months later due to arthrofibrosis, stiffness, pain, difficulty walking, and adls. The patient's issues resolved approximately 2 weeks later. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices; femur cemented cruciate retaining (cr) standard nitrided right size 9, item# 42572606602, lot#64983080; tibia cemented 5 degree stemmed right size e, item# 42532007102, lot# 64831150; articular surface medial congruent (mc) right 10 mm height use with tibia sizes e-f/cr femur sizes 8-11, item# 42522100810, lot# 64781936; all poly patella cemented 32 mm diameter, item# 42540000032, lot# 64843478. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15637160&pc=MBB
3006946279-2022-00105 | 09/08/2022 | 10/19/2022 | 41 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CMNT R 1X40 JP | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that during surgery, the bone cement hardened within 1 minute after mixing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15630591&pc=MBB
3006946279-2022-00104 | 09/08/2022 | 10/19/2022 | 41 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CMNT R 1X40 JP | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that during surgery, the bone cement hardened within 1 minute after mixing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign ¿ japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15630588&pc=MBB
3006946279-2022-00101 | 09/12/2022 | 10/12/2022 | 30 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN CEMENT | CEMENT KNEE | - | It was reported that the patient underwent a revision surgery due to subsidence of the tibial baseplate 3 years post-implantation. During the surgery, it was noted that cement mantle has been broken. Patient was revised to primary tka. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical products: lps option femoral e-l, item# 00-5996-015-51, lot# 63742529; st prc tib plt size 5, item# 00-5980-047-01, lot# j6828721; lps flex art surf ef5 6 14, item# 00-5962-040-14, lot# 64971772; stem implant 15mmdx75mm, item# 00-5988-012-15, lot# 64946296; all poly pat comp 32dia, item# 00-5972-065-32, lot# 65341199; oxf twin-peg cmntd fem md pma, item# 161469, lot# 266810; oxf anat brg lt md size 4 pma, item# 159548, lot# 403220; oxf uni tib tray sza lm, item# 154718, lot# 150470. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15587396&pc=LOD
3006946279-2022-00100 | 09/06/2022 | 09/29/2022 | 23 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BIOMET BONE CEMENT R | - | It was reported that the patient underwent a revision surgery of the left knee due to loosening, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. Event Description No further event information available at the time of this report. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem lg item# 161470 lot# 933250, oxf anat brg lt lg size 3 pma item# 159554 lot# 151740, oxf uni tib tray sz d lm pma item# 154724 lot# 487670. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2022 - 00099. Event Description It was reported that the patient underwent a revision surgery of the left knee due to loosening, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507894&pc=LOD
3006946279-2022-00099 | 09/06/2022 | 09/29/2022 | 23 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BIOMET BONE CEMENT R | - | It was reported that the patient underwent a revision surgery of the left knee due to loosening, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem lg item# 161470 lot# 933250, oxf anat brg lt lg size 3 pma item# 159554 lot# 151740, oxf uni tib tray sz d lm pma item# 154724 lot# 487670. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3006946279 - 2022 - 00100. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507691&pc=LOD
3006946279-2022-00097 | 09/06/2022 | 09/28/2022 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOB PLUS BONE CEMENT 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that cement takes a long and inconsistent setting time to set and dry. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign-south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500699&pc=MBB
3006946279-2022-00098 | 02/23/2022 | 09/27/2022 | 216 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOB BONE CMT R-3 | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that during surgery, the consistency of the cement was found to be crumbly. New cement and stem were used to complete the surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Product discarded. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15493883&pc=MBB
3006946279-2022-00096 | 09/06/2022 | 09/27/2022 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOB PLUS BONE CEMENT 2X20-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that cement takes a long and inconsistent setting time to set and dry. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source foreign: south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15492825&pc=MBB
3006946279-2022-00095 | 08/18/2022 | 09/14/2022 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that during total knee arthroplasty cement hardened so fast that it was impossible to use the implant. Procedure has been finished with another implant and cement. Attempts have been made and no further information has been provided. | (b)(4) foreign - japan. Associated devices: (b)(4), product name: series a pat std 31 3 peg, lot number: 515450. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15417138&pc=MBB&device_sequence_no=1
3006946279-2022-00094 | 07/18/2022 | 09/09/2022 | 53 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN ZIMMER BIOMET REFOBACIN R+G CEMENT | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that the patient underwent an initial knee arthroplasty. Approximately 5 years post implantation, the patient was revised due to loosening of the tibial component. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign- united kingdom. Associated devices: reference number: (b)(4), item name: stemmed nonaugmentable tibial component option cr/ps/lps size 3 for cemented use only, lot number: 63569067. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15390735&pc=MBB&device_sequence_no=1
3006946279-2022-00092 | 08/11/2022 | 09/08/2022 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOB BONE CMT R-3 | ORTHOPAEDIC CEMENT, NON-ANTIMICROBIAL | - | It was reported that during the surgery cement started to harden approximately in less than 5 minutes. There was a need to use another product to finish the surgery. | (b)(4). Foreign report source: switzerland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15376320&pc=MBB&device_sequence_no=1
3006946279-2022-00090 | 08/11/2022 | 09/05/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the cement could not be processed because it solidified too quickly. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15354569&pc=MBB&device_sequence_no=1
3006946279-2022-00091 | 08/11/2022 | 09/05/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that the cement could not be processed because it solidified too quickly. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4) report source: foreign: germany. The device will not be returned for analysis; however, an investigation of the reported. Event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15354549&pc=MBB&device_sequence_no=1
3006946279-2022-00089 | 07/14/2022 | 08/18/2022 | 35 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INSERT / SIZE 44 / 22,2 | HIP PROSTHESIS | - | It was reported that expired product has been implanted. No harm or health consequences reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ spain. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15247168&pc=KWA&device_sequence_no=1
3006946279-2022-00088 | 07/16/2022 | 08/17/2022 | 32 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COCR FM HD D22,2/ 12-14/ 0/ M | HIP PROSTHESIS | - | It was reported that expired product has been implanted. Patient outcome unknown. Attempts have been made and no further information has been provided. | (b)(4). Report source ¿ foreign ¿ italy. Product was not returned or pictures not provided. Device evaluation could not be performed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Instruction for use has been reviewed and states that the expiration date is printed on the label. Do not use implants after the expiration date. This device is used for treatment. Complaints are monitored through monthly complaint review in order to identify potential adverse trends. Reported event is not related to medical condition. Review of medical record is not applicable. The root cause of the reported issue is attributed to a user error (expired product implanted). A communication to the customer has been performed. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15238946&pc=KWA&device_sequence_no=1
3006946279-2022-00087 | 07/13/2022 | 08/11/2022 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN PLUS BONE CEM 40 -3 -3 | BONE CEMENT | - | It was reported that patient has been implanted with cement approximately month and a half after expiry date. There is no indication of any patient harm at this moment. Attempts have been made and no further information has been provided. | (b)(4). Report source: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15207833&pc=LOD&device_sequence_no=1
3006946279-2022-00085 | 05/06/2022 | 07/29/2022 | 84 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT | - | It was reported that approximately two years after right knee revision, the patient needed medical intervention due experiencing pain and tenderness. Due diligence is in progress for this complaint; to date no additional information or product has been received. | Femoral distal augment cemented size 7, 7+ 10 mm thickness catalog # 42556606210 lot # 64336285. Femoral posterior augment cemented size 7, 7+ 5 mm thickness catalog # 42556806205 lot # 64661844. Femoral distal augment cemented size 7, 7+ 10 mm thickness catalog # 42556606210 lot # 64336285. Femoral posterior augment cemented size 7, 7+ 5 mm thickness catalog # 42556806205 lot # 64635880. Femur cemented standard right size 7 catalog # 42504606202 lot # 64779915. Stem extension 6mm offset splined uncemented 14 mm diameter +135 mm length catalog # 42560613514 lot # 64729231. Tibia fixed cemented right size e stem extension use required catalog # 42542007102 lot # 64454890. Tibial augment cemented half block right medial size ef 10 mm thickness catalog # 42555805410 lot # 64648417. Articular surface fixed bearing constrained condylar knee (cck) right 12 mm height use with tibia sizes e, f / revision femur sizes 7, 7+, 9, 9+ with catalog # 42522800712 lot # 64372524. Tibial augment cemented half block right lateral size ef 5 mm thickness catalog # 42555805205 lot # 64594693. Tibial central cone size x-small catalog # 42545000510 lot # 64500861. Customer has indicated that the product will not be returned because it remains implanted. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15127495&pc=MBB&device_sequence_no=1
3006946279-2022-00086 | 06/17/2022 | 07/27/2022 | 40 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported that bone cement hardened within one 1 minute after mixing during the surgery. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15109649&pc=MBB&device_sequence_no=1
3006946279-2022-00084 | 05/18/2022 | 07/26/2022 | 69 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED CUP S46 | HIP PROSTHESIS | - | It was reported that patient underwent revision surgery of the hip approximately four (4) month post-op due to the cup loosening after patient's fall. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15102316&pc=JDL&device_sequence_no=1
3006946279-2022-00083 | 06/08/2022 | 07/26/2022 | 48 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED CUP S46 | HIP PROSTHESIS | - | It was reported that patient underwent revision surgery of the hip approximately one (1) month post-op due to the cup loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15102211&pc=JDH&device_sequence_no=1
3006946279-2022-00082 | 06/14/2022 | 07/25/2022 | 41 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that a patient underwent initial right knee surgery. Subsequently, the patient was revised due to loosening approximately 2 years later. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: vanguard cr por fem-rt 67. 5; item# 183050; lot# 363630, biomet cc cruciate tray 67mm; item# 141232; lot# j6521694, series a pat std 31 3 peg; item# 184764; lot# 102270, vngd ant stblzd brg 10x67; item# 189040; lot# 374820. Due diligence is in progress for this complaint; to date no additional information or product has been received. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15092480&pc=LOD&device_sequence_no=1
3006946279-2022-00080 | 01/29/2019 | 07/22/2022 | 1270 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN REFOBACIN BONE CEMENT R | ORTHOPAEDIC CEMENT, ANTIMICROBIAL | - | It was reported via a clinical study that a patient underwent an initial left total knee arthroplasty. Approximately one year and two months post implantation, the patient underwent revision surgery due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: vngd ps tib brg 12x71/75mm; item#: 183642; lot#: 430510, tnbn pol cocr fin tib tray 71; item#: 141253tnbn; lot#: 2017051899, tinbn vg int ps anat fm l 70mm; item#: 183132tnbn; lot#: 3560859, tibial stem; item#: unknown; lot#: unknown, patella component; item#: unknown; lot#: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15081932&pc=MBB&device_sequence_no=1
3006946279-2022-00081 | 11/23/2021 | 07/21/2022 | 240 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the cement envelope was opened and sterility was lost. No adverse events have been reported as a result of the malfunction. Attempts have been made and no further information has been provided. | (b)(4) multiple mdr reports were filed for this event, please see associated reports: 3006946279-2021-00200 investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15074391&pc=MBB&device_sequence_no=1
3006946279-2022-00077 | 06/13/2022 | 07/20/2022 | 37 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BC R | BONE CEMENT, ANTIBIOTIC | - | It was reported that during the surgery when the bone cement was put into the femoral cavity, the patient's blood pressure dropped rapidly, and the patient has died. Attempts have been made and no further information has been provided. | (b)(4). Medical device: allen plug 20mm flange; item# 00-8011-020-20; lot# 65337365. Report source- foreign: japan. Customer has indicated that the product will not be returned to zimmer biomet for investigation. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15064912&pc=MBB&device_sequence_no=1
3006946279-2022-00079 | 06/24/2022 | 07/20/2022 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | ORTHOPAEDIC CEMENT, ANTIBIOTIC | - | It was reported that the cement hardened too fast. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15062227&pc=MBB&device_sequence_no=1
3006946279-2022-00078 | 06/24/2022 | 07/19/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that the cement was setting up super fast. No known impact or consequences to the patient. Attempts have been made and no further information has been provided. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15049919&pc=MBB&device_sequence_no=1
3006946279-2022-00075 | 12/14/2021 | 07/14/2022 | 212 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that a patient underwent an initial right total knee arthroplasty. Subsequently, the patient was revised approximately one year and two months post-implantation due to pain, swelling, and possible loosening of the tibial component. During revision, no gross loosening was found in the tibial or femoral components. Mild synovitis was noted. The patella was left intact, all other components were revised without complication. | (b)(4). Medical products: articular surface fixed bearing posterior stabilized (ps) right 11 mm height: catalog# 42521400511, lot # 63750891; tibia cemented 5 degree stemmed right size d: catalog# 42532006702, lot# 64686661; femur cemented posterior stabilized (ps) narrow right size 6: catalog# 42500006002, lot# 64472820; all poly patella cemented 29 mm diameter: catalog# 42540000029, lot# 64740327. Multiple mdr reports have been filed for this event. Please see associated report: 3006946279-2022-00043. Device evaluation could not be performed as no product was returned nor were pictures provided. Device history record was reviewed and no discrepancies related to the reported event were found. Medical records were reviewed and show that patient started to have pain 2 years after the initial surgery. The revision surgery notes states that no loosening was found. Nothing which could explain the pain was found. Root cause was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15013139&pc=LOD&device_sequence_no=1
3006946279-2022-00076 | 12/24/2019 | 07/14/2022 | 933 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that a patient underwent an initial left total knee arthroplasty without complications. Subsequently, the patient has reported ongoing pain and swelling for over two years since implantation. Additionally, the patient reported issues with bending and the inability to kneel on one of the knees. It is unknown at this time which extremity is experiencing range of motion complications as the patient has had bilateral knee replacements. | This is a combination product (b)(4). Multiple mdr reports have been filed for this event. Please see associated report: 3006946279-2022-00046. Medical product: femur cemented cruciate retaining (cr) standard left size 7: catalog#42502606201, lot#64228083; natural tibia cemented 5 degree stemmed left size e: catalog#42532007101, lot#64417330; articular surface fixed bearing ultracongruent (uc) left 14 mm height: catalog#42512200514, lot#64279930; all-poly patella cemented 32 mm diameter#42540200032, lot#64540883. Device evaluation could not be performed as no product was returned nor were pictures provided. Device history record was reviewed and no discrepancies related to the reported event were found. Root cause was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15013121&pc=MBB&device_sequence_no=1
3006946279-2022-00073 | 11/23/2021 | 07/13/2022 | 232 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the cement envelope was found opened and sterility was lost. No adverse events have been reported as a result of the malfunction. | Cmp-(b)(4). Concomitant medical products - medical devices: biomet bone cement r 40 -3; item# 3035890011-3; lot# a849ah1704 biomet bone cement r 40 -3; item# 3035890011-3; lot# a849ah1704. Report source: foreign:(b)(4). Multiple mdr reports were filed for this event, associated reports: multiple mdr reports were filed for this event, associated reports: 3006946279-2021-00199, 3006946279-2021-00074. Device analysis: the product was returned and lab analysis was performed. The product analysis shows that: 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 1cm discontinuous at the upper left side of the inner pouch. 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 4cm discontinuous at the upper left side of the inner pouch. 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 1,5cm discontinuous at the upper left side of the inner pouch. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14998133&pc=LOD&device_sequence_no=1
3006946279-2022-00074 | 11/23/2021 | 07/13/2022 | 232 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the cement envelope was found opened and sterility was lost. No adverse events have been reported as a result of the malfunction. | (b)(4). Medical devices: biomet bone cement r 40 -3; item# 3035890011-3; lot# a849ah1704 biomet bone cement r 40 -3; item# 3035890011-3; lot# a849ah1704. Foreign: italy. Multiple mdr reports were filed for this event, please see associated reports: multiple mdr reports were filed for this event, please see associated reports: 3006946279-2021-00199 and 3006946279-2021-00073. Device analysis: the product was returned and lab analysis was performed. The product analysis shows that: 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 1cm discontinuous at the upper left side of the inner pouch. 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 4cm discontinuous at the upper left side of the inner pouch. 3035890011-3, a849ah1704, biomet bone cement r 40 -3 the sealing is around 1,5cm discontinuous at the upper left side of the inner pouch. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14998131&pc=LOD&device_sequence_no=1
3006946279-2022-00072 | 12/03/2015 | 07/11/2022 | 2412 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient underwent a knee replacement approximately 19 months post implantation due to migration and pain. It was reported as part of a retrospective study, that the patient had suffered from malalignment of the tibia component causing external rotation of lower leg and pain. No other contributing factors other than the malalignment led to the revision. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: 00598800300 - stemmed tibial component precoat a/p wedged for cemented use only size 3 - unknown; 00599401492 - femoral component option for cemented use only size d right - unknown; 00599403012 - articular surface with locking screw size yellow/c,d 12 mm height for use with femoral c,d - unknown; unknown refobacin cement. Source foreign: denmark. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14975907&pc=MBB&device_sequence_no=1
3006946279-2022-00067 | 06/12/2022 | 07/07/2022 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE PRESSE PLATEAU SUP. | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that part of the instrument has broken off. No harm to patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942879&pc=LXH&device_sequence_no=1
3006946279-2022-00070 | 06/07/2022 | 07/07/2022 | 30 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOB BONE CMT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that during a revision surgery of a competitor's implant, it seemed as if not cemented. Investigation has been completed. Attempts have been made and no further information has been provided. | (b)(4). Foreign source: (b)(6). Product was not returned or pictures not provided. Device evaluation could not be performed. The dhr review can¿t be performed as the item/batch number was not provided. This device is used for treatment. Reported event is not related to medical condition. Review of medical record is not applicable. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942147&pc=MBB&device_sequence_no=1
3006946279-2022-00069 | - | 07/06/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported patient is experiencing pain and bone scan displayed implant loosening approximately one year post implantation. No revision is planned at this time. Attempts to obtain additional information have been made; however, no more is available. | (b)(4). Concomitant medical products: partial femur cemented; item# 42558000402; lot # 64587381; partial tibial cemented; item# 42538000702; lot# 64686959; psn pk ve ply rl sz g 9mm; item# 42528800709; lot # 64272611; biomet bc r 1x40 us; item# 110035368; lot # az44bj2604. Product was not returned or pictures not provided. Device evaluation could not be performed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. This device is used for treatment. Medical records were reviewed and no clues to the cause of the event were found. A definitive root cause cannot be determined. No further actions have been initiated as a result of the reported event investigation. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14929991&pc=LOD&device_sequence_no=1
3006946279-2022-00068 | - | 07/06/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported patient is experiencing pain and bone scan displayed implant loosening approximately one year post implantation. No revision is planned at this time. Attempts to obtain additional information have been made; however, no more is available. | (b)(4). Medical product: partial femur cemented; item# 42558000402; lot # 64587381; partial tibial cemented; item# 42538000702; lot# 64686959; psn pk ve ply rl sz g 9mm; item# 42528800709; lot # 64272611; biomet bc r 1x40 us; item# 110035368; lot # az44bj2604. Product was not returned or pictures not provided. Device evaluation could not be performed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. This device is used for treatment. Medical records were reviewed and no clues to the cause of the event were found. A definitive root cause cannot be determined. No further actions have been initiated as a result of the reported event investigation. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14929970&pc=LOD&device_sequence_no=1
3006946279-2022-00066 | 05/27/2022 | 06/30/2022 | 34 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UNKNOWN AVANTAGE 3P CUP | HIP PROTHESIS | - | It was reported that patient underwent a revision surgery due to (b)(6). Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Source foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14871124&pc=LPH&device_sequence_no=1
3006946279-2022-00060 | 05/27/2022 | 06/29/2022 | 33 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE PEG EXTRACTOR | HIP INSTRUMENT | - | It was reported that during a surgery the extraction tool broke. No known impact or consequence to the patient. Attempts have been made and no further information has been provided. Investigation has been completed. | (b)(4). Foreign: south africa. Pictures were provided. The reported event was confirmed. Device history record (dhr) was reviewed and no discrepancies related to the reported event were found. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14852601&pc=HWB&device_sequence_no=1
3006946279-2022-00065 | 02/02/2021 | 06/23/2022 | 506 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent an initial left knee arthroplasty. Subsequently, the patient underwent revision due to infection, approximately three (3) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: biomet bc r 1x40 us item# 110035368, lot# 004cac2505. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3006946279-2022-00059. 3006946279-2022-00065. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14786858&pc=LOD&device_sequence_no=1
3006946279-2022-00059 | 02/02/2021 | 06/23/2022 | 506 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent an initial left knee arthroplasty. Subsequently, the patient underwent revision due to infection, approximately three (3) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: biomet bc r 1x40 us. Item# 110035368. Lot# 004cac2505. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3006946279-2022-00059, 3006946279-2022-00065. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14786852&pc=LOD&device_sequence_no=1
3006946279-2022-00062 | 12/08/2015 | 06/22/2022 | 2388 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient experienced pain approximately 8 months after the initial procedure. Implant loosening was confirmed. The patient underwent a right knee revision due to pain and tibial loosening approximately 1 year post-operatively. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: optipac 40 refobacin revision; catalog# 4730501163-1; lot# 235ca03910 temmed tibial component precoat a/p wedged for cemented use only size 3; catalog#00598800300; lot#62834341stem extension straight 13mm dia x 100mm length(combined length 145mm); catalog#00598801013; lot#62846723articular surface with locking screw size striped yellow/e,f 14 mm height for use with femoral e,f; 00599403214; lot#62656678femoral component option for cemented use only size f right; catalog#00599401692; lot#62851943. Foreign: denmark. Multiple mdr reports were filed for this event, please see associated reports: (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14772012&pc=MBB&device_sequence_no=1
3006946279-2022-00058 | 12/08/2015 | 06/22/2022 | 2388 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient experienced pain approximately 8 months after the initial procedure. Implant loosening was confirmed. The patient underwent a right knee revision due to pain and tibial loosening approximately 1 year post-operatively. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: stemmed tibial component precoat a/p wedged for cemented use only size 3; catalog#00598800300; lot#62834341. Stem extension straight 13mm dia x 100mm length(combined length 145mm); catalog#00598801013; lot#62846723. Articular surface with locking screw size striped yellow/e,f 14 mm height for use with femoral e,f; 00599403214; lot#62656678. Femoral component option for cemented use only size f right; catalog#00599401692; lot#62851943. Report source: foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14767054&pc=MBB&device_sequence_no=1
3006946279-2022-00056 | 05/16/2022 | 06/14/2022 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that the bone cement cured faster than usual. Bone cement was in the refrigerator until use. However, it hardened in about 2 minutes being placed in the syringe after mixing. Surgery was completed using a back-up product. No patient or user harm reported. Attempts have been made and no further information has been provided. | (b)(4). Foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14683161&pc=LOD&device_sequence_no=1
MW5110257 | 02/28/2022 | 06/10/2022 | 102 | malfunction | Voluntary | BIOMET FRANCE | - | BIOMET 4.0 CANCELLOUS PARTIALLY THREADED 70MM | SCREW, FIXATION, BONE | - | A 2. 5 mm biomet screw broke off in the bone due to the density of the bone. | - | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14676208&pc=HWC&device_sequence_no=1
3006946279-2022-00054 | 05/12/2022 | 06/08/2022 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOB BONE CMT R-3 | BONE CEMENT | - | It was reported that an expired product has been used in a surgery. The expiration date was one month before the product had been used. No known impact or consequence to the patient. Attempts have been made and no further information has been provided. | (b)(4). Foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629812&pc=LOD&device_sequence_no=1
3006946279-2022-00055 | - | 06/08/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VANGUARD ROCC IMPLANT | KNEE PROSTHESIS | - | It was reported that patient underwent a knee arthroplasty surgery. The patient sent an email to check the materials included in the implant as they are suspecting an allergic reaction to the implant. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629811&pc=JWH&device_sequence_no=1
3006946279-2022-00053 | 05/04/2022 | 06/03/2022 | 30 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a revision due to the loosening of the tibial tray and pain in the left knee, approximately one year post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: oxf uni tib tray sz b lm pma; item# 154720; lot# 783570. Oxf anat brg lt sm size 3 pma; item#159540; lot#075460. Oxf twin-peg cmntd fem sm pma; item#161468; lot#953870. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14593721&pc=LOD&device_sequence_no=1
3006946279-2022-00051 | 03/25/2022 | 06/02/2022 | 69 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a right knee arthroplasty revision to address a deep, staphylococcus hominis and staphylococcus epidermidis infection seventeen days post-operatively. Attempts have been made, however, no additional information is available at this time. | (b)(4). Concomitant medical products: tibia cemented 5 degree stemmed right size c; item# 42532006402; lot# 64884536; femur cemented cruciate retaining (cr) narrow right size 5; item# 42502005802; lot# 63717984; articular surface medial congruent (mc) right 10 mm height use with tibia sizes c-d/cr; femur sizes 4-5; item# 42522100310; lot# 64557869; stem extension tapered cemented 14 mm diameter +30 mm length; item# 42557000114; lot# 65264231; all poly patella cemented 29 mm diameter; item# 42540000029; lot# 65134509; refobacin bc r 1x40 us; item# 110034355; lot# az47ba1511. The complainant has not yet indicated whether the product will be returned to zimmer biomet for investigation. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14574442&pc=MBB&device_sequence_no=1
3006946279-2022-00052 | 03/25/2022 | 06/02/2022 | 69 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a right knee arthroplasty revision to address a deep, staphylococcus hominis and staphylococcusepidermidis infection seventeen days post-operatively. Attempts have been made, however, no additional information is available at this time. | (b)(4). Medical decive: tibia cemented 5 degree stemmed right size c; item# 42532006402; lot# 64884536. Femur cemented cruciate retaining (cr) narrow right size 5; item# 42502005802; lot# 63717984. Articular surface medial congruent (mc) right 10 mm height use with tibia sizes c-d/cr femur sizes 4-5; item# 42522100310; lot# 64557869. Stem extension tapered cemented 14 mm diameter +30 mm length; item# 42557000114; lot# 65264231. All poly patella cemented 29 mm diameter; item# 42540000029; lot# 65134509. Refobacin bc r 1x40 us; item# 110034355; lot# k1903z22ca. The complainant has not yet indicated whether the product will be returned to zimmer biomet for investigation. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14574398&pc=MBB&device_sequence_no=1
3006946279-2022-00049 | 03/18/2022 | 06/02/2022 | 76 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient has been revised approximately 1 month after the initial surgery due to staph aureus infection. No additional information is available at this time. | (b)(4). Concomitant medical products: tibia cemented 5 degree stemmed right size e; catalog#42532007102; lot#65199572; femur cemented posterior stabilized (ps) narrow right size 7;catalog#42500006202; lot#65139649; articular surface fixed bearing constrained posterior stabilized (cps) right 14 mm height; use with tibia sizes e-f / ps femur sizes 6-9; catalog#42522600714; lot#64945522; stem extension tapered cemented 14 mm diameter +30 mm length; catalog#42557000114; lot#65264429; all poly patella cemented 32 mm diameter; catalog#42540000032; lot#65134405. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14573807&pc=LOD&device_sequence_no=1
3006946279-2022-00050 | 03/18/2022 | 06/02/2022 | 76 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient has been revised approximately 1 month after the initial surgery due to staph aureus infection. No additional information is available at this time. | (b)(4). Tibia cemented 5 degree stemmed right size e; catalog#42532007102; lot#65199572; femur cemented posterior stabilized (ps) narrow right size 7;catalog#42500006202; lot#65139649; articular surface fixed bearing constrained posterior stabilized (cps) right 14 mm height; use with tibia sizes e-f / ps femur sizes 6-9; catalog#42522600714; lot#64945522; stem extension tapered cemented 14 mm diameter +30 mm length; catalog#42557000114; lot#65264429; all poly patella cemented 32 mm diameter; catalog#42540000032; lot#65134405. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14573757&pc=LOD&device_sequence_no=1
3006946279-2022-00048 | 04/15/2022 | 06/01/2022 | 47 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BIOMET BONE CEMENT R | - | It was reported that the patient underwent a revision approximately 3. 5 years post implantation due to pain, swelling, and instability. During the surgery, it was noted that the patellar component had incomplete coverage and there was a small area of necrosis on the center of the patella. The articular surface and patella were exchanged and the surgery was complete without complications. No additional information is available at this time. | (b)(4). Concomitant medical products: articular surface fixed bearing posterior stabilized (ps) left 10 mm height; item#: 42512400710; lot#: 64162098. All poly patella cemented 32 mm diameter; item#: 42540000032; lot#: 64209106. 2. 5 mm female hex screw 25 mm length; item#: 42509902525; lot#: 64249167. Headless trocar drill pin 3. 2 mm diameter 75 mm length; item#: 00590102000; lot#: 64229348. Femur cemented posterior stabilized (ps) standard left size 9; item#: 42500606601; lot#: 64050573. Biomet bc r 1x40 us; item#: 110035368; lot#: 745cah2108. Femoral set screw hex driver torque limiting size 2. 0 mm; item#: 00598707100; lot#: unknown. Tibia cemented 5 degree stemmed left size f; item#: 42532007501; lot#: 64072924. Stem extension tapered cemented 14 mm diameter +30 mm length; item#: 42557000114; lot#: 64229115. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14563990&pc=LOD&device_sequence_no=1
3006946279-2022-00047 | 04/15/2022 | 06/01/2022 | 47 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a revision approximately 3. 5 years post implantation due to pain, swelling, and instability. During the surgery, it was noted that the patellar component had incomplete coverage and there was a small area of necrosis on the center of the patella. The articular surface and patella were exchanged and the surgery was complete without complications. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: articular surface fixed bearing posterior stabilized (ps) left 10 mm height; item# 42512400710; lot# 64162098. All poly patella cemented 32 mm diameter; item# 42540000032; lot# 64209106. 2. 5 mm female hex screw 25 mm length; item# 42509902525; lot# 64249167. Headless trocar drill pin 3. 2 mm diameter 75 mm length; item# 00590102000; lot# 64229348. Femur cemented posterior stabilized (ps) standard left size 9; item# 42500606601; lot# 64050573. Biomet bc r 1x40 us; item# 110035368; lot# 803bai1807; femoral set screw hex driver torque limiting size 2. 0 mm; item# 00598707100; lot# unknown. Tibia cemented 5 degree stemmed left size f; item# 42532007501; lot# 64072924; stem extension tapered cemented 14 mm diameter +30 mm length; item# 42557000114; lot# 64229115. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14563884&pc=LOD&device_sequence_no=1
3006946279-2022-00046 | 12/24/2019 | 05/24/2022 | 882 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that a patient underwent an initial left total knee arthroplasty, without complications, during which 2 items of cement were used. Subsequently, the patient reported ongoing pain and swelling. Additionally, he reported issues bending and inability to kneel on one of his knees. It is unknown at this time, which extremity is experiencing range of motion complications as the patient has had bilateral knee replacements. | This is a combination product (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14486429&pc=MBB&device_sequence_no=1
3006946279-2022-00045 | 04/20/2022 | 05/24/2022 | 34 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported, that patient had a cardiopulmonary arrest after cement use. Surgeon was used cement during operation of the hemi arthroplasty. After stem fixation, the patient became hypotensive, dyspneic, and went into cardiopulmonary arrest, resulting in the death of the patient. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14486074&pc=LOD&device_sequence_no=1
3006946279-2022-00043 | 12/14/2021 | 05/21/2022 | 158 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that a patient underwent an initial right total knee arthroplasty performed (b)(6) 2020. Subsequently, the patient was revised (b)(6) 2022 due to pain, swelling, and possible loosening of the tibial component. During revision, no gross loosening was found in the tibial or femoral components. Mild synovitis was noted. The patella was left intact, all other components were revised without complication. | (b)(4). List of associated devices: articular surface fixed bearing posterior stabilized (ps) right 11 mm height, reference 42521400511, batch 63750891. Tibia cemented 5 degree stemmed right size d, reference 42532006702, batch 64686661. Femur cemented posterior stabilized (ps) narrow right size 6, reference 42500006002, batch 64472820. All poly patella cemented 29 mm diameter, reference 42540000029, batch 64740327. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14466882&pc=LOD&device_sequence_no=1
3006946279-2022-00044 | 05/05/2022 | 05/20/2022 | 15 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that one expired product has been consumed in a surgery. The product expired on 30. 04. 2022 and was used in a surgery on (b)(6) 2022. No further information available regarding patient health consequences. | This is a combination product. (b)(4). Report source, foreign - event occurred in spain. No non conformity or deviation was observed which could be linked to the event described in the complaint. The review of the sterilization certificate indicates that the products were sterilized according to the specification. A picture has been received, showing the product label with the expiration date of april 30, 2022. On this picture, it is also indicated that the product has been implanted on (b)(6) 2022. Therefore, the reported event is confirmed. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14460710&pc=MBB&device_sequence_no=1
3006946279-2022-00042 | 04/21/2022 | 05/20/2022 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOB PLUS BONE CEMENT 1X40-3 | CEMENT PACK (E) : BPBC | - | It was reported that during a knee implantation surgery the cement paste hardened too fast, the setting time was reported as too short and therefore the patella could not be implanted; alternative devices were used. No further outcome or patient harm was reported. Attempts have been made and no further information has been provided. | (b)(4). Event happened in chile. Customer has indicated that the product was discarded so it will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14453372&pc=MBB&device_sequence_no=1
3006946279-2022-00041 | 04/07/2022 | 05/07/2022 | 30 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that one expired product has been consumed in a surgery when the product was being invoicing. The product expired on (b)(6) 2022. And was used in a surgery on (b)(6) 2022. Neither impact in patient nor claim or incidence has been informed. No adverse health consequence has been reported as result of the malfunction. | This is a combination product. Cmp- (b)(4). Report source, foreign - event occurred in spain. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14314818&pc=MBB&device_sequence_no=1
3006946279-2022-00040 | - | 05/07/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 | BONE CEMENT | - | It was reported that a patient underwent a right knee procedure on (b)(6), 2021. Subsequently, the patient claims to suffer from pain, especially at night, just above the knee on the right side. Patient states it feels like someone is stabbing him with a sharp knife. | (b)(4). List of associated devices: refobacin bone cement r 1x40, reference (b)(4), batch z46aaj2605. Psn straight hyb st 14 x +30 m, reference (b)(4), batch 65039977. Persona cem fem/cem tib/ve surf/std pat, reference (b)(4), batch not communicated. Nexgen all-poly patella, 32mm, reference (b)(4), batch 64825513. Psn tib stm 5 deg sz g r , reference (b)(4) , batch 64762099. Psn fem ps cmt ccr std sz9 r, reference (b)(4), batch 64958869. Psn asf ps 10mm ve r 6-9 gh, reference (b)(4), batch 64747050. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14314408&pc=LOD&device_sequence_no=1
3006946279-2022-00038 | - | 05/07/2022 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 | BONE CEMENT, ANTIBIOTIC | - | T was reported that a patient underwent a right knee procedure on (b)(6) 2021. Subsequently, the patient claims to suffer from pain, especially at night, just above the knee on the right side. Patient states it feels like someone is stabbing him with a sharp knife. Patient involvement-knee pain. | This is a combination product. (b)(4) concomitant medical products: list of associated devices: biomet bc r 1x40 us, reference 110035368, batch az44cj2604, psn straight hyb st 14 x +30 m, reference (b)(4), batch 65039977, persona cem fem/cem tib/ve surf/std pat, reference (b)(4), batch not communicated, nexgen all-poly patella, 32mm, reference (b)(4), batch 64825513, psn tib stm 5 deg sz g r , reference (b)(4), batch 64762099, psn fem ps cmt ccr std sz9 r, reference (b)(4), batch 64958869, psn asf ps 10mm ve r 6-9 gh, reference (b)(4), batch 64747050. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14314347&pc=MBB&device_sequence_no=1
3006946279-2022-00039 | 03/23/2022 | 05/07/2022 | 45 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient underwent an initial right total knee arthroplasty on (b)(6) 2019. Subsequently, the patient was revised on (b)(6) 2022 due to tibial plateau loosening. The tibia and polyethylene were exchanged. | This is a combination product (b)(4). List of associated devices: psn tib stm 5 deg sz e l, reference 42-5320-071-01, batch 63974344. Psn tpr st 14 x +30 mm, reference 42-5570-001-14, batch 64114697. Psn fem ps cmt ccr std sz 9 l, reference 42-5006-066-01, batch 63944679. All poly pat ve 35 mm dia, reference 42-5402-000-35, batch 63924950. Psn asf ps 10mm ve l 6-9 ef, reference 42-5124-007-10, batch 64015908. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14314323&pc=MBB&device_sequence_no=1
3006946279-2022-00037 | 04/04/2022 | 05/04/2022 | 30 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN PLUS BONE CEMENT | BONE CEMENT, ANTIBIOTIC | - | It was reported that an expired optipac was used in surgery. No known adverse health consequences has been reported as the result of the event. | This is a combination product. Cmp- (b)(4). Report source, foreign - event occurred in poland. Product was not returned or pictures were not provided. Visual and dimensional evaluations could not be performed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the product was sterilized according to the specification. This device is used for treatment. Reported event is not related to a combination of products, therefore, a compatibility review is not applicable. Complaint history identified additional similar complaints for the reported item and no additional complaints for the reported part/lot combination. Complaints are monitored through monthly complaint review in order to identify potential adverse trends. Reported event is not related to medical condition. Review of medical record is not applicable. The root cause of the reported issue is attributed to a human error (product implanted after expiry date). A communication has been done to the customer. No further actions have been initiated as a result of the reported event investigation. The complaint is closed but could be reopened if new information is received later. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14277292&pc=MBB&device_sequence_no=1
3006946279-2022-00036 | 03/21/2022 | 05/04/2022 | 44 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT | BONE CEMENT, ANTIBIOTIC | - | It was reported that an expired optipac was used in surgery. The product expired on (b)(6), 2022 and was used in a surgery on (b)(6), 2022. No adverse health outcome has been reported as the result of the event. | This is a combination product. (b)(4). Foreign source: event occurred in spain. Pictures were provided. The reported event was confirmed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the product was sterilized according to the specification. This device is used for treatment. Reported event is not related to a combination of products, therefore, a compatibility review is not applicable. Complaint history identified additional similar complaints for the reported item and no additional complaints for the reported part/lot combination. Complaints are monitored through monthly complaint review in order to identify potential adverse trends. Reported event is not related to medical condition. Review of medical record is not applicable. The root cause of the reported issue is attributed to a human error (product implanted after expiry date). A communication has been done to the customer. No further actions have been initiated as a result of the reported event investigation. The complaint is closed but could be reopened if new information is received later. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14276278&pc=MBB&device_sequence_no=1
3006946279-2022-00032 | 03/21/2022 | 04/14/2022 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported, that the cement hardened in about 4 minutes and 30 seconds after stirring. In addition, the target product was discarded. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in japan. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14110710&pc=LOD&device_sequence_no=1
3006946279-2022-00033 | 03/07/2022 | 04/14/2022 | 38 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It was reported that bone cement cured faster than usual. Bone cement was in the refrigerator until just before. However, it hardened about 2 minutes after being placed in the syringe after mixing. About total 4 ~ 5 minutes. This product was not used and was discarded. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in japan. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14110704&pc=LOD&device_sequence_no=1
3006946279-2022-00034 | 02/18/2022 | 04/14/2022 | 55 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN PLUS BONE CEMENT 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that during placement of a left xnov knee prostesis, implants had to be cemented. Cement preparation was done according to the preparation instructions. After 8 min, the cement was still liquid. The cement was unusual (very liquid. Cement drying time was too long (27 min). No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in france. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14110694&pc=MBB&device_sequence_no=1
3006946279-2022-00035 | 03/24/2022 | 04/14/2022 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported, that: bone cement hardened in about 7 minutes. Oss implant stuck during insertion with bone cement. It was immediately removed and re-cemented. No adverse events have been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in japan. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14110668&pc=MBB&device_sequence_no=1
3006946279-2022-00031 | 03/16/2022 | 04/13/2022 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS VARIZED FEMORAL STEM SIZE -2 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported an implantation of an expired gts stem. On (b)(6), 2022, a total hip endoprosthesis was on the schedule at the erding clinic. The operation has started and when it came to the implantation it was discovered that the implant had expired on (b)(6), 2022. | (b)(4). Foreign source: event occurred in germany. Product was not returned or pictures were not provided. Visual and dimensional evaluations could not be performed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. This device is used for treatment. Reported event is not related to a combination of products, therefore, a compatibility review is not applicable. Complaint history identified no other similar event for the reported item and part/lot combination. Reported event is not related to medical condition. Review of medical record is not applicable. Instruction for use has been reviewed and it is stated the following point: implants are supplied sterile and packaged individually in double wrapping. Sterilization is carried out by exposure to a minimum dose of 25 kgy of gamma irradiation. The expiration date is printed on the label. Do not use implants after the expiration date. The root cause of the reported issue is attributed to user error (expired product implanted). A communication to the customer has been done. No further actions have been initiated as a result of the reported event investigation. Complaints are monitored through monthly complaint review in order to identify potential adverse trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14101565&pc=LPH&device_sequence_no=1
3006946279-2022-00030 | 04/15/2018 | 03/30/2022 | 1445 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | STANDARD FEMORAL STEM EXCEPTION -CEMENTLESS- 5°42- 12/14 - RIGHT S4 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that during a post market clinical follow-up study on the avantage reload cup, the patient fall and has a periprostetic femoral fracture. Revision with explantation of the femoral stem (exception stem) was performed. Acetabular component was left intact (avantage reload). | (b)(4). List of associated devices: avantage reload cup ti ha s50, reference p0460p50, batch 0000861734. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13949528&pc=LPH&device_sequence_no=1
3006946279-2022-00029 | 02/07/2022 | 03/23/2022 | 44 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS HUMERAL HEAD | TOTAL SHOULDER PROSTHESIS | - | It was reported, that patient fell and broke humerus. Patient underwent revision surgery. | (b)(4). Associated device: tess corolla, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). Product was not returned or pictures were not provided. Visual and dimensional evaluations could not be performed. The dhr review can¿t be performed as the item/batch number was not provided. This device is used for treatment. The products involved in the complaint are compatible. Complaint history search could not be performed as the part and lot information was not provided. X-ray was reviewed and show that the humerus broke below the level of the implant. The most probable root cause of the issue would be the patient fall. No further actions have been initiated as a result of the reported event investigation. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13867711&pc=KWS&device_sequence_no=1
3006946279-2022-00028 | 02/07/2022 | 03/23/2022 | 44 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS COROLLA | TOTAL SHOULDER PROSTHESIS | - | It was reported, that patient fell and broke humerus. Patient underwent revision surgery. | (b)(4). Concomitant medical products: - associated device: tess humeral head, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). Product was not returned or pictures were not provided. Visual and dimensional evaluations could not be performed. The dhr review can¿t be performed as the item/batch number was not provided. This device is used for treatment. The products involved in the complaint are compatible. Complaint history search could not be performed as the part and lot information was not provided. X-ray was reviewed and show that the humerus broke below the level of the implant. The most probable root cause of the issue would be the patient fall. No further actions have been initiated as a result of the reported event investigation. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13867565&pc=KWS&device_sequence_no=1
3006946279-2022-00026 | 03/12/2022 | 03/23/2022 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that hardening time of bone cement ended already after 5 min. No adverse events have been reported as a result of the malfunction. | This is a combination product: (b)(4). Report source: (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13860160&pc=MBB&device_sequence_no=1
3006946279-2022-00027 | 03/12/2022 | 03/23/2022 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that hardening time of bone cement ended already after 5 min. No adverse events have been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in austria. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13859949&pc=MBB&device_sequence_no=1
3006946279-2022-00022 | 02/21/2022 | 03/11/2022 | 18 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CHROME COBALT FEMORAL HEAD D32 / 12-14 / -4 / S | METALLIC FEMORAL HEAD PROSTHESIS | - | It was reported that a product expired on (b)(6) 2022 and was used in a surgery on (b)(6) 2022. As per hospital statement, the implant was opened by mistake in a surgery prior the expiration date and following the instructions of the hospital it was sterilized again (the hospital aimed to maintain the sterility for 6 months more) and it was implanted in the surgery. This decision was taken by the hospital, who assessed that in no moment the sterility was lost and it did not induce any type of the infection to the patient. No adverse health consequence has been reported as the result of the event. | (b)(4). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. No other similar event was recoded on the reference. Investigation results concluded that the reported event was due to user error (expired product implanted). A summary of the investigation was sent to the complainant conveying proper surgical technique and use of the device. If any further information is found which would change or alter any conclusions or information, the complaint will be reopened. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13729656&pc=LPH&device_sequence_no=1
3006946279-2022-00025 | 02/09/2022 | 03/10/2022 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient underwent an initial knee arthroplasty on (b)(6) 2020. Subsequently, the patient was revised on (b)(6) 2022 due to pain and loosening. | This is a combination product. (b)(4). List of associated devices: vanguard cr ilok fem-rt 62. 5, reference 183006, batch j6852735, biomet cc i-beam tray 67mm, reference 141222, batch j6785127, series a pat std 31 3 peg, reference 184764, batch 739280, vngd ant stblzd brg 12x67, reference 189042, batch 542380. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13720023&pc=MBB&device_sequence_no=1
3006946279-2022-00024 | 02/25/2022 | 03/10/2022 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that during hip artroplasty, a cement cup was implanted. A first attempt was made, but the cement did not bond with the acetabulum. The acetabular cup together with the cement tipped out of the acetabulum. A second attempt was made with another cement batch, the cement did not bond with the acetabulum, and the implant again tilted out of the acetabulum. Eventually, a cementless cup was implanted. The events result in a 45 minutes delay in procedure. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13720010&pc=MBB&device_sequence_no=1
3006946279-2022-00023 | 02/25/2022 | 03/10/2022 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that during hip artroplasty, a cement cup was implanted. A first attempt was made, but the cement did not bond with the acetabulum. The acetabular cup together with the cement tipped out of the acetabulum. A second attempt was made with another cement batch, the cement did not bond with the acetabulum, and the implant again tilted out of the acetabulum. Eventually, a cementless cup was implanted. The events result in a 45 minutes delay in procedure. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13720007&pc=MBB&device_sequence_no=1
3006946279-2022-00021 | 01/14/2022 | 03/10/2022 | 55 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported, that an initial right total knee arthroplasty was performed. Subsequently, the patient was revised due to painful aseptic loosening. During the revision, the tibial and femoral components were noted to be grossly loose with minimal ingrowth/cement adhering to the bone. The poly was also noted to have pitting. The articular surface and tibial and femoral components were replaced without complication. The patellar button was well-fixed, stable, and left in place. | (b)(4). Medical have been received, mentioning a loosening of the implant, however, no x-ray has been received showing the event. Therefore, the reported event could not be confirmed. Device manufacturing quality record indicate that the released product met all requirements to perform as intended. The products involved in the complaint are compatible. No other similar event was recoded on the batch reported. The medical record of the initial hip arthroplasty has been reviewed and nothing noticeable which could explain the loosening and the pain has been found. Then, the review of the visit notes dated from (b)(6) 2021 states that the patient suffered from pain at the right knee level since (b)(6) 2021. Several x-ray have been received and a radiolucent line has been noticed on the medical side of the tibial ap view. No traces of infection have been found. Finally, the review of the revision surgery showed that it was a minimal ingrowth/cement adhering which adhered the femoral implant to the bone. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, the complaint will be reopened. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13717821&pc=LOD&device_sequence_no=1
3006946279-2022-00020 | 02/03/2022 | 03/07/2022 | 32 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that when the surgeon was mixing the optipac, the viscosity was lower than normal. This product was used as is. The surgeon confirmed after use that the implant had been well fixation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13684062&pc=LOD&device_sequence_no=1
3006946279-2022-00018 | 12/03/2021 | 03/04/2022 | 91 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that an initial left total knee arthroplasty performed in 2007 with unknown product was subsequently revised (b)(6) 2019 due to persistent pain. At the 2-year follow up, the patient reported severe pain with twisting or stairs and patellar tendinitis. Topical antiinflammatory cream was applied, and the complication resolved the same day. | This is a combination product. (b)(4). List of associated devices: - refobacin bc r 1x40 us, reference 110034355, batch 835aae2507 - articular surface fixed bearing constrained posterior stabilized (cps) left 16 mm height use with tibia sizes e-f / ps femur sizes 6-9, reference (b)(4), batch 64536583 - femur cemented plus left size 7+, reference (b)(4), batch 64429274 - tibia fixed cemented left size f stem, reference (b)(4), batch 64207423 - stem extension 3mm offset splined uncemented 13 mm diameter +135 mm length, reference (b)(4), batch 64559288 - tibial central cone size small, reference (b)(4), batch 64189640 - tibial augment cemented half block left medial size ef 15 mm thickness, reference (b)(4), batch 64363791 - tibial augment cemented half block left lateral size ef 10 mm thickness, reference (b)(4), batch 64289805 - stem extension 3mm offset splined uncemented 14 mm diameter +135 mm length, reference (b)(4), batch 64514946 - femoral distal augment cemented size 7, 7+ 5 mm thickness, reference (b)(4), batch 64295436 - femoral distal augment cemented size 7, 7+ 10 mm thickness, reference (b)(4), batch 64336285. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13665529&pc=MBB&device_sequence_no=1
3006946279-2022-00016 | 02/04/2022 | 03/03/2022 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INSERT 28 S60 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that, during a revision surgery of hip replacement, the surgeon encountered difficulties with the installation of the joint between the head and the liner. Indeed, the two items slide apart very easily. The surgeon elected to immediately revise the liner and head. These items were explanted. The surgery took around 25 minutes longer than expected due to the need to revise the implanted head and liner. | (b)(4). List of associated devices: 12/14 ceramic fem head +3. 5x28, reference (b)(4), batch 2895208; femoral head +10. 5x28mm dia, reference (b)(4), batch 62474057; cpt 12/14 size 4 cocr xext, reference (b)(4), batch 63706891; cpt centralizer pmma pkgd, reference (b)(4), batch 64590498; act artic e1 hip brg 28x54mm, reference (b)(4), batch 376190. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13659307&pc=JDI&device_sequence_no=1
3006946279-2022-00019 | 12/03/2021 | 03/03/2022 | 90 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that an initial left total knee arthroplasty performed in 2007 with unknown product was subsequently revised (b)(6) 2019 due to persistent pain. At the 2-year follow up, the patient reported severe pain with twisting or stairs and patellar tendinitis. Topical antiinflammatory cream was applied, and the complication resolved the same day. | This is a combination product. (b)(4). List of associated devices: refobacin bc r 1x40 us, reference 110034355, batch 828ead1611. Articular surface fixed bearing constrained posterior stabilized (cps) left 16 mm height use with tibia sizes e-f / ps femur sizes 6-9, reference 42512600716, batch 64536583. Femur cemented plus left size 7+, reference 42504606211, batch 64429274. Tibia fixed cemented left size f stem, reference 42542007501, batch 64207423. Stem extension 3mm offset splined uncemented 13 mm diameter +135 mm length, reference 42560313513, batch 64559288. Tibial central cone size small, reference 42545000511, batch 64189640. Tibial augment cemented half block left medial size ef 15 mm thickness, reference 42555805315, batch 64363791. Tibial augment cemented half block left lateral size ef 10 mm thickness, reference 42555805110, batch 64289805. Stem extension 3mm offset splined uncemented 14 mm diameter +135 mm length, reference 42560313514, batch 64514946. Femoral distal augment cemented size 7, 7+ 5 mm thickness, reference 42556606205, batch 64295436. Femoral distal augment cemented size 7, 7+ 10 mm thickness, reference 42556606210, batch 64336285. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13659292&pc=MBB&device_sequence_no=1
3006946279-2022-00017 | 03/02/2021 | 03/03/2022 | 366 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS FEMORAL STEM | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent revision surgery due to gts stem loosening. | (b)(4). List of associated devices: head (delta ceramic), reference and batch unknown. Liner (e-poly), reference and batch unknown. Exceed cup, reference and batch unknown. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13659099&pc=KWA&device_sequence_no=1
3006946279-2022-00015 | 02/02/2022 | 03/03/2022 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported, that the cement hardened (quickly 4 - 6 minutes) after putting the optipac in the optigun. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13658021&pc=LOD&device_sequence_no=1
3006946279-2022-00014 | 12/14/2021 | 02/08/2022 | 56 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported, that the hardening of the bone cement was faster than usual. It hardened in about four minutes. Another product was used to performed the surgery. It took a few minute to take a new product. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source: foreign: event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13480382&pc=LOD&device_sequence_no=1
3006946279-2022-00013 | 01/02/2017 | 02/01/2022 | 1856 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC SOFTPAC REFOBACIN REVISION 40 | BONE CEMENT, ANTIBIOTIC | - | It was reported the patient received a right total knee arthroplasty on (b)(6) 2015. The patient had persistent pain after surgery and all components except the patella were revised on (b)(6) 2017 due to loosening. | This is a combination product (b)(4). List of associated devices: femoral component option size d right, reference 00599401492, batch 62480325. Stemmed tibial component precoat a/p wedged size 3, reference 00598800300, batch 62781311 articular surface with locking screw size cd 20 mm height, reference 00596403020, batch 60763893. Poly petella standard cemented size 29 mm diameter 8. 0 mm thickness, reference 00597206529, batch 62549287. Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. Therefore, the reported event could not be confirmed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to pain and loosening: 1 complaint (1 product), this one included, has been recorded on optipac softpac refobacin revision 40, reference (b)(4), from january 01, 2019 to january 27, 2022. 1 complaint (1 product), this one included, has been recorded on optipac softpac refobacin revision 40, reference (b)(4), batch 235ca16970. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13412796&pc=MBB&device_sequence_no=1
3006946279-2022-00012 | 11/12/2021 | 01/28/2022 | 77 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a left total knee arthroplasty on (b)(6) 2019. Subsequently, the patient experienced pain, swelling, couldn't straighten leg, couldn't walk without falling down, couldn't use leg, weight gain, became diabetic and had loosening of the implant. The patient underwent a revision procedure on (b)(6) 2021 where large amounts of metal debris were found. | (b)(4). List of associated devices: biomet bone cement r 1x40 us, reference (b)(4), batch 743cah1407; offset tib tray 2. 5mm adaptor, reference (b)(4), batch 752210; series a pat std 31 3 peg, reference (b)(4), batch 674220; vngd ti fem ssk 70mm rt, reference cp113126, batch 486890; bmt splined knee stm 18x120, reference (b)(4), batch 082360; biomet offset tibial tray 71mm, reference (b)(4), batch 154000; bmt splined knee stm 80x16mm, reference (b)(4), batch 901020; vngd ssk psc tib brg 10x71/75, reference (b)(4), batch 740600. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13382142&pc=LOD&device_sequence_no=1
3006946279-2022-00011 | 11/12/2021 | 01/28/2022 | 77 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the patient underwent a left total knee arthroplasty (b)(6) 2019. Subsequently, the patient experienced pain, swelling, couldn't straighten leg, couldn't walk without falling down, couldn't use leg, weight gain, became diabetic and had loosening of the implant. The patient underwent a revision procedure on (b)(6) 2021 where large amounts of metal debris were found. | (b)(4). List of associated devices: biomet bone cement r 1x40 us, reference 110035368, batch 826aac2203. Offset tib tray 2. 5mm adaptor, reference (b)(4), batch 752210. Series a pat std 31 3 peg, reference (b)(4), batch 674220. Vngd ti fem ssk 70mm rt, reference (b)(4), batch 486890. Bmt splined knee stm 18x120, reference (b)(4), batch 082360. Biomet offset tibial tray 71mm, reference (b)(4), batch 154000. Bmt splined knee stm 80x16mm, reference (b)(4), batch 901020. Vngd ssk psc tib brg 10x71/75, reference (b)(4), batch 740600. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13382081&pc=LOD&device_sequence_no=1
3006946279-2022-00010 | 04/09/2019 | 01/28/2022 | 1025 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported through a clinical study that a patient underwent a primary total knee arthroplasty on (b)(6) 2019. Subsequently, the patient was revised on (b)(6) 2019 date due to deep infection. | This is a combination product. (b)(4). Concomitant medical products - list of associated devices: psn asf cr 10mm ply l 7-12 gh, reference 42-5110-006-10, batch 63561707a17. Psn all poly pat ply 32mm, reference 42-5400-000-32, batch 64173391i18. Psn tib stm 5 deg sz g l, reference 42-5320-079-01, batch 63931138b18. Psn fem cr cmt ccr std sz 9 l, reference 42-5026-066-01, batch 63487999i16. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13382075&pc=MBB&device_sequence_no=1
3006946279-2022-00009 | 12/01/2020 | 01/28/2022 | 423 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC-S 60 REFOBACIN BONE CEMENT R | BONE CEMENT, ANTIBIOTIC | - | It was reported through a clinical study that a patient underwent a primary total knee arthroplasty on (b)(6) 2019. Subsequently, the patient was revised on (b)(6) 2020 date due to deep infection. Manufacturer Narrative This is a combination product: (b)(4). List of associated devices: psn asf cr 10mm ply l 3-11 ef, reference 42-5110-005-10, batch 63535021; psn tib stm 5 deg sz f l, reference 42-5320-075-01, batch 63560762; psn fem cr cmt ccr std sz 11 l, reference 42-5026-070-01, batch 63567263; psn all poly pat ply 38mm, reference 42-5400-000-38, batch 64100059. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | This is a combination product: (b)(4). List of associated devices: psn asf cr 10mm ply l 3-11 ef, reference 42-5110-005-10, batch 63535021; psn tib stm 5 deg sz f l, reference 42-5320-075-01, batch 63560762; psn fem cr cmt ccr std sz 11 l, reference 42-5026-070-01, batch 63567263; psn all poly pat ply 38mm, reference 42-5400-000-38, batch 64100059. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Manufacturer Narrative This is a combination product: (b)(4). List of associated devices: psn asf cr 10mm ply l 3-11 ef, reference 42-5110-005-10, batch 63535021; psn tib stm 5 deg sz f l, reference 42-5320-075-01, batch 63560762; psn fem cr cmt ccr std sz 11 l, reference 42-5026-070-01, batch 63567263; psn all poly pat ply 38mm, reference 42-5400-000-38, batch 64100059. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13382068&pc=MBB&device_sequence_no=1
3006946279-2022-00007 | 09/07/2018 | 01/24/2022 | 1235 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient underwent a knee procedure on (b)(6) 2014. Subsequently, the patient was revised due to pain and infection after multiple diagnostics on (b)(6) 2018. | This is a combination product (b)(4). Concomitant medical products - list of associated devices: optipac 80 refobacin revision, reference 4732501165, batch 341cal0301. Kne-nexgen-tibial trays, reference and batch not communicated. Kne-nexgen-bearings, reference and batch not communicated. Kne-nexgen-femorals, reference and batch not communicated. Kne-nexgen-stems, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to infection and pain: 1 complaint (1 product), this one included, has been recorded on optipac 40 refobacin revision, reference (b)(4), from january 01, 2018 to january 21, 2022. 1 complaint (1 product), this one included, has been recorded on optipac 40 refobacin revision, reference (b)(4), batch 211aai1803. According to available data, root cause of the event was unable to be determined. It is not excluded that the contamination come from hospital environment or patient health condition. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13333933&pc=MBB&device_sequence_no=1
3006946279-2022-00006 | 09/07/2018 | 01/24/2022 | 1235 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that the patient underwent a knee procedure on (b)(6) 2014. Subsequently, the patient was revised due to pain and infection after multiple diagnostics on (b)(6) 2018. | This is a combination product. (b)(4). List of associated devices: optipac 40 refobacin revision, reference 4730501163, batch 211aai1803; kne-nexgen-tibial trays-unk, reference and batch not communicated; kne-nexgen-bearings-unk, reference and batch not communicated; kne-nexgen-femorals-unk, reference and batch not communicated; kne-nexgen-stems-unk, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to infection and pain: 1 complaint (1 product), this one included, has been recorded on optipac 80 refobacin revision, reference 4732501165, from january 01, 2018 to january 21, 2022. 1 complaint (1 product), this one included, has been recorded on optipac 80 refobacin revision, reference 4732501165, batch 341cal0301. According to available data, root cause of the event was unable to be determined. It is not excluded that the contamination come from hospital environment or patient health condition. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13333603&pc=MBB&device_sequence_no=1
3006946279-2022-00008 | 02/28/2013 | 01/24/2022 | 3252 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN REVISION | BONE CEMENT, ANTIBIOTIC | - | It was reported that during a retrospective study of the nexgen ap wedge tibial prosthesis patient underwent an initial implant of (b)(6) 2012. Subsequently, the patient underwent revision procedure on (b)(6) 2015 due to pain and loosening. The pi listed the event as uncertainly related to the instrument, uncertainly related to the device and probably related to the procedure. | This is a combination product (b)(4). List of associated devices: lcck fem implant sz c-rt, reference 00-5994-013-92, batch 61522741. Ng a/p wgd prct tib plt sz 4, reference 00-5988-004-00, batch 61916850. Lcck art surf cd 3-4/yel 17, reference 00-5994-030-17, batch 61903523. Greport source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13333101&pc=MBB&device_sequence_no=1
3006946279-2022-00005 | 05/05/2021 | 01/21/2022 | 261 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R-3 1X40 | BONE CEMENT | - | It was reported that the patient had an initial right total knee arthroplasty performed (b)(6) 2020. On (b)(6) 2021, it has been noticed a prepatellar bursitis inflammation treated with anti-inflammatories. On (b)(6) 2021, the patient suffered from a severe sepsis and acute organ dysfunction. Subsequently, the patient underwent simultaneous and bilateral knee irrigation and debridement with poly exchange on (b)(6) 2021 due to acute hematogenous infection. | (b)(4). List of associated devices: cr-flex gsf pct sz f-r minus, reference 00-5750-016-06, batch 64459337; peg prc tib plt size 5, reference 00-5970-045-01, batch 64740565; ng ac art surf 5-6/ grn 17mm, reference 00-5976-040-17, batch 64597174; all poly pat comp 35dia, reference 00-5972-065-35, batch 64774932. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13313384&pc=LOD&device_sequence_no=1
3006946279-2022-00004 | 12/14/2021 | 01/03/2022 | 20 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET PLUS BONE CEMENT 1X40-3 | BONE CEMENT | - | It was reported that during the surgery of vanguard knee prosthesis, when opening the cements they realized that between the last bag and the previous one there was powder loose. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13141439&pc=LOD&device_sequence_no=1
3006946279-2022-00003 | 12/06/2021 | 01/03/2022 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported, that during placement of cement on tibia, with the tibia subluxed and soft tissues protected, bovie cauterization (electrosurgical cauterization) was utilized in the medial parapatellar area. At this point, an assistant noticed at the center part of the cement, that a low flame was noted, was tamped, put out immediately glowing on surface of cement. The cement was removed. The wound was copiously irrigated once again. Witnessed no tissues compromised or flame adjacent to tissues. Cementing technique was performed. The implants were put in atraumatically. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13141121&pc=MBB&device_sequence_no=1
3006946279-2022-00002 | 12/06/2021 | 01/03/2022 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported, that during placement of cement on tibia, with the tibia subluxed and soft tissues protected, bovie cauterization (electrosurgical cauterization) was utilized in the medial parapatellar area. At this point, an assistant noticed at the center part of the cement, that a low flame was noted, was tamped, put out immediately glowing on surface of cement. The cement was removed. The wound was copiously irrigated once again. Witnessed no tissues compromised or flame adjacent to tissues. Cementing technique was performed. The implants were put in atraumatically. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13141116&pc=MBB&device_sequence_no=1
3006946279-2021-00216 | 11/16/2021 | 12/30/2021 | 44 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that the seal on the pouch was partially opened and the cement polymer was leaking. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13118241&pc=LOD&device_sequence_no=1
3006946279-2021-00215 | 10/26/2021 | 12/30/2021 | 65 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that the seal on the pouch was partially opened and the cement polymer was leaking. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13117447&pc=LOD&device_sequence_no=1
3006946279-2021-00214 | 12/15/2021 | 12/29/2021 | 14 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COBALT CHROME FEMORAL HEAD 5°42 Ø28 / 0 / 12-14/ COL MOYEN | METALLIC FEMORAL HEAD PROSTHESIS | - | It was reported, that there was an intraprosthetic dislocation of the right total hip prosthesis. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: avantage inlay s54 / 28, reference p0561054, batch 000151683. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13106225&pc=KWA&device_sequence_no=1
3006946279-2021-00213 | 12/15/2021 | 12/29/2021 | 14 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE INLAY S54 / 28 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported, that there was an intraprosthetic dislocation of the right total hip prosthesis. | (b)(4). Report source, foreign - event occurred in (b)(6). Concomitant medical products - list of associated devices: cobalt chrome femoral head 5°42 ø28 / 0 / 12-14/ col moyen, reference (b)(4), batch j6787299. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13106207&pc=KWY&device_sequence_no=1
3006946279-2021-00212 | 11/16/2021 | 12/29/2021 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AURA II HIP HA COATED RIGHT SIZE 6 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that, following a patient fall on (b)(6) 2021, the stem broke and all of the devices initially implanted had to be revised on (b)(6) 2021. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13105939&pc=JDI&device_sequence_no=1
3006946279-2021-00210 | 06/22/2021 | 12/17/2021 | 178 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC REFOBACIN BONE CEMENT R | BONE CEMENT, MEDICATED | - | It was reported that the patient underwent an initial partial knee surgery on (b)(6) 2020, with optipac refobacin bone cement r was used during surgery. After that, the patient suffered from pain, swelling and progressive loss of range of motion of the knee. Despite medical treatment and physiotherapy, the pain persisted. On (b)(6) 2021, patient had an increased inflammation value (crp = 6 mg/ l). On (b)(6) 2021, patient underwent a revision surgery due to a suspicion of septic loosening. As reported, the operation report stated that there was practically no mechanical connection in the implant - cement interface - and that the implant could be easily retrieved. It was also reported that the femur component was no longer frictional connected. No issue was noticed during the revision surgery. No pathological examination of the tissue in the region of the prosthesis was performed. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). The product analysis can't be performed as the product was not returned. Therefore, the reported event could not be confirmed. The review of the device manufacturing quality can't be performed as the product reference and batch number was not communicated. A complaint extract was done regarding revision due to pain, implant loosening and inflammation: (b)(4) complaints ((b)(4) products), this one included, have been recorded on optipac refobacin bone cement r, from (b)(6) 2018 to (b)(6) 2021. The complaint history, regarding the reference number, can't be performed as it was not communicated. The complaint history, regarding the batch number, can't be performed as it was not communicated. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13026048&pc=MBB&device_sequence_no=1
3006946279-2021-00211 | 12/07/2021 | 12/15/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the bag of bone cement is broken even before the bone cement box was open. When cement powder bag was taken out off the box, the cement powder splatter. Then we had to open another bone cement box instead. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=13009486&pc=LOD&device_sequence_no=1
3006946279-2021-00207 | 09/08/2016 | 12/13/2021 | 1922 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC-S 60 REFOB BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the patient underwent an initial knee arthroplasty on (b)(6) 2015 and was revised due to pain and loosening of the femoral component on (b)(6) 2018. Subsequently, patient experienced pain and lock knee (inability to either bend or straighten the knee) 9 months after total knee replacement. There was no explanation for this on the x-ray. Patient had arthroscopy (surgery) of the right knee and the prosthesis and surrounding area had been cleaned. No infection was detected. Patient was revised two years after the first revision due to pain and loosening of the tibial component. The current complaint investigates the first revision. | This is a combination product. (b)(4). Report source, study - event occurred in (b)(6). List of associated products : item number 42521400710, item name articular surface fixed bearing posterior stabilized (ps) right 10 mm height, lot # 63051212. Item number 42500606402, item name hfemur cemented posterior stabilized (ps) standard right size 8, lot # 63174617. Item number 42532007102, item name rtibia cemented 5 degree stemmed right size e, lot # 63119287. Item number 42540000035, item name all poly patella cemented 35 mm diameter, lot # 63029234. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. No dhr available for now. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12988833&pc=MBB&device_sequence_no=1
3006946279-2021-00209 | 11/11/2021 | 12/13/2021 | 32 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT | - | It was reported that the bone cement used in bipolar hip arthroplasty hardened in 4 minutes, it could not be used for this surgery. The surgery was prolonged of no more than 3 minutes to take another cement. No adverse patient outcome was reported related to this event. | (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Batch unknown. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12987478&pc=LOD&device_sequence_no=1
3006946279-2021-00208 | 11/12/2021 | 12/13/2021 | 31 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the bone cement hardened in 1 minute during a total knee arthroplasty, it could not be used for this surgery. Another cement was used. No patient adverse consequence was reported related to this event. | (b)(4). Report source, foreign, health professional - event occurred in (b)(6). Concomitant medical products - list of associated products : item number 11-5049-011-01 item name cement mixing bowl lot # 65194894. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12984114&pc=MBB&device_sequence_no=1
3006946279-2021-00206 | 11/09/2021 | 12/13/2021 | 34 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that: there was no problem to mix with optipac, but the bone cement came out like a water. It could not be mixed correctly, so not used at the surgeon's decision. No adverse patient consequence was reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12982752&pc=MBB&device_sequence_no=1
3006946279-2021-00205 | 11/15/2021 | 12/09/2021 | 24 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S50 / 28 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that the patient underwent an initial hip arthroplasty on (b)(6) 2019. After that, he was on (b)(6) 2020 due to dislocation between the head and the insert (event handle in (b)(4). Subsequently, the patient was revised again on (b)(6) 2021 due dislocation between the head and the insert (even handle in the current complaint and (b)(4). | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: femoral head non-skirted 12/14 taper, reference 00801802814, batch 64054134. Femoral stem 12/14 neck taper ext. Offset size 1 130 mm stem length, reference 00811400110, batch 64493868. Allen medullary cement plugs 1-28 mm diameter flange/14 mm diameter core store in cool dry place, reference 00801102028, batch 64545194. Avantage cemented cup s50, reference p0463050, batch 0001354635. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12966389&pc=KWA&device_sequence_no=1
3006946279-2021-00202 | 10/25/2021 | 12/08/2021 | 44 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that the bone cement hardened faster than usual in about 8 minutes in this surgery. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12958984&pc=LOD&device_sequence_no=1
3006946279-2021-00203 | 11/26/2021 | 12/08/2021 | 12 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the bone cement hardened faster than usual in about 5 minutes in this surgery. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12958982&pc=MBB&device_sequence_no=1
3006946279-2021-00204 | 11/30/2021 | 12/08/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that shortly after bringing the two components together, the cement immediately became hard. It was no longer possible to finish the mixing process. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12958978&pc=MBB&device_sequence_no=1
3006946279-2021-00201 | 11/23/2021 | 12/07/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the cement envelope was opened and sterility was lost. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12951250&pc=LOD&device_sequence_no=1
3006946279-2021-00200 | 11/23/2021 | 12/07/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the cement envelope was opened and sterility was lost. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12950639&pc=LOD&device_sequence_no=1
3006946279-2021-00199 | 11/23/2021 | 12/07/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the cement envelope was found opened and sterility was lost. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12950629&pc=LOD&device_sequence_no=1
3006946279-2021-00198 | 10/26/2021 | 11/24/2021 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the cement was flaky and not pasted as usual. No known adverse event was reported. | This is a combination product. (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12867213&pc=MBB&device_sequence_no=1
3006946279-2021-00197 | 10/26/2021 | 11/24/2021 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the cement was flaky and not pasted as usual. No known adverse event was reported. | This is a combination product. (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12867198&pc=MBB&device_sequence_no=1
3006946279-2021-00196 | 10/25/2021 | 11/17/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It was reported that the sterile cement package was opened inside of unsterile outer package. No adverse events have been reported as a result of the malfunction. | (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12828687&pc=LOD&device_sequence_no=1
3006946279-2021-00194 | - | 11/09/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a surgery has been performed with rosa knee and persona cps in a patient with rhumatoid arthritis performed bilaterally at one time. The surgery took place as expected, without complications. Doctor reported that this patient would be a challenge due to reduced movement amplitude. She was able to extend her leg only up to 30 degrees and flex her leg up to 70 degrees. According to the doctor, the postoperative period still presented pain of unknown origin which he imagined was due to the surgery. However, as the condition did not evolve, 3 months after surgery he asked the patient to undergo a xray, where he verified the loosening of the prosthesis on both sides, apparently only on the femoral side. As a result, he underwent a series of tests, including mri scans and will soon underwent tests to assess whether there was an infection. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). List of associated products : item number 42500606202 item name pfemur cemented posterior stabilized (ps) standard right size 7lot # 64302103; item number 19090127af1 item name hoscillating saw blade zimmerâ®/synthesâ® lot # 372069; item number 12076100ur1 item name reciprocating saw blade zimmer biomet recip lot # 348875; item number 00515048200 item name hip kit lot # 64796937; item number 2557000114 item name pstem extension tapered cemented 14 mm diameter +30 mm length lot # 64359223; item number 42540000032 item name pall poly patella cemented 32 mm diameter lot # 64367722; item number 42532007101 item name pnatural tibia cemented 5 degree stemmed left size e lot # 64367521; item number 42512600812 item name particular surface fixed bearing constrained posterior stabilized (cps) left 12 mm height use with tibia sizes e-f / ps femur sizes 10-11lot # 63904814; item number 42557000114 item name pstem extension tapered cemented 14 mm diameter +30 mm length lot # 64832711; item number 42540000035 item name all poly patella cemented 35 mm diameter lot # 64374887; item number 42532006702 item name ptibia cemented 5 degree stemmed right size d lot # 64355617; item number 42522600510 item name particular surface fixed bearing constrained posterior stabilized (cps) right 10 mm height use with tibia sizes c-d / ps femur sizes 6-9 lot # 64461648. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12777885&pc=MBB&device_sequence_no=1
3006946279-2021-00193 | - | 11/09/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a surgery has been performed with rosa knee and persona cps in a patient with rheumatoid arthritis performed bilaterally at one time. The surgery took place as expected, without complications. Doctor reported that this patient would be a challenge due to reduced movement amplitude. She was able to extend her leg only up to 30 degrees and flex her leg up to 70 degrees. According to the doctor, the postoperative period still presented pain of unknown origin which he imagined was due to the surgery. However, as the condition did not evolve, 3 months after surgery he asked the patient to undergo a xray, where he verified the loosening of the prosthesis on both sides, apparently only on the femoral side. As a result, he underwent a series of tests, including mri scans and will soon underwent tests to assess whether there was an infection. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). Concomitant medical products: list of associated products : item number 42500606202 item name pfemur cemented posterior stabilized (ps) standard right size 7lot # 64302103. Item number 19090127af1 item name hoscillating saw blade zimmerâ®/synthesâ® lot # 372069. Item number 12076100ur1 item name reciprocating saw blade zimmer biomet recip lot # 348875. Item number 00515048200 item name hip kit lot # 64796937. Item number 2557000114 item name pstem extension tapered cemented 14 mm diameter +30 mm length lot # 64359223. Item number 42540000032 item name pall poly patella cemented 32 mm diameter lot # 64367722. Item number 42532007101 item name pnatural tibia cemented 5 degree stemmed left size e lot # 64367521. Item number 42512600812 item name particular surface fixed bearing constrained. Posterior stabilized (cps) left 12 mm height use with tibia sizes e-f / ps femur sizes 10-11lot # 63904814. Item number 42557000114 item name pstem extension tapered cemented 14 mm diameter +30 mm. Length lot # 64832711. Item number 42540000035 item name all poly patella cemented 35 mm diameter lot # 64374887. Item number 42532006702 item name ptibia cemented 5 degree stemmed right size d lot # 64355617. Item number 42522600510 item name particular surface fixed bearing constrained posterior stabilized (cps) right 10 mm height use with tibia sizes c-d / ps femur sizes 6-9 lot # 64461648. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12777841&pc=MBB&device_sequence_no=1
3006946279-2021-00195 | 10/15/2021 | 11/09/2021 | 25 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN PLUS BONE CEMENT 40 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that: a patient underwent a knee arthroplasty on (b)(6) 11, 2017. Patient was revised due to infection on (b)(6) 2021. Patient involvement- revision surgery. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). List of associated products : item 0059620421 _ articular surface with insert and locking screw use with lps/lps-flex 51 or 52 suffix femoralssize gh 17 mm height _ batch : 62920054; item 4020830401_ refobacin plus bone cement _batch ag10dl0103. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12777639&pc=MBB&device_sequence_no=1
3006946279-2021-00192 | 10/15/2021 | 11/09/2021 | 25 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIVAC M | MANUAL ORTHOPAEDIC CEMENT MIXER, REUSABLE | - | It was reported that a patient underwent a knee arthroplasty. Patient was revised due to infection 4 years after the initial surgery. | (b)(4). Report source, foreign - event occurred in (b)(6). Associated product:- item 0059620421 _ articular surface with insert and locking screw use with lps/lps-flex 51or 52 suffix femoralssize gh 17 mm height _ batch : 62920054; item 4020830401_ refobacin plus bone cement _batch ag10dl0103. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12777627&pc=JDZ&device_sequence_no=1
3006946279-2021-00191 | 09/10/2021 | 11/08/2021 | 59 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that: a patient was operated on (b)(6) 2021 for a right intermediate prosthesis. He had an intraoperative fat embolism, remained in intensive continuous care until (b)(6) 2021. The hospital decided not to use this product anymore. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12766506&pc=MBB&device_sequence_no=1
3006946279-2021-00190 | 10/11/2021 | 11/08/2021 | 28 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that: a patient was operated on (b)(6) 2021 for a left intermediate prosthesis. He had several intraoperative cardiac arrests, and went in the intensive care unit. Poor prognosis , the patient is in intensive care. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12766469&pc=MBB&device_sequence_no=1
3006946279-2021-00189 | 10/08/2021 | 11/04/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It was reported, that before the operation, it was found that white powder was leaked out. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12750230&pc=LOD&device_sequence_no=1
3006946279-2021-00187 | 10/11/2021 | 11/03/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported, that bone cement powder was not fully sealed at the bottom of the packaging and it spilled out during the opening of the packaging. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12743375&pc=MBB&device_sequence_no=1
3006946279-2021-00188 | 10/15/2021 | 11/03/2021 | 19 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It was reported that at opening, the powder bag of the cement was already opened and the powder spilled on the table. To continue the surgery it was opened another cement with the same reference. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12743362&pc=LOD&device_sequence_no=1
3006946279-2021-00186 | 09/30/2021 | 10/26/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET PLUS BONE CEMENT 40 -3 | BONE CEMENT | - | It was reported that during a knee prosthesis surgery, when opening the cement in the auxiliar table, they realized that there was loose powder of polymethacrylate between the 2nd and 3rd bag of the product. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12696244&pc=LOD&device_sequence_no=1
3006946279-2021-00185 | 10/01/2021 | 10/25/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported by the nurse an optipac 40g was mixed as usual, with all surgical theatre temperatures same as in the past on average 16 degrees. 2min after the mixing, the staff tried to press the cement paste out of the nozzle and it was very sticky and only became workable at about 5min. The doctor complained about the cement consistency. The cement paste took about 15-16min to set. No patient health consequences or impact has been reported. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). List of associated products : item number 405940 item name comp rvs hmrltrl 44std-36std 4std-36std lot # unknown. Item number 405907 item name 25mm compr rvs taper adapt trl lot # unknown. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12688404&pc=MBB&device_sequence_no=1
3006946279-2021-00184 | 09/10/2021 | 10/25/2021 | 45 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported, that a patient underwent revision due to loosening of the femoral cemented stem and wear with fracture of the polyethylene. The cement initially implanted was an optipac 80 refobacin bone cement r-3. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). Concomitant medical products: list of associated products : item number 150372 item name oss cemented im stem 18x150 50 lot # 216090. Item number 150411 item name oss tibial poly bearing 14mm lot # 822750. Item number unknown item name unknown femoral component lot # unknown. Item number unknown item name unknown hinge assembly lot # unknown. Item number unknown item name unknown tibial tray and stem lot # unknown. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12688377&pc=MBB&device_sequence_no=1
3006946279-2021-00183 | 09/29/2021 | 10/25/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | ORTHOPAEDIC CEMENT, MEDICATED | - | I was reported that: after stirring, cement hardened in about 6 minutes. It usually cures in about 13 minutes. The surgeon reported concerns about the fixation of femoral implant. To date, no adverse event has been reported. Manufacturer Narrative (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Manufacturer Narrative (b)(4). 2 complaints on cement paste too short setting time were recorded on the batch since ever, and 5 complaints on cement paste too short setting time were recorded on the reference from oct 01, 2018 to nov 3, 2021. Reserve sample from the same lot was tested under standardized conditions. The product did not show any unusual behavior during mixing, handling or setting. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12687714&pc=LOD&device_sequence_no=1
3006946279-2021-00182 | 09/30/2021 | 10/21/2021 | 21 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIVAC M | MANUAL ORTHOPAEDIC CEMENT MIXER, REUSABLE | - | It was reported that during the application of cement into a patient's femoral canal, at the point of snapping off the cement applicator nozzle, the whole length of the applicator cartridge nozzle snapped and broke. An attempt was made to reconnect the nozzle to cement applicator cartridge. However, on inspection, the two connections were broken. As some cement has already been applied to the patient's femoral canal, the consultant decided to proceed and improvise the rest of the application of cement. No known adverse event was reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12672467&pc=JDZ&device_sequence_no=1
3006946279-2021-00181 | 09/14/2021 | 10/21/2021 | 37 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the bone cement was hardening in 3 minutes used with total knee arthroplasty. It could not be used for this surgery. No known adverse event was reported. | (b)(4). Report source, foreign and health professional: event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12670210&pc=LOD&device_sequence_no=1
3006946279-2021-00180 | 08/05/2021 | 10/21/2021 | 77 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported, that the bone cement was hardening in 8 minutes used with total knee arthroplasty, it could not be used for tibial component but only for femoral component. No known adverse event was reported. | (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12670195&pc=LOD&device_sequence_no=1
3006946279-2021-00179 | 09/17/2021 | 10/18/2021 | 31 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 JP | BONE CEMENT, ANTIBIOTIC | - | It was reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. This surgery was finished with backup product. No adverse health consequence has been reported as the result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12646078&pc=MBB&device_sequence_no=1
3006946279-2021-00178 | 09/15/2021 | 10/18/2021 | 33 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 JP | BONE CEMENT, ANTIBIOTIC | - | It was reported, that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. This surgery was finished with backup product. No adverse health consequence has been reported as the result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information isfound which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12646062&pc=MBB&device_sequence_no=1
3006946279-2021-00177 | 09/24/2021 | 10/18/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the inner package leaked. The issue was found before the surgery started by zb staff. The product will not be returned due to the cost of shipment. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12646045&pc=LOD&device_sequence_no=1
3006946279-2021-00176 | 09/22/2021 | 10/18/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the inner package leaked. The issue was found before the surgery started by (b)(6) staff. The product will not be returned due to the cost of shipment. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information isfound which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12645939&pc=LOD&device_sequence_no=1
3006946279-2021-00175 | 09/21/2021 | 10/18/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the inner package leaked. The issue was found before the surgery started by zb staff. The product will not be returned due to the cost of shipment. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information isfound which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12645928&pc=LOD&device_sequence_no=1
3006946279-2021-00173 | - | 10/15/2021 | - | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that the cement started to harden about 2 minutes after the start of stirring, so it could not be used. No known adverse event was reported. | (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12638006&pc=MBB&device_sequence_no=1
3006946279-2021-00174 | 09/15/2021 | 10/15/2021 | 30 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that after stirring the cement, the cylinder was attached to the cement gun, but it was reported that it hardened in about 4 minutes. No known adverse event was reported. | (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12637976&pc=MBB&device_sequence_no=1
3006946279-2021-00172 | 09/15/2021 | 10/12/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the inner package leaked. The issue was found before the surgery started by (b)(6) staff. The product will not be returned due to the cost of shipment. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12613258&pc=LOD&device_sequence_no=1
3006946279-2021-00171 | 09/17/2021 | 10/11/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It was reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. The surgery was finished with backup product. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604969&pc=LOD&device_sequence_no=1
3006946279-2021-00170 | 09/15/2021 | 10/11/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It was reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. This surgery was finished with backup product. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604924&pc=LOD&device_sequence_no=1
3006946279-2021-00169 | 08/25/2021 | 10/11/2021 | 47 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 JP | BONE CEMENT, ANTIBIOTICS | - | It was reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. The surgery was finished with backup product. No adverse event has been reported as the result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604891&pc=MBB&device_sequence_no=1
3006946279-2021-00168 | 08/17/2021 | 10/11/2021 | 55 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. This surgery was finished with backup product. No adverse event has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604841&pc=LOD&device_sequence_no=1
3006946279-2021-00167 | 09/27/2021 | 10/11/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that when the nurse opened a box of refobacin bone cement, the pouch containing powder was not sealed. No adverse health consequence has been reported as the result of the malfunction. It was reported as well, that the product will not be returned as the cement was destroy on site. | This is a combination product (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604807&pc=MBB&device_sequence_no=1
3006946279-2021-00166 | 09/23/2021 | 10/11/2021 | 18 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 58MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that while dr (b)(6) was impacting the avantage cup into the acetabular after it seated and he removed the avantage impactor instrument noticed it was broken. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12604787&pc=HWA&device_sequence_no=1
3006946279-2021-00165 | 09/13/2021 | 10/07/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It has been reported that reported that, before the operation, it was found that the inner sterile packaging was opened. No adverse event has been reported as a result of the malfunction. | (b)(4). (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12588097&pc=LOD&device_sequence_no=1
3006946279-2021-00163 | 06/03/2020 | 10/07/2021 | 491 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CEMENT | CEMENT | - | It was reported that a patient had right elbow pain for years and advanced arthrosis. In (b)(6) 2020, an initial intervention to put a cemented implant for the elbow was performed, not involving a zimmer biomet cement. The cement solidified in a very fast way and mechanical complications appeared after the surgery. A revision surgery was therefore performed on (b)(6) 2021. The revision surgery resulted in complication post operative with paralysis of the right upper extremities and loss of usage of right hand. The current complaint investigated the revision surgery on (b)(6) 2020. This report is communicated as a conservative approach, as to date, it is not confirmed that a zimmer biomet cement was used during this revision surgery. | (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record can't be performed since the reference and batch number of the product was not communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. No reference or batch number communicate. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12588045&pc=LOD&device_sequence_no=1
3006946279-2021-00164 | 09/08/2021 | 10/05/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET PLUS BONE CEMENT 40-3 | BONE CEMENT | - | It was reported that the inner cement pouch was opened. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12574541&pc=LOD&device_sequence_no=1
3006946279-2021-00162 | 09/23/2021 | 10/01/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT | - | It was reported that the cement powder leaked out through the unsealed bottom packaging upon removing it from the unsterile packaging. No adverse health consequence has been reported as the result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12561798&pc=LOD&device_sequence_no=1
3006946279-2021-00161 | 09/15/2021 | 10/01/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the inner package leaked. This issue was found before the surgery started. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12560413&pc=LOD&device_sequence_no=1
3006946279-2021-00159 | 08/30/2021 | 09/28/2021 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the doctor performed an initial surgery at an unknown date. Then patient underwent infection, and he was implanted with a spacer to treat infection. The doctor performed a knee surgery to remove the knee implant used as temporary spacer and implant revision knee. No additional patient consequences were reported. | (b)(4). Report source, health professional - event occurred in united states. Concomitant medical devices: vanguard cr ilok femoral-lt 67. 5 - reference : 183030 - lot : j6666239; vangard ant stabilized brg12x67 - reference : 189042 - lot : 747800. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12538495&pc=LOD&device_sequence_no=1
3006946279-2021-00160 | 08/30/2021 | 09/28/2021 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the doctor performed an initial surgery at an unknown date. Then patient underwent infection, and he was implanted with a spacer to treat infection. The doctor performed a knee surgery to remove the knee implant used as temporary spacer and implant revision knee. No additional patient consequences were reported. | (b)(4). Report source, health professional - event occurred in united states. Concomitant medical devices: vanguard cr ilok femoral-lt 67. 5 - reference : 183030 - lot : j6666239; vanguard antero stabilized brg12x67 - reference : 189042 - lot : 747800. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12538491&pc=LOD&device_sequence_no=1
3006946279-2021-00158 | 09/02/2021 | 09/17/2021 | 15 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported, that: "after opening the outer package of bone cement during the operation, it was found that the sterile package was broken and white powdery objects leaked out. There was only one batch of bone cement involved, four packets. One of them:bc corner+dc corner+da corner. One of them:bc corner +d side. One of them:cd corner. The other of them: ad corner+ c side. The bone cement will not be returned because of customs regulations". As reported there was a patient involvement. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12487231&pc=MBB&device_sequence_no=1
3006946279-2021-00157 | 09/01/2021 | 09/17/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported, that after opening the outer package of bone cement during the operation, it was found that the sterile package was broken and white powder leaked, which was bc corner. In addition it was reported that the product will not be return because of customs regulations. As reported, there is a patient involvement. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12487224&pc=MBB&device_sequence_no=1
3006946279-2021-00156 | 08/27/2021 | 09/17/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that: "on (b)(6) 2021, bone cement was removed during knee surgery in (b)(6) hospital of (b)(6) medicine, and the packaging of bone cement powder was found to be leaking powder. As reported, the event occurred during the operation, and consequently the patient is involved. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12487206&pc=MBB&device_sequence_no=1
3006946279-2021-00155 | - | 09/15/2021 | - | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN CEMENT | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported on 18-aug-2021, a journal article was retrieved from international orthopaedics (2021) that reported a database search from the (b)(6) that looked at rates of aseptic loosening between two types of total knee designs. The purpose of the study was to ascertain the rates of aseptic tibial loosening for the two implant types within five years of implantation and to investigate the causes for any difference observed. Overall, 2172 tkas were performed in the study period with 759 augmentable tibial components inserted between 05-jan-2009 and jan-2011 and 1413 non-augmentable tibial components between jan-2011 and 31-mar-2015. All the patients underwent primary tka using the nexgen lps-fixed bearing prosthesis. All of the implants were posterior stabilised and were cemented without stems or augments. All of the augmentable tibial components had a pmma pre-coating. The non-augmentable tibial components had no pmma pre-coating. Palacos r&g (heraeus medical) cement was used until january 2012 until it was phased out over a two month period to refobacin (zimmer biomet). The study population had a mean age of 68 years at time of surgery (range 21-92). A five year follow-up is noted without specified intervals provided. The study reported 22 patients underwent revision due to aseptic loosening. 'patients were classified as having aseptic loosening if they had symptoms including pain, instability or swelling; had radiographic evidence of loosening; and did not meet the msis criteria for infection. ' of the 22 patients, 17 had refobacin cement. Reference: michael brown, rohan ramasubbu, mark jenkinson, james doonan, mark blyth, bryn jones (2021). Significant differences in rates of aseptic loosening between two variations of a popular total knee arthroplasty design. International orthopaedics. | This is a combination product. (b)(4). Report source, foreign, literature - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. No reference and batch was communicated. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12474187&pc=MBB&device_sequence_no=1
3006946279-2021-00154 | 08/20/2021 | 09/13/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 58MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that during impaction of definitive avantage reload acetabular cup (total hip replacement), part of the avantage 58mm impator tip broke off. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12459302&pc=KWA&device_sequence_no=1
3006946279-2021-00154 | 08/20/2021 | 09/13/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 58MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that during impaction of definitive avantage reload acetabular cup (total hip replacement), part of the avantage 58mm impator tip broke off. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12459302&pc=QBO&device_sequence_no=1
3006946279-2021-00153 | 08/27/2021 | 09/09/2021 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It has been reported that, before the surgery, it was found that cement powder was leaking out from the package. As reported, this described event occurred in hong kong, and the product will not be returned due to customs regulations. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12446714&pc=LOD&device_sequence_no=1
3006946279-2021-00151 | 08/16/2021 | 09/09/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | T was reported that the customer was using an optipac 40g indoors. He mixed it up as usual. Temperatures in the theater were the same as in the past with an average of sixteen degrees. Two minutes after mixing, he tried to press out the nozzle and it was very sticky and only became usable after about five minutes. The doctor complained about the consistency. It took about 15-16min to set. No known adverse event was reported. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Batch not available. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12446574&pc=MBB&device_sequence_no=1
3006946279-2021-00150 | 08/16/2021 | 09/09/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that an optipac 40g was mixed in theatre as normal. All theater temperatures were the same as in the past an average of sixteen degrees. Two minutes after mixing the customer tried to get out of the nozzle and it was very sticky and became usable after about five minutes. The total setup time was twenty-one minutes. In the same theater, on previous occasions, it took 15-16min. No known adverse event was reported. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12445965&pc=MBB&device_sequence_no=1
3006946279-2021-00152 | 08/19/2021 | 09/09/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, during the surgery, the powder pouch was poorly sealed. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12444980&pc=LOD&device_sequence_no=1
3006946279-2021-00149 | 08/22/2021 | 09/02/2021 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was found opened before the surgery. As reported the product will no be returned due to customs regulations. In addition it was reported that no patient was involved. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12413592&pc=MBB&device_sequence_no=1
3006946279-2021-00146 | 08/22/2021 | 09/02/2021 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was found opened before the surgery. As reported the product will no be returned due to customs regulations. In addition it was reported that no patient was involved. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12412258&pc=MBB&device_sequence_no=1
3006946279-2021-00147 | 08/22/2021 | 09/02/2021 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was found opened before the surgery. As reported the product will not be returned due to customs regulations. In addition it was reported that no patient was involved. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the products were not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12412254&pc=MBB&device_sequence_no=1
3006946279-2021-00148 | 08/19/2021 | 09/02/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that a pre-operative examination of bone cement revealed powder leakage from the inner cement pouch. As reported the product will no be returned due to customs regulations. In addition it was reported that no patient was involved. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12411930&pc=MBB&device_sequence_no=1
3006946279-2021-00144 | 08/01/2021 | 09/01/2021 | 31 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. As reported, the product will not be returned due to custom regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign: event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12405862&pc=MBB&device_sequence_no=1
3006946279-2021-00142 | 08/11/2021 | 09/01/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, it was found that he polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. As reported, the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12405604&pc=LOD&device_sequence_no=1
3006946279-2021-00143 | 08/09/2021 | 08/31/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CMNT R 1X40 JP | BONE CEMENT, ANTIBIOTIC | - | It was reported that this cement could not be used due to a defective packaging seal on the polymer pouch. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12391470&pc=MBB&device_sequence_no=1
3006946279-2021-00145 | 08/18/2021 | 08/31/2021 | 13 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It was reported that the inner sterile packaging was found opened before the surgery. As reported the product will not be returned due to customs regulations. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign: event occurred in (b)(6). The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12391508&pc=LOD&device_sequence_no=1
3006946279-2021-00139 | 04/27/2021 | 08/27/2021 | 122 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | ENDOBON® XENOGRAFT GRANULES 0.5ML | BONE MATRIX IMPLANT, ANIMAL-DERIVED, NON-ABSORBABLE | - | It has been reported that the patient had a loss of integration and infection with pain and inflammation after a endobon xenograft implantation. Doctor performed reconstruction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. The review of the raw material certificate shows no non conformity or deviation. The label of the product is stuck on the zper and confirms the expiration date. Complaint history : a complaint extract was done regarding medical : infection and non integration: 12 complaints (12 products), this one included, have been recorded on endobon® xenograft granules 0. 5ml, reference rox05, from 01 jan 2018 to 18 aug 2021. 1 complaint (1 product), this one included, has been recorded on endobon® xenograft granules 0. 5ml, reference rox05, batch y0098304. A complaint extract was done regarding medical : expired product implanted. 1 complaint (1 product), this one included, has been recorded on endobon® xenograft granules 0. 5ml, reference rox05, from 01 jan 2018 to 27 aug 2021. 1 complaint (1 product), this one included, has been recorded on endobon® xenograft granules 0. 5ml, reference rox05, batch y0098304. Investigation results concluded that the reported event was due to user error (the product has been used in spite of the expiration date passed). A summary of the investigation was sent to the complainant conveying proper use of the device. There is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12376477&pc=LYC&device_sequence_no=1
3006946279-2021-00140 | - | 08/27/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 JP | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that an anaphylactic phenomenon occurred in patient who used refobacin bone cement. The patient recovered the next day. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12376392&pc=MBB&device_sequence_no=1
3006946279-2021-00141 | 08/20/2021 | 08/26/2021 | 6 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +2 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that the stem implanted 8 years ago has broken. As reported by the patient, at the time of the event he was running as he did regularly. In addition it was reported that, as a result of the reported event, a medical intervention/revision was performed on (b)(6) 2021. Also, as reported, at the moment it is not possible to have the return ,as communicated the representative of the area that will give feedback as soon as it is possible to have it. It was also reported that there were no injuries reported as a result of the event (no impairment after the intervention) and that no delays in intervention were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: maxera cup 40mm x 50mm; ref#00-1515-050-40; lot#62207415. Biolox-delta ceramic femoral head; ref#00-8775-040-02; lot#2691418. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12371153&pc=OQG&device_sequence_no=1
3006946279-2021-00136 | 07/28/2021 | 08/23/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. As reported, the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12352480&pc=LOD&device_sequence_no=1
3006946279-2021-00137 | 07/27/2021 | 08/23/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because a part of the inner polymer pouch was unsealed. As reported, the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12352442&pc=LOD&device_sequence_no=1
3006946279-2021-00138 | 08/05/2021 | 08/23/2021 | 18 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because a part of the inner polymer pouch was unsealed. As reported, the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12352436&pc=LOD&device_sequence_no=1
3006946279-2021-00134 | 07/30/2021 | 08/19/2021 | 20 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that on 30. 07 morning a patient underwent a surgery to implant an avenir stem. The (b)(6) patient with a history of stroke suffered a cement shock when introducing refobacin cement into the femoral shaft. The patient went into cardiac arrest on the table, forcing the anesthetists to perform cardiac massage and then several shockings to try to resuscitate her. In the end, it ended well as the patient was then referred (after closing the incision) to the intensive care unit. We received the information that the patient died in the afternoon. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12342168&pc=LOD&device_sequence_no=1
3006946279-2021-00135 | 07/28/2021 | 08/19/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported cement hardens too quickly in the distal femur. A second portion of cement could therefore no longer be administered. Used articles were disposed of intraoperatively. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12342083&pc=MBB&device_sequence_no=1
3006946279-2021-00132 | 07/23/2021 | 08/13/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that customer is experienced in using the product : hardening time is too short. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 1 complaint (1 product), has been recorded on optipac 60 refobacin bone cement r-3, reference (b)(4), batch az16bb2411 since ever. According to available data, the root cause could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12317936&pc=MBB&device_sequence_no=1
3006946279-2021-00131 | 07/26/2021 | 08/13/2021 | 18 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | BONE CEMENT | - | It has been reported that the surgeon complaint of cement not feeling as sticky as usual. Patella did not adhere to bone. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source - event occurred in united states. The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 1 complaint (1 product), has been recorded on biomet bone cement r 1*40 us, reference 110035368, batch f2701z34ca since ever. No further information provided (x-rays, surgical report, photographs, lab test). According to available data, the root cause could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12317795&pc=LOD&device_sequence_no=1
3006946279-2021-00133 | 09/01/2019 | 08/06/2021 | 705 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CEMENT PACK -1 | BONE CEMENT | - | It was reported by the surgeon a minimal swelling of the ipsilateral knee after partial knee arthroplasty with a tinbn-coated implant. Surgeon categorized the event as an adverse event, with uncertain relation to the device and probably related to the procedure. Patient was treated with icing the knee and conservative treatment. Swelling is now not persistent anymore. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12286743&pc=LOD&device_sequence_no=1
3006946279-2021-00127 | 07/12/2021 | 08/03/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the bag containing the cement powder was not correctly sealed and that as a result, the powder came out from the bag. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12264867&pc=LOD&device_sequence_no=1
3006946279-2021-00126 | 05/24/2021 | 08/03/2021 | 71 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12264903&pc=MBB&device_sequence_no=1
3006946279-2021-00128 | 07/12/2021 | 08/03/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It was reported that the inner sterile package was found opened before the surgery. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12264863&pc=LOD&device_sequence_no=1
3006946279-2021-00129 | 07/16/2021 | 08/03/2021 | 18 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that the polymer was leaking from the inner pouch because a part of the inner polymer pouch was unsealed. As reported the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12264832&pc=LOD&device_sequence_no=1
3006946279-2021-00130 | 01/28/2021 | 08/02/2021 | 186 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the cement would not mix and the powder was still present. It was also reported later, that the paste had clumps so the consistency failed during the mixing process (stryker advanced mixing system). Another cement was prepared and surgery was finished. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Event occurred in united states. The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. No available photos have been provided by the customer which confirm the issue, but the received picture with the labels confirms the references of the two involved products. 1 complaint (1 product), this one included, has been recorded on refobacin bone cement r 1x40 us, reference 110034355, batch 936bae2505 since ever. No information on the stryker mixing system and mixing process followed by the customer has been provided. According to available data, the root cause could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12259081&pc=MBB&device_sequence_no=1
3006946279-2021-00125 | 01/28/2021 | 08/02/2021 | 186 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the cement would not mix and the powder was still present. It was also reported later, that the paste had clumps so the consistency failed during the mixing process (stryker advanced mixing system). Another cement was prepared and surgery was finished. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Event occurred in united states. The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. No available photos have been provided by the customer which confirm the issue, but the received picture with the labels confirms the references of the two involved products. 1 complaint (1 product), this one included, has been recorded on refobacin bone cement r 1x40 us, reference 110034355, batch 840cag1903 since ever. According to available data, the root cause could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12258918&pc=MBB&device_sequence_no=1
3006946279-2021-00124 | 07/20/2021 | 07/29/2021 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, before the surgery, it was found that the powder of the inner package leaked out. As reported, the product will not be returned due to the cost of shipment. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12249687&pc=LOD&device_sequence_no=1
3006946279-2021-00119 | 03/02/2020 | 07/29/2021 | 514 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. As reported, the position of powder leakage is ad angle. As reported the product cannot be returned to (b)(6) as it does not comply customs regulations. No adverse event was reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files ¿(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12249676&pc=MBB&device_sequence_no=1
3006946279-2021-00123 | 07/20/2021 | 07/29/2021 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, before the surgery, it was found that the powder of the inner package leaked out. As reported, the product will not be returned due to the cost of shipment. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12249604&pc=LOD&device_sequence_no=1
3006946279-2021-00122 | 07/26/2021 | 07/29/2021 | 3 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was found opened during the surgery. As reported the doctor used the spare bone cement to complete the operation and the surgery was delayed for 5 minutes, due to the issue. In addition, it was reported that the product shall not be returned as per customs regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. It was noticed that the surgeon tried to use an expired product. The customer has therefore, been informed of this point. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12245608&pc=MBB&device_sequence_no=1
3006946279-2021-00121 | 07/26/2021 | 07/29/2021 | 3 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was opened before the surgery. As reported the doctor used the spare bone cement to complete the operation. It was reported as well that the product shall not be returned due to customs regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12245442&pc=MBB&device_sequence_no=1
3006946279-2021-00120 | 04/22/2021 | 07/29/2021 | 98 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. The product cannot be returned to france, as it does not comply customs regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. One picture was received but the reported event could not be confirmed on the basis of that picture. The product analysis was not able to be performed as the product was not returned. A complaint extract was done regarding packaging : inner_cement_pouch_open_sealing: 54 complaints (involving 155 products), this one included, have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, from 01-jan-2018 to 28-jul-2021. 1 complaint (involving 2 products), this one included, has been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, batch z18aag0103, since ever. So this is a single event for this product lot. The reported event was not able to be confirmed as the product was not returned and no picture showing the reported event was available. Therefore, the root cause of the event was unable to be determined. Regarding the inner cement pouch open sealing issues, these are well-known events, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). And a corrective action has been initiated to address this issue. A summary of the investigation has been sent to the customer. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12245337&pc=MBB&device_sequence_no=1
3006946279-2021-00118 | 06/15/2021 | 07/28/2021 | 43 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because a part of the inner polymer pouch was unsealed. In addition, it was reported that the surgery was finished with backup product. No adverse event was reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). The product analysis was not performed as the product was not returned to the manufacturer. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. A sealing force test has been performed on a product received from a similar complaint ((b)(4) - same item reference) and showed that the pouch sealing force a was compliant with zimmer biomet valence specification. A complaint extract was done regarding packaging : inner_cement_pouch_open_sealing: 12 complaints (involving 12 products), this one included, have been recorded on biomet bone cement r 1x40 jp, reference 110035372, from (b)(6) 2018 to (b)(6) 2021. 8 complaints (involving 8 products), this one included, have been recorded on biomet bone cement r 1x40 jp, reference 110035372, batch z34eaf1509, since ever. Investigation results concluded that the reported event was due to manufacturing (sealing process). A scar ((b)(4)) has been initiated in order to deeper investigate the cause of the event with the cement pouch supplier osartis and to consider to implement actions to avoid the recurrence of this type of issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12238203&pc=LOD&device_sequence_no=1
3006946279-2021-00117 | 07/06/2021 | 07/26/2021 | 20 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. In addition, it was reported that the surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12223615&pc=LOD&device_sequence_no=1
3006946279-2021-00116 | 07/02/2021 | 07/23/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It has been reported that, upon opening of cement packet tech noticed cement powder pouring out of side of cement packet. As reported, the product was disqualified and a replacement product was opened, all prior to patient being in the surgical suite. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, health professional - event occurred in united states. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12216252&pc=LOD&device_sequence_no=1
3006946279-2021-00115 | 06/21/2021 | 07/23/2021 | 32 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, during the surgery, the inner cement pouch was found not sealed: as reported while visual examination, the inner pouch containing the powder was not sealed at certain points. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12216274&pc=LOD&device_sequence_no=1
3006946279-2021-00114 | 04/04/2021 | 07/23/2021 | 110 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. Product cannot be returned to france as it does not comply customs regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12216236&pc=MBB&device_sequence_no=1
3006946279-2021-00110 | - | 07/15/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | IMPLANT CABLE 2 MM STAINLESS STEEL | ORTHOPAEDIC BONE WIRE | - | It has been reported a suspicion of metallosis on a patient according to the test results mentioned below. Moreover, it was noticed that the scar has a grey-blue-brownish tint. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12171962&pc=LRN&device_sequence_no=1
3006946279-2021-00111 | - | 07/15/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION STEM HAP SIZE 4 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was requested by the poison control center of (b)(6) the composition of products biomet exception hap n°4 - cupule avantage biomet e1 50 22 col metal court. Indeed, these products were recorded with undesirable events. No more information provided. | (b)(4). List of associated devices: avantage cup, reference and batch not communicated; report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12171863&pc=KXA&device_sequence_no=1
3006946279-2021-00112 | - | 07/15/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CUP | ACETABULAR SHELL | - | It was requested by the (b)(6) the composition of products biomet exception hap n°4 - cupule avantage biomet e1 50 22 col metal court. Indeed, these products were recorded with indesirable events. No more information provided. | Cmp-(b)(4). Concomitant medical products: exception stem hap size 4, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12171823&pc=KWA&device_sequence_no=1
3006946279-2021-00113 | 07/06/2021 | 07/13/2021 | 7 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the cement hardened too fast. No adverse events have been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The review of the device manufacturing quality can't be performed as the product batch number was not communicated. 14 complaints (18 products), this one included, have been recorded on optipac 40 refobacin bone cement r-3, reference (b)(4) from 01 january 2018 to date (13 july 2021). No added information are available in order to investigate the involved product. With the available information, the exact root cause of the event could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Lot number not communicated. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12158634&pc=MBB&device_sequence_no=1
3006946279-2021-00109 | 02/10/2021 | 07/09/2021 | 149 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that patient underwent right total knee revision (b)(6) 2020 due to tibial loosening. As reported, subsequently, the patient developed pain and underwent a second right tka revision on (b)(6) 2021 due to tibial loosening: the poly and tibial component were revised. In addition it was reported that the tibial component was found grossly loose with varus alignment (related complaint is (b)(4)). As reported the cement used is refobacin bc r 1x40 us ref 110034355 lot 907cai1801. | This is a combination product (b)(4). Report source heath professional - event occurred in united states. This is a combination product. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12143134&pc=MBB&device_sequence_no=1
3006946279-2021-00108 | - | 07/06/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a patient experienced pain, surgeon states mild intra-articular effusion. Surgeon classified the relation to device as uncertain, not related to both procedure and instrumentation. It was not a revision surgery, but a complication reported as part of a post market clinical follow-up study. The product is still in place in the patient. | This is a combination product (b)(4). Report source, foreign and study. Event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. No reference and batch number. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12120177&pc=MBB&device_sequence_no=1
3006946279-2021-00107 | 09/01/2020 | 07/05/2021 | 307 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN REVISION 40 -3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the patient developed allergic reaction after having a surgery related to refobacin cement. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12115206&pc=MBB&device_sequence_no=1
3006946279-2021-00104 | 06/08/2021 | 07/02/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, before the surgery, it was noticed that the packaging of the cement powder was badly sealed. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12105898&pc=LOD
3006946279-2021-00105 | 06/15/2021 | 07/01/2021 | 16 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that an optipac 80g was used in tha after removal of femoral nail. It was injected intramedullally 4 minutes and 30 seconds after the agitation was completed. About 12 minutes after inserting the stem, the cement hardened. Immediately afterwards, the patient's blood pressure dropped to 20. As a result of recovery treatment by doctors, vitals recovered in about 20 minutes. The patient was transferred to the intensive care unit in the intubated state. | (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12098818&pc=MBB
3006946279-2021-00106 | 06/15/2021 | 07/01/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the cement setting too quickly. The scrub nurse had difficulty in mixing it for the required time as it was already setting. There was no patient involvement. | This is a combination product (b)(4). Report source, foreign and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12098927&pc=MBB
3006946279-2021-00103 | 05/18/2021 | 06/30/2021 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S50 / 28 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that a patient presented on emergency board case for dislocated hip. Surgeon noticed that the head was disengaged from the liner again. Moreover, he noticed some trauma to metal head and liner probably caused after disengagement and contact with the cup. Stem and cup were solid. New liner and 28mm head were implanted. This dislocation happened after a first one, handle in (b)(4). | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: stryker 28mm metal +4mm head, reference and batch unknown. No further information provided (x-rays, surgical report, photographs, lab test). The product analysis can't be performed as the product was not returned. The review of the device manufacturing quality can't be performed as the product batch number was not communicated. A complaint extract was done regarding revision due to dislocation: 4 complaints (4 products), this one included, have been recorded on avantage e1 inlay s50 / 28, reference p0561e50, from january 01, 2018 to june 28, 2021. The complaint history, regarding the batch number, can't be performed as it was not communicated. The review of the instruction for use of the avantage brand sated that only legacy zimmer or legacy biomet femoral heads without skirt, 22. 2mm and 28mm diameter, manufactured from stainless steel, alumina ceramic (biolox® delta) or cobalt chromium with a 12/14 taper, a type 1 taper or an 8/10 taper can be used. However, in the current event, the head used with the inlay was manufactured by stryker. Thus, the compatibility between these two components cannot be guaranteed. According to available data, root cause of the event was unable to be determined. However, it can be noticed that one of the potential root cause could be the incompatibility between the inlay and the head used. A summary of the investigation has been sent to the customer. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12089947&pc=KWA
3006946279-2021-00102 | 02/06/2021 | 06/29/2021 | 143 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE E1 INLAY S50 / 28 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that a patient presented on emergency board case for dislocated hip. Surgeon discovered that the liner and metal head had disassociated. Both head and inlay were replaced with same size head/liner. Some posterior structure was repaired to augment the capsul. The range of motion was good and no impingement was detected. After that, patient underwent a new revision surgery due to dislocation, handled in (b)(4). | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: stryker 28mm metal head, reference and batch not communicated. No further information provided (x-rays, surgical report, photographs, lab test). The product analysis can't be performed as the product was not returned. The review of the device manufacturing quality can't be performed as the product batch number was not communicated. A complaint extract was done regarding revision due to dislocation: 4 complaints (4 products), this one included, have been recorded on avantage e1 inlay s50 / 28, reference p0561e50, from january 01, 2018 to june 28, 2021. The complaint history, regarding the batch number, can't be performed as it was not communicated. The review of the instruction for use of the avantage brand sated that only legacy zimmer or legacy biomet femoral heads without skirt, 22. 2mm and 28mm diameter, manufactured from stainless steel, alumina ceramic (biolox® delta) or cobalt chromium with a 12/14 taper, a type 1 taper or an 8/10 taper can be used. However, in the current event, the head used with the inlay was manufactured by stryker. Thus, the compatibility between these two components cannot be guaranteed. According to available data, root cause of the event was unable to be determined. However, it can be noticed that one of the potential root cause could be the incompatibility between the inlay and the head used. A summary of the investigation has been sent to the customer. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12083955&pc=JDI
3006946279-2021-00101 | 06/02/2021 | 06/25/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the sterile package of bone cement product was inspected and powder leakage was found when it entered the hospital system. In addition, it was reported that no surgery was involved. As reported, the product will not be returned as it does not comply customs regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12065126&pc=MBB
3006946279-2021-00099 | 05/12/2021 | 06/22/2021 | 41 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the patient underwent a revision due to infection. There are a total of 5 complaints associated with this event, as 5 products are involved (b)(4) related to bone cement (current report) batch f2701z12ba. (b)(4) related to bone cement batch 927caa1602. (b)(4) related to freedom constrained liner and freedom constrained head. (b)(4) related to bone screw. (b)(4) related to freedom cup. No additional patient consequences were reported. | (b)(4). Report source: health professional - event occurred in united states. List of associated devices: freedom all poly cup 54mm, reference (b)(4), batch 075510, bone screw self-tapping 6. 5 mm dia. 35 mm length, reference (b)(4), batch j6829869, ball nose guide wire 80cm, reference (b)(4), batch 444030, frdm cnstr hd 36mm t12/14 +9mm, reference (b)(4), batch 675390, bone screw self-tapping 6. 5 mm dia. 50 mm length, reference (b)(4), batch 64143387, 3. 2mmx40mm rnglc+ acet drl bit, reference (b)(4), batch 068790, wagn cone prosthesis 125 21, reference (b)(4), batch 3023874. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12038648&pc=LOD
3006946279-2021-00100 | 05/12/2021 | 06/22/2021 | 41 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the patient underwent a revision due to infection. There are a total of 5 complaints associated with this event, as 5 products are involved. (b)(4) related to bone cement (current report) batch f2701z12ba. (b)(4) related to bone cement batch 927caa1602. (b)(4) related to freedom constrained liner and freedom constrained head. (b)(4) related to bone screw. (b)(4) related to freedom cup. No additional patient consequences were reported. | (b)(4). Report source, health professional - event occurred in united states. List of associated devices: freedom all poly cup 54mm, reference (b)(4), batch 075510. Bone screw self-tapping 6. 5 mm dia. 35 mm length, reference (b)(4), batch j6829869. Ball nose guide wire 80cm, reference (b)(4), batch 444030. Frdm cnstr hd 36mm t12/14 +9mm, reference (b)(4), batch 675390. Bone screw self-tapping 6. 5 mm dia. 50 mm length, reference (b)(4), batch 64143387. 3. 2mmx40mm rnglc+ acet drl bit, reference (b)(4), batch 068790. Wagn cone prosthesis 125 21, reference (b)(4), batch 3023874. The investigation is in progress. Once the investigation has been completed, a follow-up mdr. Will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12038643&pc=LOD
3006946279-2021-00098 | 06/09/2021 | 06/18/2021 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 2X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner packaging of powder of the product was unsealed. The powder fell out the package and the quantity of the powder was not enought to start the mixing process. A backup product was used to complete the surgery. There was no patient involvement. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the lot number was not communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12023923&pc=MBB
3006946279-2021-00094 | 04/27/2021 | 06/16/2021 | 50 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, during the surgery, the package of the cement was damaged, hence cement was not used. In addition, it was reported that the internal package was badly sealed (see (b)(4). No adverse event has been reported as a result of the malfunction. | Cmp-(b)(4). Report source, foreign - event occurred in la reunion. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12006604&pc=LOD
3006946279-2021-00095 | 05/24/2021 | 06/16/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, during the surgery, the user realized that a part of cement was in package, and that due to doubt on product's sterility it was not used. As reported, the surgery was completed with alternative cement. In addition it was reported that the inner sterile cement pouch was damaged and that the cement powder leaked out of this package. No adverse event has been reported as a result of the malfunction. | This is a combination product cmp- (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12006602&pc=MBB
3006946279-2021-00096 | 06/02/2021 | 06/16/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, it was noticed that the cement powder was leaking from the inner pouch, because part of the inner cement pouch sealing was damaged. As reported, this surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | Cmp- (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12006600&pc=LOD
3006946279-2021-00097 | 06/02/2021 | 06/16/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, it was noticed that the cement powder was leaking from the inner pouch, because part of the inner cement pouch sealing was damaged. As reported, this surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | Cmp-(b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12006599&pc=LOD
3006946279-2021-00093 | 05/24/2021 | 06/15/2021 | 22 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that immediately after injecting optipac cement into the medullary cavity with tha surgery, the patient's blood pressure dropped sharply. Eventually, the patient was reported dead. The surgeon commented that the patient was in poor condition and there was no problem with the product. | (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The datasheet on the composition of the product were reviewed and no anomalies were found. No other complaint on patient death following bone cement implantation syndrome was recorded on the batch since ever, and no other complaint on patient death following bone cement implantation syndrome was recorded on the reference since ever. The ifu of the product has been reviewed and it was found that the ifu for optipac contains risk -reducing warning and different precautions regarding severe complications: "after preparation of the prosthesis bed or directly after the implantation of the cement and prosthesis, pressure rise in the medullary canal may cause a temporary fall in blood pressure. In addition to hypotension, pulmonary embolism and cardiac arrest, with their potentially fatal consequences, have been encountered in rare cases". Reserve sample from the same lot was tested under standardized conditions. The product did not show any unusual behavior during mixing, handling or setting. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11997295&pc=MBB
3006946279-2021-00092 | 05/08/2021 | 06/11/2021 | 34 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 2X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that when the cement was put into femur, the saturation in the blood pressure dropped in both cases. (separate incidents). First patient passed away, 3 minutes into recovery. Second patient was reintegrated and sent to critical care. The current report handles the second patient. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The ifu of the product has been reviewed and it contains risk -reducing warning and different precautions regarding severe complications): "after preparation of the prosthesis bed or directly after the implantation of the cement and prosthesis, pressure rise in the medullary canal may cause a temporary fall in blood pressure. In addition to hypotension, pulmonary embolism and cardiac arrest, with their potentially fatal consequences, have been encountered in rare cases". The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The datasheet on the composition of the product were reviewed and no anomalies were found. No other complaint bone cement implantation syndrome was recorded on the batch since ever, and one other complaint bone cement implantation syndrome was recorded on the reference since ever. Reserve sample from the same lot was tested under standardized conditions. The product did not show any unusual behavior during mixing, handling or setting. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11983706&pc=LOD
3006946279-2021-00091 | 05/06/2021 | 06/11/2021 | 36 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 2X40-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that when the cement was put into femur, the saturation in the blood pressure dropped in both cases. First patient passed away, 3 minutes into recovery. Second patient was reintegrated and sent to critical care. (b)(4) was created to investigate on the second patient. The current report corresponds to the first patient (death). | This is a combination product. (b)(4). Foreign report source: (b)(6). The ifu of the product has been reviewed and it was found that the ifu for cement pack contains risk -reducing warning and different precautions regarding severe complications: "after preparation of the prosthesis bed or directly after the implantation of the cement and prosthesis, pressure rise in the medullary canal may cause a temporary fall in blood pressure. In addition to hypotension, pulmonary embolism and cardiac arrest, with their potentially fatal consequences, have been encountered in rare cases". The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The datasheet on the composition of the product were reviewed and no anomalies were found. No other complaint on patient death following bone cement implantation syndrome was recorded on the batch since ever, and no other complaint on patient death following bone cement implantation syndrome was recorded on the reference since ever. Reserve sample from the same lot was tested under standardized conditions. The product did not show any unusual behavior during mixing, handling or setting. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11983150&pc=MBB
3006946279-2021-00090 | - | 06/09/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VARIZED FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - LEFT S5 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that patient underwent a revision surgery on (b)(6) 2021 due to aseptic loosening 3 years post-operatively. It was also reported by the sales representative who assisted to the revision surgery that the metaphysal loosening was probably due to a malpositioning of the stem. | (b)(4). List of associated devices: cobalt chrome femoral head 5°42 ø28 / 0 / 12-14/ col moyen, reference p0206m28, batch j6214037; avantage reload acetabular cup / cementless / size 56, reference p0460p56, batch 0001165540; avantage inlay size 56 ø 28, reference p0561056, batch 0001180858. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11964848&pc=KWY
3006946279-2021-00089 | 04/26/2021 | 06/08/2021 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BC R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a patient underwent a revision total knee due to loose tibia, pain and instability. | Cmp-(b)(4). Report source, health professional - event occurred in united states of america. List of associated products : item number : 96-0100 - item name : oval dome patella 3-peg - lot : d19072003. Item number : 96-0013 - item name : femoral cruciate retaining cemented size 3 right - lot : d19090535. Item number : 1581-30-000 - item name : tibial tray fixed bearing mod cocr - lot : 9285147. Item number : 1581-13-108 - item name : tibial insert fixed bearing curved - lot : 9299157. Item number : ar-1927bcf-3 - item name : suture anchor biocomposite corkscrew ft tripleplay vented - lot : 10319263. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11954498&pc=LOD
3006946279-2021-00086 | 05/12/2021 | 06/03/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It has been reported that the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. In addition it was reported that this surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11929897&pc=LOD
3006946279-2021-00087 | 05/12/2021 | 06/03/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because part of the inner polymer powder pouch was unsealed. In addition it was reported that this surgery was finished with backup product. No adverse event has been reported as a result of the malfunction. | Cmp- (b)(4) report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11929865&pc=LOD
3006946279-2021-00088 | 05/18/2021 | 06/03/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, the powder bag was not well sealed and that the cement powder was therefore partly spilled into the first package. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11929829&pc=LOD
3006946279-2021-00084 | - | 05/28/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CEMENT | CEMENT | - | It was reported through a clinical study of a removal of 2 free cement fragments in the patient in 2020 due to a bicycle fall in 2019 that resulted in right knee contusions and then joint locking. The cement reference was not communicated to date. | (b)(4). Report source, study and health professional - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11901395&pc=LOD
3006946279-2021-00085 | 05/04/2021 | 05/26/2021 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | STANDARD FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - RIGHT S8 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a product has been implanted on (b)(6) 2021 whereas the expiration date was (b)(6) 2021. It has also been reported that the stock control has not been performed by zimmer biomet sales representative since 6 months which leads to several stock-out. The instrumentalist was aware of the surgery but forgot to verify expiration dates one week before the surgery. The day of the surgery, the surgeon planned to implant an exception femoral stem size 7, however, during the trials, he realized that size 7 would not fit and that it would take size 8. However, size 8 was expired and surgeon had no other choice to implant the expired stem. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusion or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11886552&pc=KXA
3006946279-2021-00083 | 02/18/2021 | 05/12/2021 | 83 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COCR-HEAD 28MM 12/14 S NECK-3.5 | METALLIC FEMORAL HEAD PROSTHESIS | - | It was reported that the patient underwent an initial hip arthroplasty on (b)(6) 2006 performed by dr (b)(6) at the (b)(6) hospital with an metal-on-metal hip prosthesis. A few years later, patient began to complaint of severe groin pain and suffered two miscarriages. In addition, in 2014, patient suffered from an ovarian cancer for which she was operated. Than, the groin pain worsened with such intensity as it prevents her from walking. In (b)(6) 2020, blood and urine tests were performed and revealed an abnormal presence of cobal in her body fluids. Finally, patient was revised on (b)(6) 2021 as the ct scan revealed cracks in the biomet prosthesis. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11810914&pc=JDI
3006946279-2021-00082 | 04/14/2021 | 05/11/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported patient underwent a left knee arthroplasty on (b)(6) 2020. Subsequently, the patient was revised on (b)(6) 2021, due to a looseporous stem and femoral component. Dr. (b)(6) revised with an oos femur and a cemented stem. | (b)(4). List of associated devices: natural tibia cemented 5 degree stemmed left size e, reference 42532007101, batch 64794770; reference :150464 oss 3cm diaphysel segment lot : 119830; reference :150493 oss reinforced yoke batch : 293590; reference :150477 oss poly femoral bushings 2pk batch : 869290; reference :130615 intramedullary plug xl batch : 372180; reference :150480 oss axle batch : 171210; reference :150355 oss 7cm segmental femoral lt batch : 188880; reference :150388 oss porous im stem 17. 5 x 90 batch : 427060; reference :150476 oss poly tibial bushing batch : 783180; reference :150411 oss tibial poly bearing 14mm batch : 274210; reference :130613 intramedullary plug lge batch : 003670; reference :150413 oss tibial poly bearing 18mm batch : 948210; reference :150355 oss 7cm segmental femoral lt batch : 188940; reference :150493 oss reinforced yoke batch : 898000; reference :150477 oss poly femoral bushings 2pk batch : 533050; reference :150480 oss axle batch : 826040; reference :150482 oss 4cm diaphyseal segment batch : 289650; reference :150366 oss cemented im stem 12x150 batch : 994210; reference :150478 oss poly lock pin batch : 207110. Report source, foreign - health professional - event occurred in united states. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11805523&pc=LOD
3006946279-2021-00081 | - | 05/11/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS FEMORAL STEM | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported by dr. (b)(6) that a revision surgery will be needed due to progressive varization of the gts stem after covid-19. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11804342&pc=LPH
3006946279-2021-00080 | 04/20/2021 | 05/11/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 56MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that during the impaction of the cup, a piece of the impactor broke off upon impact. Nothing fell into the patient and they were able to finish the surgery without any delay. | (b)(4) report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11804324&pc=LXH
3006946279-2021-00079 | 03/25/2021 | 05/07/2021 | 43 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | ENDOBON® XENOGRAFT GRANULES LARGE PARTICLE 8.0ML | BONE MATRIX IMPLANT, ANIMAL-DERIVED, NON-ABSORBABLE | - | It has been reported that the patient experienced a bone loss. Sinus lift was performed at tooth site 16 and there was dehiscense. Bone regeneration was necessary as there was bone loss and loss of the membrane. No additional patient consequences were reported. | (b)(4). Concomitant medical products: copios membrane ref (b)(4) lot nza18490169. Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the lot number was not correct. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11789444&pc=LYC
3006946279-2021-00078 | 04/14/2021 | 05/07/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE 3P ACETABULAR CUP / CEMENTLESS/ SIZE 62 | ACETABULAR SHELL | - | It was reported that 1 expired product has been consumed in a surgery when the product was being invoicing. Indeed, the product expired on 22 feb. 2021 and was used in a surgery on (b)(6) 2021. The sales agent reported that some days before the surgery, he removed from the stock products expired or near to expired, which were introduced in a box, however, someone else put them back to the shelves by mistake or ignorance thinking that they were new products. No adverse health consequences has been reported for the patient but it can be noticed that the department manager of the hospital leaves the information written in the history of the patient in case he will show any problem in the future. | (b)(4). Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The instructions for use have been reviewed and it is mentioned that the expiry date is printed on the label. It was also mentioned not to use implants after the expiry date. A complaint extract was done regarding expired product implanted: 1 complaint (1 product), this one included, has been recorded on avantage 3p acetabular cup / cementless/ size 62, reference (b)(4), from january 01, 2018 to may 06, 2021. 1 complaint (1 product), this one included, has been recorded on avantage 3p acetabular cup / cementless/ size 62, reference (b)(4), batch 0001100130. Investigation results concluded that the reported event was due to user error (product implanted after expiry date). A summary of the investigation was sent to the complainant conveying proper instructions for use of the device. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11790210&pc=LPH
3006946279-2021-00077 | 04/14/2021 | 05/05/2021 | 21 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | BONE CEMENT | - | It has been reported that, the surgical tech was told to mix cement, (two boxes of cement were to be used in the case) the first vial of monomer was opened followed by the second. As reported, the second vial broke in an odd way to where part of the glass vial was sticking below the plastic sleeve and the glass ended up cutting the surgical tech¿s thumb. It was also reported that the contaminated items were taken off the field and the surgical tech left the case to get the cut addressed. As reported, a new surgical tech scrubbed in and two new boxes of cement were opened, mixed, and the case proceeded. | (b)(4). Report source: event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11776619&pc=LOD
3006946279-2021-00076 | 04/24/2021 | 04/30/2021 | 6 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that, before the surgery , the inner package was leaking no adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11751647&pc=LOD
3006946279-2021-00075 | 04/26/2021 | 04/30/2021 | 4 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that, before the surgery , the inner package was leaking no adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11751646&pc=LOD
3006946279-2021-00074 | 04/24/2021 | 04/30/2021 | 6 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product. Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11751642&pc=MBB
3006946279-2021-00072 | 12/10/2019 | 04/27/2021 | 504 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN PLUS BONE CEMENT 1X40 | BONE CEMENT, ANTIBIOTICS | - | It was reported that the patient had yellowish secretion after elbow arthroplasty surgery. The patient was prophylactically and proactively treated for infection, however infection had never been confirmed nor diagnosed as all cultures taken resulted negative for bacteria/infection, and the surgeon suspects reaction to the cement. Patient underwent an additional procedure to remove graft and wires that may have contributed to the event. | This is a combination product. (b)(4). Concomitant medical products: humeral assembly for cemented use - ref 32810502506 - lot 11022221. Ulnar assembly for cemented use - ref 32810505302 - lot 64122517. Report source, foreign and consumer - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11730872&pc=MBB
3006946279-2021-00073 | 08/14/2020 | 04/27/2021 | 256 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | CEMENT WITH GENTAMYCINE | BONE CEMENT, ANTIBIOTICS | - | It was reported that a patient who was part of polar-total knee arthroplasty clinical study has died. The cause of death is due to massive brain hemorrhage. The exact dat of patient death is not known. | This is a combination product. (b)(4). Report source, foreign and study - event occurred in (b)(6). List of associated products : ps femoral porous, standard, size 9, right; ref : 42500806602 batch : 63111358; persona tibial base plate (right) size f; ref : 42532007502 batch : 63198951; persona ps conv (right) size 7, 10mm + thick; ref : 42521400710 batch : 63470553; persona all poly 35mm patella; ref : 42540000035 batch : 63426281. The device manufacturing quality record can't be done since the batch wasn't communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Reference - batch number wasn't shared. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11730385&pc=MBB
3006946279-2021-00069 | 03/23/2021 | 04/26/2021 | 34 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. No adverse event has been reported as a result of the malfunction. | Cmp- (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11723551&pc=LOD
3006946279-2021-00068 | 03/12/2021 | 04/26/2021 | 45 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. No adverse event has been reported as a result of the malfunction. | Cmp- (b)(4). Report source, foreign - event occurred in (b)6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11723522&pc=LOD
3006946279-2021-00067 | 03/04/2021 | 04/26/2021 | 53 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It has been reported that, during the surgery, the polymer was leaking from the inner pouch because part of the inner polymer pouch was unsealed. No adverse event has been reported as a result of the malfunction. | Cmp-(b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11723299&pc=LOD
3006946279-2021-00071 | 03/26/2021 | 04/23/2021 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BC R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that refobacin cement (lot 919faa0807 handle in the current complaint and lot az11ba1304 handle in cmp- (b)(4) hardened at around 7 mins and customer had to waste a femur and uc surface. | This is a combination product cmp-(b)(4). Event occurred in united states of america. Item number 42-5099-025-25 item name psn 2. 5mm female scr 25mm 2pk lot # 65003862. Item number 00-5901-020-00 item name headless trocar drill pin 75mm lot # 64920190. Item number 42-5026-064-01 item name psn fem cr cmt ccr std sz 8 l lot # 64824129. Item number 42-5320-071-01 item name psn tib stm 5 deg sz e l lot # 64780504. Item number 42-5400-000-29 item name psn all poly pat ply 29mm lot # 64843484. Item number 42-5112-005-12 item name psn asf uc 12mm ply l 4-11 ef lot # 64461197. Item number 42-5026-064-01 item name psn fem cr cmt ccr std sz 8 l lot # 64807940. Item number 42-5112-005-12 item name psn asf uc 12mm ply l 4-11 ef lot # 64461295. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11716174&pc=MBB
3006946279-2021-00070 | 03/26/2021 | 04/23/2021 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that refobacin cement (lot 919faa0807 handle in the current complaint and lot az11ba1304 handle in (b)(4)) hardened at around 7 mins and customer had to waste a femur and uc surface. | This is a combination product (b)(4) report source, foreign and health professional: event occurred in united states of america. List of associated products: item number 42-5099-025-25 item name psn 2. 5mm female scr 25mm 2pk lot # 65003862, item number 00-5901-020-00 item name headless trocar drill pin 75mm lot # 64920190, item number 42-5026-064-01 item name psn fem cr cmt ccr std sz 8 l lot # 64824129, item number 42-5320-071-01 item name psn tib stm 5 deg sz e l lot # 64780504, item number 42-5400-000-29 item name psn all poly pat ply 29mm lot # 64843484, item number 42-5112-005-12 item name psn asf uc 12mm ply l 4-11 ef lot # 64461197, item number 42-5026-064-01 item name psn fem cr cmt ccr std sz 8 l lot # 64807940, item number 42-5112-005-12 item name psn asf uc 12mm ply l 4-11 ef lot # 64461295. This is a combination product. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11715955&pc=MBB
3006946279-2021-00066 | 02/01/2021 | 04/23/2021 | 81 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CMENT R 1X40 JP | BONE CEMENT | - | It was reported that, during the surgery, it was noticed that the polymer (cement powder) was leaking because the polymer pouch had a hole. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11714813&pc=LOD
3006946279-2021-00065 | 04/06/2021 | 04/22/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 52MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | During the impaction of the cup, the impactor broke off in 2 pieces. Nothing fell into the patient and surgeon was able to finish the surgery without any delay. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source: foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description During the impaction of the cup, the impactor broke off in 2 pieces. Nothing fell into the patient and surgeon was able to finish the surgery without any delay. No adverse health consequence has been reported as the result of the malfunction. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The following section has been updated. Received pictures shows that the impactor tip is broken in different parts. The product was returned and lab analysis was performed. The product analysis shows that the product has been returned in spare parts. Indeed, it was noticed that two parts of the piece were broken on the top of the product. Therefore, the reported event has been confirmed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the raw material certificate shows no non conformity or deviation. The instruction for use has been reviewed and it was found that cleaning and sterilization instructions are described in it, like storage and use. A complaint extract was done regarding instrument fracture: 2 complaints (2 products), this one included, have been recorded on avantage impactor tip 52mm, reference a4640052, from (b)(6) 2018 to (b)(6) 2021. 1 complaint (1 product), this one included, has been recorded on avantage impactor tip 52mm, reference a4640052, batch zb160502, since ever. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11705236&pc=HWA
3006946279-2021-00063 | 12/07/2020 | 04/20/2021 | 134 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | STANDARD FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - RIGHT S5 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent an initial hip arthroplasty in 2010. Subsequently, a revision procedure due to major bone destruction at the cotyle level following a massive metallosis was performed on (b)(6) 2020. It was also noticed during this revision surgery that the implant was loosen. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: exceed standard shell ha/pc 052 mm, reference123952ha, batch 2042350, handle in cmp-0685194. Exceed inlay, reference and batch not communicated, handle in cmp-0685194. Ceramic head, reference unknown, batch 564016. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11693873&pc=LPH
3006946279-2021-00062 | 03/19/2021 | 04/20/2021 | 32 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the bone cement became warm after only about 3 minutes, and after 8 minutes the cement was completely hardened. Due to the experience of the surgeon and the surgical technique used, the cement could still be used temperature in the operating room approx. 21°c, the cement was not stored near a heat source. No adverse events have been reported as a result of the malfunction. | This is a combination product cmp-(b)(4). Event occurred in germany. This is a combination product. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11691717&pc=MBB
3006946279-2021-00060 | - | 04/09/2021 | - | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 60MM | ORTHOPAEDIC PROSTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported, by the hospital, that the instrument broke during surgery. No adverse health consequence has been reported as the result of this malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). No further information provided (x-rays, surgical report, photographs, lab test). The product was returned and lab analysis was performed. The product analysis shows that the product has been returned in spare parts. Indeed, it can be noticed that two parts of the piece was broken on the top of the product. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the raw material certificate could not be performed as it was not available. A complaint extract was done regarding instrument fracture: 4 complaints (4 products), this one included, have been recorded on avantage impactor tip 60mm, reference a4640060, from january 01, 2018 to april 08, 2021. 2 complaints (2 products), this one included, have been recorded on avantage impactor tip 60mm, reference a4640060, batch 1229119. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. A summary of the investigation has been transmitted to the customer. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11639909&pc=HWA
3006946279-2021-00061 | 04/01/2021 | 04/09/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. As reported, the product will not be returned to valence due to country regulations. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product. Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11639904&pc=MBB
3006946279-2021-00059 | 12/10/2020 | 04/08/2021 | 119 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS ANATOMIQUE | TOTAL SHOULDER PROSTHESIS | - | This patient had his first surgery several years ago for a tess anatomic (with a cemented glenoid). Doctor rio revised him for a glenoid loosening, with a tess reversed. The patient was doing very well postoperatively, but felt a sudden pain one day, several weeks post op, and it appeared that the glenosphere was dissociated from the baseplate. This dislocation has been reported in (b)(4). No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the lot number was not communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information isfound which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11634603&pc=KWS
3006946279-2021-00058 | - | 04/08/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS GLENOID REVERSE HEAD 41MM | TOTAL SHOULDER PROSTHESIS | - | After a tess reverse implantation, the patient experienced a glenoid sphere disassembly and was revised (this first disassembly has been reported under the (b)(4)). Unfortunately, the patient experienced again a glenosphere disassembly after this revision. The sphere is still implanted. No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information isfound which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11634394&pc=PHX
3006946279-2021-00057 | 02/11/2021 | 04/08/2021 | 56 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS GLEN REVERSE HEAD 41MM | TOTAL SHOULDER PROSTHESIS | - | It has been reported that the patient underwent his first surgery several years ago for a tess anatomic (with a cemented glenoid). Doctor rio revised him for a glenoid loosening, with a tess reversed. The patient was doing very well postoperatively, but felt a sudden pain, several weeks post op, and it appeared that the glenosphere was dissociated from the baseplate. No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11632875&pc=KWS
3006946279-2021-00056 | 02/11/2021 | 04/08/2021 | 56 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | TESS GLENOID CEMENTLESS BASEPLATE S2 | TOTAL SHOULDER PROSTHESIS | - | It has been reported that the patient underwent his first surgery several years ago for a tess anatomic (with a cemented glenoid). Doctor rio revised him for a glenoid loosening, with a tess reversed. The patient was doing very well postoperatively, but felt a sudden pain, several weeks post op, and it appeared that the glenosphere was dissociated from the baseplate. No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11632861&pc=KWS
3006946279-2021-00055 | 03/30/2021 | 04/07/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that the polymer powder contained in the inner pouch leaked. The product is not returned due to the cost of shipment. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11626193&pc=LOD
3006946279-2021-00054 | 04/02/2021 | 04/07/2021 | 5 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It was reported that, during the surgery, the inner sterile packaging was found opened. In addition it was reported that the doctor used a new product to complete the surgery. No adverse events have been reported as a result of the malfunction. | This is a combination product. (b)(4). Foreign - event occurred in (b)(6). This is a combination product. Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11626174&pc=MBB
3006946279-2021-00053 | 03/09/2021 | 04/06/2021 | 28 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a surgeon had implanted this initial persona primary knee (cr / mc combination) together with 1x optipac 40 refobacin bone cement - on (b)(6) 2019 into this patient. However, the implants were revised on (b)(6) 2021 at about 16h45 - at (b)(6) hospital. And was replaced with the nexgen lcck system. Intra-operatively, while revising the components - the bone cement was visually still perfectly intact and very strongly attached to both prostheses. However, both persona components were loosened from the bone, with no bondage interface to the bone. The surgeons could offer no explanation to the loosening, but could not rule out possible allergic reaction from the patient on a physiological level, only as speculation? the patient¿s bone seemed to have resorbed at the interface level on visual inspection? and it was also noted that the bone quality underneath the removed prostheses ¿ seemed to be very sclerotic indeed, especially on the femoral side. Upon intra-operative visual inspection and according to prior blood tests before the revision - had suggested & shown no infection markers. For reference purposes: a post-operative x-ray a few months after the primary implantation taken on (b)(6) 2019. As well as a more recent post-operative x-ray on (b)(6) 2021 - which clearly show the components¿ loosening. Implant failure cannot be explained by the surgeon at this stage. And the theory of a potential infection shortly after the primary knee implantation in 2019 could not be ruled out / however, it could not be confirmed? although, recent tests have shown no infection markers. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). This is a combination product list of associated devices: femur cr, cemented, left 9, standard, reference 42-5026-066-01 and lot no. 64193285; tibia, cemented, left, size f, reference 42-5320-075-01 and lot no. 64111819; articular surface mc, 10mm (tib. Ef, fem. Cr 8-11) reference 42-5121-008-10 and lot no. 63949882. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11619121&pc=MBB
3006946279-2021-00052 | 01/13/2021 | 04/06/2021 | 83 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 JP | BONE CEMENT | - | It was reported that, during the surgery, when the outer bag of powder was opened, the inner bag had a crack. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11619017&pc=LOD
3006946279-2021-00050 | - | 04/02/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | T.E.S.S. REVERSED COROLLA S2 | TOTAL SHOULDER PROSTHESIS | - | It has been reported that the patient underwent a dislocation after a revision surgery. A second revision surgery is planed. | (b)(4). List of associated devices: baseplate, s, 15mm, reference 0104440002, batch 3019007; tess hum insrt 6x41 sz 2, reference p1700188, batch unknown; glenosphere dia 41 deg 0, reference 0104440070, batch unknown; dpsc screw 3. 5/4. 5x30mm, reference 0104440030, batch unknown; dpsc screw 3. 5/4. 5x30mm, reference 0104440030, batch unknown. Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the lot number was not communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11606998&pc=KWS
3006946279-2021-00051 | 03/02/2021 | 04/01/2021 | 30 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the powder package had a hole and consequently the cement powder was leaking. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11601221&pc=MBB
3006946279-2021-00045 | 07/28/2020 | 03/31/2021 | 246 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COBALT CHROME FEMORAL HEAD 542 DIAMETER 22,2/ 12-14/ +2/ L | METALLIC FEMORAL HEAD PROSTHESIS | - | It was reported that the patient was received on (b)(6) 2020 at the clinic and underwent on (b)(6) 2020 a cleaning and washing of the joint following a bacterial infection to staphylococcus epidermidis, which is a normal colonizer of skin. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: avantage inlay s44 / 22,2 ; item # p0560044; lot # 0001439499; manufactured by zb valence. Avantage reload cup ti ha s44; item #p0460p44; lot # 0001219576; manufactured by zb valence. Exception std stem left size 1; item #ps125y01; lot #0001270155; manufactured by zb valence. The device manufacturing quality records all indicate that the released products met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11593995&pc=JDI
3006946279-2021-00044 | 07/28/2020 | 03/31/2021 | 246 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | STANDARD FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - LEFT SIZE 1 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that the patient was received on (b)(6) 2020 at the clinic and underwent on (b)(6) 2020 a cleaning and washing of the joint following a bacterial infection to staphylococcus epidermidis, which is a normal colonizer of skin. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: avantage inlay s44 / 22,2 ; item # p0560044; lot # 0001439499; manufactured by zb valence. Avantage reload cup ti ha s44; item #p0460p44; lot # 0001219576; manufactured by zb valence. Cocr fm hd d22,2/ 12-14/ +2/ l ; item #p0206l22; lot #j6382918; manufactured by zb valence. The device manufacturing quality records all indicate that the released products met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11593986&pc=LZO
3006946279-2021-00028 | 07/28/2020 | 03/31/2021 | 246 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE RELOAD CUP TI HA S44 | ACETABULAR SHELL | - | It was reported that the patient was received on (b)(6) 2020 at the clinic and underwent on (b)(6) 2020 a cleaning and washing of the joint following a bacterial infection to staphylococcus epidermidis, which is a normal colonizer of skin. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: avantage inlay s44 / 22,2 ; item # p0560044; lot # 0001439499; manufactured by zb valence. Exception std stem left size 1; item #ps125y01; lot #0001270155; manufactured by zb valence. Cocr fm hd d22,2/ 12-14/ +2/ l ; item #p0206l22; lot #j6382918; manufactured by zb valence. The device manufacturing quality records all indicate that the released products met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11593982&pc=KWA
3006946279-2021-00049 | 11/06/2019 | 03/30/2021 | 510 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 1X40 US | ORTHOPAEDIC CEMENT, NON-MEDICATED | - | It was reported that the patient received an initial left total knee arthroplasty. The patient was revised 5 years later for an unknown reason. Subsequently, the patient was revised again 3 years after the first revision. The patient was revised again as part of a two stage revision due to infection less than one year post operative. | This is a combination product (b)(4). Report source, foreign - event occurred in united states. Relevant accessories : biomet cc cruciate tray - ref : 141236 et lot : n/a; vngd ps open intl fem rt 75 - ref : 183114 et lot : n/a; series a pat std 37 3 peg - ref : 184768 et lot : n/a; vngd ps tib brg 10x79/83mm - ref : 183660 et lot : n/a. The device manufacturing quality record can't be performed as the batch number is not communicated. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Batch number not communicated. | No Answer Provided | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11589263&pc=LOD
3006946279-2021-00048 | 03/04/2021 | 03/30/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It has been reported that the optipac cement stays rather liquid for a long time after mixing. No known adverse event was reported. | This is a combination product (b)(4). Report source, foreign, health professional - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11587309&pc=MBB
3006946279-2021-00047 | 03/03/2021 | 03/30/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported cement was already hardened completely after 5 minutes. Product not available for return (discarded by customer at hospital). No patient information provided. No known adverse event was reported. | This is a combination product (b)(4). Report source (b)(4). The device manufacturing quality record can't be performed as the correct batch number wasn't communicate. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. The batch number wasn't correct. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11587292&pc=MBB
3006946279-2021-00046 | 03/03/2021 | 03/30/2021 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the cement was already hardened completely after 5 minutes. Product not available for return (discarded by customer at hospital). No patient information provided. No known adverse event was reported. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11587282&pc=MBB
3006946279-2021-00043 | 02/15/2021 | 03/29/2021 | 42 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AURA II HIP HA COATED RIGHT SIZE 6 | PROSTHESIS, HIP | - | It was reported that a patient underwent a revision surgery due to a fracture of the implant following a false movement and a fall. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the lot number was not communicated / not correct. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11580929&pc=LZO
3006946279-2021-00039 | 03/10/2021 | 03/26/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that during the surgery, the inner cement sterile packaging was found to be open. As reported as well, the surgery was delayed for 26 minutes. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). This is a combination product. Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11573690&pc=MBB
3006946279-2021-00040 | 03/10/2021 | 03/26/2021 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that during the surgery, the inner cement sterile packaging was found to be open. In addition it was reported that the surgery was delayed for 26 minutes. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11573674&pc=MBB
3006946279-2021-00041 | 03/22/2021 | 03/26/2021 | 4 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found open. Three product boxes were opened and all have shown this same problem. The products will not be returned due to applicable regulations of the country. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11573478&pc=MBB
3006946279-2021-00042 | 03/16/2021 | 03/26/2021 | 10 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the inner sterile packaging was found open. Nine product boxes were opened and all have shown this same problem. All the products will not be returned due to applicable regulations of the country. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11573471&pc=MBB
3006946279-2021-00032 | 02/23/2021 | 03/23/2021 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | During the later part of surgery, the nurse began to mix the cement, but when she was done mixing, she was unable to break the mixing rod. Then she tried to squeeze the cement out of the mixing container and the cement gun, but the cement hardened in 1 minute. The normal recommended cement mixing procedure was followed in the mixing sequence and all steps were followed. No known adverse event was reported. | This is a combination product (b)(4). Report source, foreign event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11548653&pc=MBB
3006946279-2021-00033 | 02/24/2021 | 03/22/2021 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the plunger has not moving up automatically after releasing. The consistency of the product was different than expected. There was not enough adhesion to the femoral component. The surgery was completed successfully with new product. There was no patient involvement. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11541386&pc=MBB
3006946279-2021-00035 | 03/05/2021 | 03/19/2021 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, during the surgery, a pack of refobacin bone cement was opened to the scrub nurse in or. As reported, the pack was made of 2 packages, one non sterile paper package and one sterile package. It was reported as well that the paper package was peeled in order to take out the sterile package. As reported, when the non sterile package was opened from the top and revealed the sterile cement pack, the scrub nurse grabbed and pull it out. As reported as well, it was then noticed that there was a leakage of cement powder at the bottom of the inner package. In addition, a video showing the issue has been provided by the complainant. It was also reported that the cement package has been thrown away due to contamination. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). One video and one picture (extracted from the video) have been received. The reported event was confirmed as it can be seen that the cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11522983&pc=MBB
3006946279-2021-00036 | 02/03/2021 | 03/19/2021 | 44 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that 44 bone cement products were sent from the warehouse on february 3 for the experimental test of montagne (used to simulate the tensile and shear forces of knee joint products), and for 20 of the 44 products, the inner pouches leaked powder in total. As reported, the powder leaking products have been scrapped, and the customer service has reissued 20 packages of products with the same batch number. In addition, it was reported that no surgery was involved (i. E. R&d testing). No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11522960&pc=MBB
3006946279-2021-00038 | 03/10/2021 | 03/18/2021 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that, before the surgery, it was found (by zb staff) that the inner package had leaked. In addition it was reported that the product will not be returned due to the cost of shipment. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). One picture has been received and the reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11517176&pc=LOD
3006946279-2021-00037 | 03/09/2021 | 03/18/2021 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, during the surgery, the inner cement sterile packaging was found to be opened. As reported, the doctor performed the surgery using a spare bone cement product. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source foreign: (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11517119&pc=MBB
3006946279-2021-00034 | 02/22/2021 | 03/18/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the inner package leaked. As reported, this issue was found before the surgery (by zb staff). In addition, it was reported that the item will not be returned due to the cost of shipment. No adverse event has been reported as a result of the malfunction. | (b)(4) report source, foreign - event occurred in (b)(6). One picture has been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11517058&pc=LOD
3006946279-2021-00026 | 02/18/2021 | 03/15/2021 | 25 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN PLUS BONE CEMENT-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that during left total knee arthroplasty following implant of the femoral component, in deep flexion position, the femur moved distally. Cement had not adhered to patients bone or femoral implant. Cement time was around 11-12 minutes, when femoral implant became loose and was removed. Surgeon assessed the knee and required to downsize femur to a size 7 left standard trabecular metal implant, which was a better fit. Ultra congruent poly and down sized femur to match the correct size for the tibia. There was a 60 minute delay in the procedure trouble shooting for new femur and re trialing implants. Patients posterior cruciate ligament was sacrificed and used ultra congruent poly instead of cruciate retaining. | Cmp-(b)(4). Natural tibia cemented 5 degree stemmed left size e, reference (b)(4), batch 64794770. Event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11485453&pc=MBB
3006946279-2021-00030 | 02/22/2021 | 03/15/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the inner package leaked. As reported, this issue was found before surgery. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source: foreign - event occurred in (b)(6). One picture has been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11479387&pc=LOD
3006946279-2021-00031 | 02/22/2021 | 03/15/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the inner package leaked. As reported, this issue was found before surgery. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). One picture has been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11479385&pc=LOD
3006946279-2021-00029 | 12/15/2015 | 03/12/2021 | 1914 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN REVISION | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported in a retrospective study of the nexgen ap wedge tibial prosthesis that one of the patients underwent a revision on the (b)(6) 2015 of the study product following the prosthesis loosening, involving cement. The surgeon listed the event as not related to the instrument, uncertainly related to the device and uncertainly related to the procedure. | This is a combination product (b)(4). Report source, foreign, study - event occurred in (b)(6). List of associated devices: femoral component option for cemented use only size d right, reference 00599401492, batch unknown. Stemmed tibial component precoat a/p wedged for cemented use only size 3, reference 00598800300, batch 62266219. Articular surface with locking screw "size yellow/c d 12 mm height", reference 00599403012, batch unknown. All poly patella standard size 32 mm diameter 8. 5 mm thickness cemented, reference 00597206532, batch unknown. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11473131&pc=MBB
3006946279-2021-00025 | 02/09/2021 | 03/04/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that, before the surgery, the inner package leaked. In addition, as reported, product will not be returned due to the cost of shipment. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source: foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11408928&pc=LOD
3006946279-2021-00024 | 02/09/2021 | 03/04/2021 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It was reported that, before the surgery, the inner package leaked. In addition, as reported, product will not be returned due to the cost of shipment. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11408913&pc=LOD
3006946279-2021-00027 | - | 03/02/2021 | - | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS FEMORAL STEM | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported through the article fixation of the short global tissue-sparing hip stem from the bone & joint journal that gts stems showed an initial migration pattern of varization (x-translation combined with z-rotation) and posterior translation (z translation). No adverse event has been reported for patients involved in this study. | (b)(4). Report source: foreign - event occurred in (b)(6). Report source: literature - this event has been identified from the journal article of m. J. Nieuwenhuijse, s. B. W. Vehmeijer, n. M. C. Mathijsen and s. B. Keizer (2020) "fixation of the short global tissue-sparing hip stem - two-year follow-up results of a randomized controlled clinical and roentgen stereophotogrammetric analysis study",the bone and joint journal, vol. 102-b(6); 699¿708. Doi: 10. 1302/0301-620x. 102b6. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11398708&pc=JDI
3006946279-2021-00021 | 03/04/2020 | 03/01/2021 | 362 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that patient had pain when bearing weight. In addition, it was reported that the cup could be involved. No medical treatment has been provided to patient. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: g7 bispherical shell 58g; ref#110017335; lot#6527436; g7 neutral e1 liner 36mm e; ref#010000857; lot#6524503; delta cer fm hd 036/0mm12/14 m 12/14; ref#650-0837; lot#2019071399. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11391276&pc=KWL
3006946279-2021-00023 | 09/08/2020 | 02/25/2021 | 170 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | FEMORAL STEM AURA II TOTAL HA / SIZE 8 RIGHT / CEMENTLESS / 542 / 12-14 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | As reported, the patient is a male individual operated on the right side on (b)(6) 2013 with a primary total hip arthroplasty for a primary osteoarthritis. It was reported as well that the patient suffered from a periprosthetic neck fracture after a fall on the hip and was operated for a stem-head revision with explantation of the femoral component, followed by osteosynthesis and implantation of a revision femoral stem on (b)(6) 2020. In addition, as reported, the acetabular component is intact and remains in the patient. | (b)(4). List of associated devices: avantage reload acetabular cup / cementless / size 58, reference p0460p58, batch unknown; avantage e1 insert size 58 ø 28, reference p0561e58, batch unknown; biomet cobalt chrome, size 28, reference and batch unknown. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11378083&pc=KWA
3006946279-2021-00022 | 02/01/2021 | 02/25/2021 | 24 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE INLAY SIZE 54 28 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that a revision surgery was required for hip dislocation between the head and the avantage dual mobility liner. A new liner and head were implanted. It can be noticed that the surgical technique for the primary procedure was followed by a surgeon who uses many avantage per year (50+). | (b)(4). List of associated devices: femoral head +3. 5x28mm dia, reference 00-8018-028-03, batch 64482571. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11377627&pc=LZO
3006946279-2021-00020 | 07/28/2020 | 02/25/2021 | 212 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE INLAY SIZE 44 22,2 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported that the patient underwent a revision surgery following an unspecified infection. No other information available. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11377511&pc=KWA
3006946279-2019-00562 | 07/04/2019 | 02/23/2021 | 600 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CURVED HANDLE SHAFT | HIP, INSTRUMENT | - | It was reported that during the surgery the instrument was broken due to impaction. No adverse event was reported due to the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Reported event was confirmed as the pictures provided show the product broken. The product evaluation showed that the avantage curved handle shaft broke at the weld between the straight part and the curved part. It is possible that the two parts were not sufficient welded or that the weld was damaged by the polishing. Therefore, the product evaluation confirmed the reported event and the most probable cause of the event is a manufacturing issue. Dhr review : the review of the device manufacturing quality record indicates that 10 products avantage curved handle shaft, reference 110031338, batch 056790 were manufactured on 01 february 2018. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. No non conformity or deviation was observed which could be linked to the event described in the complaint. Only this complaint has been recorded for avantage curved handle shaft, batch 056790 within one year. According to available data, the most probable root cause could be a manufacturing issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions; if any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11363991&pc=LXH
3006946279-2021-00018 | 01/29/2021 | 02/22/2021 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that, during the surgery, the product was examined and the inner sterile packaging was opened. As reported, the surgery was delayed for 15 minutes and the doctor performed the surgery using spare bone cement. No adverse event has been reported as a result of the malfunction. | This is a combination product (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11357046&pc=MBB
3006946279-2021-00017 | 01/22/2021 | 02/22/2021 | 31 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | It was reported that during the surgery, the product was examined and it was found that the powder leaked out of the inner pouch. It was reported as well that the doctor performed the surgery using a spare product. No adverse event has been reported as a result of the malfunction. | This is a combination product. (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11357033&pc=MBB
3006946279-2020-00238 | 11/26/2020 | 02/19/2021 | 85 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that during the operation, it was found that the bottom of the aseptic package of bone cement was opened and no powder leaked. No adverse events have been reported as a result of the malfunction. | This is a combination product (b)(4). Report source: foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed as it remains implanted. Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11347510&pc=MBB
3006946279-2021-00019 | 10/31/2020 | 02/19/2021 | 111 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM CEMENTLESS S+3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that patient had pain in the left hip when starting to walk. In addition it was reported that no medical treatment was provided to the patient. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: g7 bispherical shell 54f; ref#110017333; lot#6705272; g7 neutral e1 liner 36mm f; ref# (b)(4); lot#6744914; delta cer fm hd 036/0mm 12/14; ref#(b)(4); lot#2020021626. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11347420&pc=MEH
3006946279-2021-00012 | 06/01/2011 | 02/12/2021 | 3544 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC | OPTIPAC 40 REFOBACIN BONE CEMENT R | - | It was reported that the patient underwent an initial left knee arthroplasty. Subsequently, it was discovered that free cement fragment was blocking the joint. An arthroscopic surgery was performed to remove the loosened cement fragment. No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. With the available information, the event could not be confirmed. The review of the device manufacturing quality record indicates that optipac 40 refobacin bone cement r, reference (b)(4), lot number 001bal1403 were manufactured on 2010, 23 april. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. No non conformity or deviation was observed which could be linked to the event described in the complaint. No other similar complaint has been recorded on optipac 40 refobacin bone cement r, reference (b)(4), from (b)(6) 2018 to (b)(6) 2021. No other similar complaint has been recorded has been recorded on optipac 40 refobacin bone cement r, reference (b)(4), batch 001bal1403. With the available information, the exact root cause of the event could not be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11322109&pc=JDZ
3006946279-2021-00016 | 12/08/2020 | 02/10/2021 | 64 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, during the surgery, two products were examined and for both of them the powder leaked out the inner pouches. It was reported as well that the surgeon performed the surgery using spare product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event can't be confirm as it can't be seen any damage of the inner pouch or cement powder leak. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11306019&pc=MBB
3006946279-2021-00015 | 11/27/2020 | 02/10/2021 | 75 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, during the surgery, two products were examined and for both of them the powder leaked out the inner pouches. It was reported as well that the doctor performed the operation using a spare product. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11306004&pc=MBB
3006946279-2021-00014 | 06/14/2020 | 02/09/2021 | 240 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | FEMORAL STEM AURA II HA / SIZE 7 RIGHT / CEMENTLESS / 542 / 12-14 | PROSTHESIS, HIP | - | It was reported that patient was implanted on an unknown date estimated between 2001 and 2015 and underwent revision surgery on (b)(6) 2020 following a fracture of the left femoral stem on (b)(6) 2020. Please note that this event was initialy reported by zimmer (b)(4) manufacturing gmbh on december 03, 2020 and registered in your database under the number 0009613350-2020-00565. However, additional information has been received concerning the product brand involved and the legal manufacturer had to be changed to biomet france. Therefore, report number 0009613350-2020-00565 has to be voided and the case is now handled in report number 3006946279-2021-00014. | (b)(4). Foreign report source: event occurred in (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11301815&pc=JDI
3006946279-2021-00011 | 12/23/2020 | 02/09/2021 | 48 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, before the surgery, the product was examined and it was found that the powder leaked out of the inner pouch. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11299773&pc=MBB
3006946279-2021-00013 | 11/24/2020 | 02/04/2021 | 72 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | STANDARD FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - LEFT S4 | PROTHESIS, HIP | - | It was reported that a patient underwent revision surgery on (b)(6) 2020 following the fracture of its exception stem prothesis at the distal level while he was at home. Patient suffered from pain. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the raw material certificate shows no non conformity or deviation. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported that a patient underwent revision surgery on (b)(6) 2020 following the fracture of its exception stem prothesis at the distal level while he was at home. Patient suffered from pain. Manufacturer Narrative (b)(4). List of associated devices: avantage reload cup ti ha s54, reference (b)(4), batch 0000848839. Avantage e1 inlay s54 / 28, reference (b)(4), batch 0000835979. Alumina femoral head d28/0/5 42/12-14, reference (b)(4), batch 0000849679. This follow-up report is being submitted to relay additional information. Both initial and revision surgeries notes have been received. Regarding the initial surgery, nothing abnormal has been noticed. Regarding the notes of the revision surgery, it has been noticed that the stem was loosen in its upper part but well osteointegrated in its lower part which might explain the fracture of the stem at the distal level due to mechanical forces exerted on the lower part. Finally, an anatomopathology report has been received dated before the initial hip arthroplasty, which stated that the patient had a lymphoma treated with chemotherapy and corticosteroids. Moreover, it is noticed that the patient had coxarthroses which led to important osteonecrosis. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the raw material certificate shows no non conformity or deviation. A complaint extract was done regarding revision due to implant fracture following loosening of the upper part of the stem: 1 complaint (1 product), this one included, has been recorded on standard femoral stem exception -cementless- 5°42- 12/14 - left s4, reference (b)(4), from (b)(6) 2018 to (b)(6) 2021. 1 complaint (1 product), this one included, has been recorded on standard femoral stem exception -cementless- 5°42- 12/14 - left s4, reference (b)(4), batch 0000858950. The product was returned and lab analysis was performed. The product analysis shows that the product has been returned in spare parts. Indeed, it was noticed that the stem was broken in two parts. The root cause of the breakage could not be determined. The review of the instruction for use of the exception stem shows that avantage cup, avantage inlay and alumina femoral head used are compatible with the stem. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. Indeed, the fact that the stem was loosen in its upper part but well osteointegrated in its lower part might explain the fracture of the stem at the distal level due to mechanical forces exerted on the lower part. However, the reason why the stem did not osteointegrates in its upper part could not be determined. The patient has been requested to provide x-rays. The investigation will be re-opened upon receipt of these x-rays. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11277393&pc=LPH
3006946279-2021-00010 | 01/15/2021 | 02/01/2021 | 17 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | It has been reported that, before the surgery, the product was exanimated and it was found that the powder leaked out of the inner pouch. No adverse event has been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11257917&pc=MBB
3006946279-2021-00009 | - | 01/28/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COBALT CHROME FEMORAL HEAD 542 22,2 / 0 / 12-14/ COL MOYEN | PROTHESIS, HIP | - | It was reported that patient underwent initial hips arthroplasties of the left and the right sides in (b)(6) and then (b)(6) 2018 following cortisone osteonecrosis. Since the implantation of the both prosthesis, patient suffers from chronic pain at both sides. Patient can not walk longer than 2000m or hold 2-4kgs. The pain is also affecting his professional career. It can be noticed that pain is at the same place as the prosthesis. The current report handle the head component of the right hip. | (b)(4). List of associated devices: femoral stem cmk oc modular cemented 12/14 s303 l140, reference 3553. 303, batch 0001128937, handle in (b)(4). Novae e 53 th, reference rm45050007, batch 1607282a. Vis corticale autotaraudeuse de 5, reference rm65150047, batch 1610417a. Universal cement restrictor-disposable inserter, reference b000-1240, batch 6m9409. Ci 55/22,2 e, reference rm51100007, batch 1712139a. Palacos r+g 1x40, reference 66017750, batch 89324686. (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11248180&pc=JDI
3006946279-2021-00008 | - | 01/28/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | FEMORAL STEM CMK OC MODULAR CEMENTED 12/14 S303 L140 | PROTHESIS, HIP | - | It was reported that patient underwent initial hips arthroplasties of the left and the right sides in (b)(6) and then (b)(6) 2018 following cortisone osteonecrosis. Since the implantation of the both prosthesis, patient suffers from chronic pain at both sides. Patient can not walk longer than 2000m or hold 2-4kgs. The pain is also affecting his professional career. It can be noticed that pain is at the same place as the prosthesis. The current report handle the stem component of the right hip. | (b)(4). Concomitant medical products: cocr fm hd d22,2/ 12-14/ 0/ m, reference p0206m22, batch raj6166223, handle in (b)(4). Novae e 53 th, reference rm45050007, batch 1607282a. Vis corticale autotaraudeuse de 5, reference rm65150047, batch 1610417a. Universal cement restrictor-disposable inserter, reference b000-1240, batch 6m9409. Ci 55/22,2 e, reference rm51100007, batch 1712139a. Palacos r+g 1x40, reference 66017750, batch 89324686. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11248119&pc=LZO
3006946279-2021-00006 | 12/30/2020 | 01/28/2021 | 29 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that a patient using bbc optipac 80g developed a cerebral infarction 2 days after surgery. The doctor in charge said that it was not possible to determine whether the cause was due to the use of cement, the long operation time, or the old age of (b)(6). | (b)(4). Report source foreign: (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11248101&pc=MBB
3006946279-2021-00007 | - | 01/28/2021 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COBALT CHROME FEMORAL HEAD 542 22,2 / 0 / 12-14/ COL MOYEN | PROTHESIS, HIP | - | It was reported that patient underwent initial hips arthroplasties of the left and the right sides in (b)(6) and then (b)(6) 2018 following cortisone osteonecrosis. Since the implantation of the both prosthesis, patient suffers from chronic pain at both sides. Patient can not walk longer than 2000m or hold 2-4kgs. The pain is also affecting his professional career. It can be noticed that pain is at the same place as the prosthesis. The current report handle the head component of the left hip. | (b)(4). Concomitant medical products: femoral stem cmk oc modular cemented 12/14 s303 l140, reference 3553. 303, batch 0001172941, handle in (b)(4). Novae e 55 th, reference rm45050008, batch 1501124a. Vis corticale autotaraudeuse de 6, reference rm65150047, batch 1703172a. Universal cement restrictor-disposable inserter, reference b000-1240, batch 6m9531. Ci 55/22,2 e, reference rm51100008, batch 1412305a. Palacos r+g 1x40, reference 66017750, batch 88654786. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11248091&pc=LPH
3006946279-2021-00005 | 01/14/2021 | 01/25/2021 | 11 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It was reported that before the operation, the sterile package of bone cement product was inspected and powder leakage was found. When other bone cements were opened, powder leakage was found in the same batch. 5 boxes of bone cement of the same batch are impacted. No adverse events have been reported as a result of the malfunction. | Cmp(b)(4). Report source, foreign - event occurred in (b)(6). 5 pictures were received. On the first picture, it can be seen the aluminum pouch with the cement pouch batch number (a925c), which comply with the batch reported in the complaint. The inner pouch sealing is discontinuous and powder leaked. On the 2nd picture, it can be seen the folding box with the batch number (a925ca2503) which comply with the batch reported in the complaint, the aluminium pouch, the monomer ampoule and the patient labels. The inner pouch is not visible but powder is visible. On the 3rd picture, it can be seen the folding box with the batch number (a925ca2503) which comply with the batch reported in the complaint, the aluminium pouch, the monomer ampoule and the patient labels. The inner pouch is not visible. On the 4th picture, it can be seen the aluminum pouch with the cement pouch batch number (a925c), which comply with the batch reported in the complaint. The inner pouch sealing is damaged and seems to be discontinuous on the corner. On the 5th picture, it can be seen the aluminum pouch with the cement pouch batch number (a925c), which comply with the batch reported in the complaint. The product analysis can't be performed as the product was not returned the complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11221775&pc=MBB
3006946279-2021-00004 | 01/15/2021 | 01/25/2021 | 10 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that before the surgery, over eleven products, the powder leaked out the pouches. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11221656&pc=MBB
3006946279-2021-00003 | 01/06/2021 | 01/25/2021 | 19 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that during the surgery, the cement hardened too quickly. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it was not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11221649&pc=MBB
3006946279-2021-00002 | 01/04/2021 | 01/25/2021 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation. A delay of 46 minutes has been reported. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures has been received. The reported event can't be confirm as it can't be seen any damage of the inner pouch or cement powder leak. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11221642&pc=MBB
3006946279-2021-00001 | 12/21/2020 | 01/05/2021 | 15 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION STANDARD STEM CEMENTED LEFT SIZE 2 | HIP, PROTHESIS | - | It was reported that a patient was reviewed during follow-up visit for retrospective study on the exception cemented stems. It was therefore found that patient suffered from femur cortical hypertrophy of zone 5-6 without implant loosening. The hypertrophy was seen on x-ray. No symptoms either pain were found. No treatment was provided. | (b)(4). List of associated devices avantage reload cup ti ha s48, reference p0460p48, batch 0000803673, avantage e1 inlay s48 / 22,2, reference p0560e48, batch 0000822612, cocr femoral head d22,2/ 12-14/ -2/ s, reference p0206c22, batch 00j2970687, refobacin bone cement r 40x1, reference 3003940001, batch 136eaa3012, optiplug 10mm, reference 4361, batch 200849b, (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11116295&pc=JDI
3006946279-2020-00243 | 01/24/2020 | 12/22/2020 | 333 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AURA II FEMORAL STEM SIZE 6 | PROTHESIS, HIP | - | It was reported through patient's legal representative that the patient was implanted with an aura ii stem, a crco head and a magnum acetabular component on (b)(6) 2010. Following this total hip arthroplasty, the patient suffered from pain, however, he got back to professional life. On (b)(6) 2017, the patient fell and the prosthesis has decompensated. It was therefore seen through x-ray that the patient suffered from bipolar geodic reaction following a release of ions by the prothesis. It was reported that the metal friction between the prosthesis components has led to inflammatory granulomas with a phenomen of bone resorption. On (b)(6) 2020, the patient underwent a revision surgery to remove the prothesis. Based on current data it cannot be determined whether both the acetabular system and the hip stem were removed. | (b)(4). List of associated devices : magnum cup, reference and batch unknown, handle in (b)(4), manufactured by biomet (b)(4) ltd; cobalt chrome head, reference and batch unknown, handle in (b)(4), manufactured by biomet (b)(4) ltd. Establishment name : (b)(6). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11055851&pc=OQI
3006946279-2020-00244 | - | 12/22/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | FEMORAL STEM CMK OC MODULAR CEMENTED 12/14 S303 L140 | PROTHESIS, HIP | - | It was reported that patient underwent initial hips arthroplasties of the left and the right sides in (b)(6) and then (b)(6) 2018 following cortisone osteonecrosis. Since the implantation of the both prosthesis, patient suffers from chronic pain at both sides. Patient can not walk longer than 2000m or hold 2-4kgs. The pain is also affecting his professional career. It can be noticed that pain is at the same place as the prosthesis. At this stage of investigation, no detail was provided concerning the parts implanted. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11055945&pc=JDI
3006946279-2020-00242 | 09/14/2020 | 12/22/2020 | 99 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 2X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that when opening refobacin bone cement outer powder pouch the welding was to strong and it was not possible to open in a correct sterile manner. They had to take a new one to finish the surgery. After investigation, the category has been updated as inner cement pouch open sealing and the event becomes reportable. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Three pictures have been received. One of the picture shows that a outer pouch is damaged. The reported event seems to be confirmed. The product was returned with only one pouch instead of two and lab analysis was performed. The product analysis shows that the inner cement pouch sealing is damaged. However, the outer pouch has been open in a correct sterile manner without any damaged of the sealing. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11055395&pc=MBB
3006946279-2020-00237 | 11/24/2020 | 12/17/2020 | 23 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 2X40G | BONE CEMENT | - | It has been reported that the issue is related to the inner cement pouch open sealing and the powder leaked when they opened the external box. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11026866&pc=LOD
3006946279-2020-00239 | 12/08/2020 | 12/17/2020 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CMT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It has been reported that the cement was crumbly after mixing, not as smooth as usual. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer as it has been discarded. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11026769&pc=LOD
3006946279-2020-00241 | 12/09/2019 | 12/16/2020 | 373 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE INLAY PE DUAL MOBILITY | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported through ansm report (b)(4) that the patient underwent a second revision surgery on (b)(6) 2019 due to a dislocation between the inox head from biomet and the avantage pe inlay dual mobility. It can be noticed that the patient previously underwent a first revision surgery at annecy also due to a multiple luxation of the total hip prothesis with a major intra articular metallosis with a non negligeable septic risk. | (b)(4). Concomitant medical products: avanatge cup, reference and batch not communicated, inox head, reference and batch not communicated. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11017759&pc=KWA
3006946279-2020-00240 | 10/08/2020 | 12/16/2020 | 69 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that when using optipac, an attempt was made to insert it into the medullary cavity 5 minutes after mixing. However, the cement was hardened and its use was stopped. No adverse event has been reported as a consequence of this event. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=11017690&pc=LOD
3006946279-2020-00234 | 11/17/2020 | 12/13/2020 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 2X40G | BONE CEMENT | - | It has been reported that when the user opened the external box, there was powder spread into it. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10999581&pc=LOD
3006946279-2020-00236 | - | 12/10/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION FEMORAL STEM | PROTHESIS, HIP | - | It was reported through siris outlier report 2020 on exception + exceed cups there was a revision due to periprosthetic fracture of the femure within two years after the first implantation. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10983657&pc=JDI
3006946279-2020-00235 | - | 12/10/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | EXCEPTION FEMORAL STEM | PROTHESIS, HIP | - | It was reported through siris outlier report 2020 on exception + eceed cups there was a revision due to infection within two years after the first implantation. No other health consequence has been reported for the patient. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10983506&pc=JDI
3006946279-2020-00233 | 01/21/2019 | 12/09/2020 | 688 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | VARIZED FEMORAL STEM EXCEPTION RIGHT SIZE 6 | HIP, PROTHESIS | - | It was reported that the patient fell and broke his femur around the implant (periprostetic fracture) on the (b)(6) 2019. On the (b)(6) 2019, the stem was revised in another center ((b)(6)) and replaced by a locked revision stem. The cup was not revised. On the (b)(6) 2019, the surgeon saw the patient again. The patient is very satisfied since the revision, walk and go his daily activities normally and without pain. He says that there was no relation with the implant, instrumentation, or procedure. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: alu femoral head d28/+3. 5/5 42/12-14, reference p0209l28, batch 0000399442; avantage reload cup ti ha s54, reference p0460p54, batch 0000349027; avantage inlay s54 / 28, reference p0561054, batch 0000374935. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10974251&pc=KWY
3006946279-2020-00232 | 11/06/2020 | 12/07/2020 | 31 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 BIOMET BONE CEMENT R | BONE CEMENT | - | It was reported that the cement hardened before the stem was set in the target position. The surgeon attempted to insert the stem in 5 minutes, but was informed that it did not reach the target position. It was removed because it could not be reduced. The cement could not be completely removed. No additional patient consequences were reported. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10957228&pc=LOD
3006946279-2020-00231 | 11/17/2020 | 12/02/2020 | 15 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that two optipac 80 were mixed in succession, in both cases the cement was difficult to mix. After 4 minutes, the cement had already hardened and had to be discarded. This mixing issue has led to a delay of 40 min during the surgery. No adverse event has been reported for the patient. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10934456&pc=MBB
3006946279-2020-00227 | 11/02/2020 | 12/01/2020 | 29 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation on four bone cements. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4); foreign source: event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event was confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Ten complaints have been recorded on refobacin bone cement r 1x40-3, reference#: (b)(4), over the batch#: a926ba2506. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be sent regarding the reported issue. If any further information found, which would change, alter any conclusions, or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10926323&pc=MBB
3006946279-2020-00230 | 11/04/2020 | 12/01/2020 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It was reported that the inner sterile packaging was found opened during the operation. The doctor use new product complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4) report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouches. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10926301&pc=MBB
3006946279-2020-00229 | 10/23/2020 | 12/01/2020 | 39 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It was reported that the cement powder pack wasn't sealed well and leaked. The event happened before the surgery. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Picture has been received. The reported event was confirmed as it can be seen that the inner cement pouch sealing is damaged and that cement powder leaked outside the inner pouch. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10926239&pc=MBB
3006946279-2020-00228 | 10/28/2020 | 11/30/2020 | 33 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTICS | - | It was reported that the inner sterile packaging was found opened during the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). Pictures have been received. The reported event can't be confirm as it can't be seen any damage of the inner pouch or cement powder leak. The product analysis can't be performed as the product was not returned. The complaint is related to a well-known event, previously evaluated and investigated. No further investigation is required as per 21 cfr part 820. 198 complaint files §(b) and (c). Investigation results concluded that the reported event was due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A summary of the investigation was sent to the complainant conveying zimmer biomet conclusions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10919240&pc=MBB
3006946279-2020-00226 | 10/30/2020 | 11/25/2020 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source: foreign: (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10902586&pc=MBB
3006946279-2020-00225 | 10/28/2020 | 11/25/2020 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTICS | - | It has been reported that it was found that the inner sterile packaging was opened during the operation, the doctor use an other product to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10902356&pc=MBB
3006946279-2020-00224 | 06/10/2020 | 11/19/2020 | 162 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE ACETABULAR CUP CEMENTED SIZE 50 | HIP, PROTHESIS | - | Dislocation of advantage liner that was previously cemented in after failure of the previous freedom constrained cup. No other adverse event has been reported for the patient. According to the research and development department, the dislocation could have occurred between the insert and the cup. Therefore, the advantage cup has been investigated in the same complaint but in the decision tree number (b)(4). | (b)(4). List of associated devices: avantage e1 inlay s50/28, reference p0561e50, batch 0001453965. Product pictures have been received. However, they are difficult to interpret as there is some reflections on it. The insert did not seems to be prematurly worn out. No x-ray showing the dislocation has been received. Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to dislocation: 2 complaints (2 products), this one included, have been recorded on avantage cemented cup s50, reference p0463050, from january 02, 2017 to november 10, 2020. 1 complaint (1 product), this one included, has been recorded on avantage cemented cup s50, reference p0463050, batch 0001460174. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10868294&pc=KWA
3006946279-2020-00223 | 09/26/2020 | 11/19/2020 | 54 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation. On the four packages of bone cement, the powder was leaked. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The product analysis cannot be performed as the product was not returned. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Eight complaints have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, over the batch a925ca2503. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10866160&pc=MBB
3006946279-2020-00222 | 11/20/2019 | 11/19/2020 | 365 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation. On the six packages of bone cement, the powder leaked. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The product analysis cannot be performed as the product was not returned. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Eight complaints have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, over the batch a925ca2503. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10866151&pc=MBB
3006946279-2020-00221 | 10/29/2020 | 11/19/2020 | 21 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation. On the two packages of bone cement, the powder leaked. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The product analysis cannot be performed as the product was not returned. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Eight complaints have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, over the batch a925ca2503. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10866140&pc=MBB
3006946279-2020-00220 | 10/22/2020 | 11/19/2020 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened during the operation. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | Cmp (b)(4). Report source, foreign - event occurred in (b)(6). The product analysis cannot be performed as the product was not returned. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Ten complaints have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, over the batch a926ba2506. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10866103&pc=MBB
3006946279-2020-00219 | 10/15/2020 | 11/10/2020 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the bone cement powder was spilled out from the sterile packaging and was found not completely sealed. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The product analysis shows that on the inner pouch, the left sealing and right sealing were opened on 1/6 of the length on each side. The reported event is confirmed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 8 complaints have been recorded on refobacin bone cement r 1x40-3, reference 3003940001-3, over the batch a925ca2503 since ever. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10816305&pc=MBB
3006946279-2020-00218 | 10/29/2020 | 11/06/2020 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 56MM | ORTHOPAEDIC PROTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that during the impaction of the cup, on the pieces on the tip of the impactor broke off. Nothing fell into the patient and they were able to finish the surgery without any delay. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10800261&pc=HWA
3006946279-2020-00217 | 10/05/2020 | 11/04/2020 | 30 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT 1X40G | BONE CEMENT | - | It has been reported that during the surgery, the last bag was found not properly sealed and the powder came out, compromising the sterility of the cement. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10784764&pc=LOD
3006946279-2020-00215 | 10/01/2020 | 10/27/2020 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the inner cement pouch was found not sealed correctly. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10739950&pc=LOD
3006946279-2020-00216 | 10/03/2020 | 10/27/2020 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that during the surgery, the cement come off from the bag which was found not properly sealed when the packaging has been opened. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10739899&pc=LOD
3006946279-2020-00214 | 10/13/2020 | 10/23/2020 | 10 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE INLAY SIZE 44 22,2 | NON-CONSTRAINED POLYETHYLENE ACETABULAR LINER | - | It was reported patient underwent hip arthroplasty on (b)(6) 2012. Subsequently was revised on (b)(6) 2020 due to hip pain. Surgeon had xray, showed movement of the cup and cage. Surgeon operated to remove cup and cage. | (b)(4). List of associated devices: avantage cemented cup s44, reference (b)(4), batch 0000710631. Max ti cage 52mm right m right, reference (b)(4) , batch 31093. 22. 2mm dia cocr modular head standard neck, reference (b)(4) , batch 428170. Ti low profile screw 6. 5x30mm, reference (b)(4) , batch 138500. Report source, foreign - event occurred in (b)(4). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10725659&pc=LPH
3006946279-2020-00213 | 10/13/2020 | 10/23/2020 | 10 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED ACETABULAR CUP SIZE 44 | PROTHESIS, HIP | - | It was reported patient underwent hip arthroplasty on (b)(6) 2012. Subsequently was revised on (b)(6) 2020 due to hip pain. Surgeon had xray, showed movement of the cup and cage. Surgeon operated to remove cup and cage. | (b)(4). List of associated devices: avantage inlay size 44 ø 22,2, reference (b)(4), batch 0000714004. Max ti cage 52mm right m right, reference (b)(4), batch 31093. 22. 2mm dia cocr modular head standard neck, reference (b)(4), batch 428170. Ti low profile screw 6. 5x30mm, reference (b)(4), batch 138500. Foreign report source: (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10725619&pc=JDI
3006946279-2020-00210 | 10/12/2020 | 10/21/2020 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN REVISION 1X40G | BONE CEMENT, ANTIBIOTICS | - | It has been reported that during the surgery, the cement blister was found opened when it was given to the instrumentalist. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event was confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 5 complaints have been recorded over the batch a808ag2705. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10710068&pc=MBB
3006946279-2020-00209 | 10/14/2020 | 10/21/2020 | 7 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN REVISION 1X40G | BONE CEMENT, ANTIBIOTICS | - | It has been reported that during the second surgery. The cement blister was found opened when it was given to the instrumentalist. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event was confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 5 complaints have been recorded over the batch a808ag2705. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found, which would change, alter any conclusions, or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10710053&pc=MBB
3006946279-2020-00212 | - | 10/21/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that patient underwent right total knee arthroplasty (b)(6) 2019. Subsequently, patient has had pain, difficulty walking (limps) and limited range of motion since her procedure. She had allergy testing and was told she is quite possibly allergic to the bone cement. She requested material composition for all devices to provide to her allergist. | (b)(4). The device will not be returned to the manufacturer as cement remains implanted. Therefore it will not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the raw material certificate shows no non conformity or deviation. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10709846&pc=MBB
3006946279-2020-00211 | - | 10/21/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported that patient underwent right total knee arthroplasty (b)(6) 2019. Subsequently, patient has had pain, difficulty walking (limps) and limited range of motion since her procedure. She had allergy testing and was told she is quite possibly allergic to the bone cement. She requested material composition for all devices to provide to her allergist. | (b)(4). The device will not be returned to the manufacturer as cement remains implanted. Therefore it will not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10709820&pc=MBB
3006946279-2020-00208 | 09/28/2020 | 10/20/2020 | 22 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that during operation, the cement powder packet was found not sealed properly, and the powder has leaked inside the packaging. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10703212&pc=MBB
3006946279-2020-00206 | 09/23/2020 | 10/20/2020 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been that during the initial surgery two sterile packs of the cement of polymer powder have leaked. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer yet. Therefore, it could not be analyzed yet. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is pending. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10703205&pc=MBB
3006946279-2020-00205 | 09/22/2020 | 10/20/2020 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the inner cement pouch leaked. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event was confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 4 complaints (included 6 products) have been reported over the batch a848dh1404. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10703202&pc=LOD
3006946279-2020-00207 | 08/04/2020 | 10/16/2020 | 73 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | UPTION COMPLETE RECONSTRUCTION STEM CEMENTLESS / 11 / S3 / LEFT | PROTHESIS, HIP | - | It was reported a breakage of the stem implanted 10 month before ((b)(6) 2019) at the level of the first interlocking pins hole. A revision surgery has been made on (b)(6) 2020 in order to remove the stem and replaced by osteosynthesis by plate. Event Description It was reported a breakage of the stem implanted 10 months before ((b)(6) 2019) at the level of the first interlocking pins hole. A revision surgery has been made on (b)(6) 2020 in order to remove the stem and replaced by osteosynthesis by plate. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported a breakage of the stem implanted 10 month before ((b)(6) 2019) at the level of the first interlocking pins hole. A revision surgery has been made on (b)(6) 2020 in order to remove the stem and replaced by osteosynthesis by plate. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10689483&pc=LPH
3006946279-2020-00204 | - | 10/14/2020 | - | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported that the cement of the optipac has not the correct consistency during surgery. The product has not been used and it has been discarded. There was no patient involvement. | (b)(4). Foreign report source: (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10676255&pc=MBB
3006946279-2020-00203 | 03/09/2020 | 10/13/2020 | 218 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT | - | It has been reported that the inner sterile packaging was found opened during the operation. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 2 complaints have been recorded over the batch a742bl1109. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10675721&pc=MBB
3006946279-2020-00199 | 09/14/2020 | 10/12/2020 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner sterile packaging was found opened before the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that the inner sterile packaging was found opened before the operation. No adverse events have been reported as a result of the malfunction. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The product analysis shows that on the inner pouch, the left supplier sealing is opened on each side of the length and it stops just before the junction with the biomet sealing. The right supplier sealing was also opened on each side of the length. The reported event is confirmed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 8 complaints have been recorded on refobacin bone cement r 1x40-3, reference (b)(4), over the batch a926ba2506. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10662171&pc=MBB
3006946279-2020-00202 | 09/29/2020 | 10/12/2020 | 13 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN PLUS BONE CEMENT-3 | BONE CEMENT, ANTIBIOTIC | - | The patient underwent a primary knee prosthesis with vanguard, where signature molds were used. The patient started with pain shortly after the surgery, which became disabling. The surgeon decided to review the tibial component to correct a possible excessive posterior slope. During the revision surgery, a total detachment of the vanguard tibial plate was detected, without any cement residue at the base. | (b)(4). Concomitant medical products: polished finned tib tray 63mm, reference 141251, batch 2019040413. Vanguard cr ilok fem-lt 60, reference 183024, batch j6371325. Vanguard ant stblzd brg 12x63, reference 189022, batch 376110. Report source, foreign - event occurred in portugal. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10661727&pc=MBB
3006946279-2020-00201 | 02/05/2019 | 10/12/2020 | 615 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40G | BONE CEMENT, ANTIBIOTIC | - | The patient was treated for a periprosthetic fracture in (b)(6) 2016 initially. In (b)(6) 2019 a revision for aseptic loosening of the femur had to be performed. A second revision had to be carried out on (b)(6) 2020 for a new removal of the femoral component because of aseptic loosening. This complaint handle the first revision surgery which occured on (b)(6) 2019. | (b)(4). List of associated devices: oss poly tibial bushing, reference 150476, batch 678220; oss mod tib baseplate 79mm, reference 150424, batch 811420; oss cemented im stem 13mmx90mm, reference 150362, batch 169330; biomet arcom ap patella 34mm, reference 11-150822, batch 801050; oss axle, reference 150480, batch 936720; oss poly femoral bushings 2pk, reference 150477, batch 166910; oss reinforced yoke, reference 150493, batch 190020; oss 7cm segmental femoral rt, reference 150354, batch 171250; oss poly lock pin, reference 150478, batch 879820; oss 3cm ellip diaphyseal seg, reference 150461, batch 627580; oss cemented im stem 13x150, reference 150367, batch 379450; oss poly femoral bushings 2pk, reference 150477, batch 420580. Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10661317&pc=MBB
3006946279-2020-00200 | 08/28/2020 | 10/12/2020 | 45 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 60 REFOBACIN PLUS BONE CEMENT-3 | BONE CEMENT, ANTIBIOTIC | - | The patient was treated for a periprosthetic fracture in (b)(6) 2016 initially. In (b)(6) 2019 a revision for aseptic loosening of the femur had to be performed. A second revision had to be carried out on (b)(6) 2020 for a new removal of the femoral component because of aseptic loosening. This complaint handle the second revision surgery which occured on (b)(6) 2020. | (b)(4). List of associated devices: oss 7cm segmental femoral rt, reference 150354, batch 804700. Oss cemented im stem 15x150, reference 150369, batch 613330. Oss 5cm diaphyseal segment, reference 150465, batch 925300. Oss tibial poly bearing 14mm, reference 150411, batch 560730. Oss poly tibial bushing, reference 150476, batch 315070. Oss poly lock pin, reference 150478, batch 668620. Oss poly femoral bushings 2pk, reference 150477, batch 742240. Oss reinforced yoke, reference 150493, batch 061560. Oss axle, reference 150480, batch 258450. Report source, foreign - event occurred in (b)(6). The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10661219&pc=MBB
3006946279-2020-00198 | 09/10/2020 | 10/06/2020 | 26 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that the inner package leaked. The issue was found before the surgery. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 4 complaints have been recorded on biomet bone cement r 40 -3, reference (b)(4), over the batch a848dh1404. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found, which would change, alter any conclusions, or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10634657&pc=MBB
3006946279-2020-00197 | 09/12/2020 | 10/06/2020 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTICS | - | It has been reported that he inner sterile pouch of the cement was found opened. The event happened during the initial surgery. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Complaints have been recorded on refobacin bone cement r 1x40-3; reference (b)(4), over the batch a925ca2505. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A final response will be send to the customer. If any further information is found which would change or alter any conclusions, or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10633438&pc=MBB
3006946279-2020-00196 | 11/06/2019 | 10/05/2020 | 334 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE -4 | PROTHESIS, HIP | - | It was reported patient underwent a right hip arthroplasty on (b)(6) 2019. The day after, she had dysesthesia in right thigh. | (b)(4). Report source, foreign - event occurred in (b)(6). List of associated devices: g7 bispherical shell 44a, reference 110017328, batch 334515; g7 neutral e1 liner 28mm a, reference 010000835, batch 3411930; biomet biolox delta ceramic, reference 6500831, batch 2018112087. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10625478&pc=KWA
3006946279-2020-00195 | 08/19/2020 | 10/05/2020 | 47 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 | BONE CEMENT, ANTIBIOTICS | - | It was reported that a patient had an initial unilateral total left knee arthroplasty on (b)(6) 2020. Subsequently, the patient fell approximately three weeks post-op. The patient began experiencing difficulty ambulating, decrease in adl¿s, pain, swelling, radiolucency, a periprosthetic fracture, loosening, subsidence, and poor bone quality. The patient was revised to a full total left knee arthroplasty on (b)(6) 2020. | (b)(4). List of associated devices: partial tibial cemented size d left medial, reference 42538000401, batch unknown, handle in (b)(4); partial femur cemented size 3 left medial, reference 42558000301, batch 64465213, handle in (b)(4); partial articular surface left medial size d 10 mm thickness, reference 42518200410, batch 63755573, handle in (b)(4). The device will not be returned to the manufacturer. Therefore it will not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10625324&pc=MBB
3006946279-2020-00194 | - | 09/30/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | UPTION COMPLETE FEMORAL STEM | COAED HIP FEMUR PROTHESIS, MODULAR | - | A fracture of the greater trochanter was osteosynthesized by hook plate at 15 days postoperatively following a fall during rehabilitation. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description A fracture of the greater trochanter was osteosynthesized by hook plate at 15 days postoperatively following a fall during rehabilitation. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. Manufacturer Narrative (b)(4). No further information provided (x-rays, surgical report, photographs, lab test), therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The review of the device manufacturing quality can't be performed as the product reference and batch number was not communicated. A complaint extract was done regarding bone fracture: 1 complaint (1 product), this one included, has been recorded on uption complete femoral stem, from (b)(6) 2017 to (b)(6) 2020. - the complaint history, regarding the reference number, can't be performed as it was not communicated. The complaint history, regarding the batch number, can't be performed as it was not communicated according to available data, the most probable root cause of the bone fracture is the patient fall during the rehabilitation protocole. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10604929&pc=KWY
3006946279-2020-00191 | 06/24/2013 | 09/30/2020 | 2655 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | UPTION COMPLETE FEMORAL STEM REVISION | COATED HIP FEMUR PROTHESIS, MODULAR | - | It was reported that an uption complete stem was implanted for aseptic loosening of the femoral component without changing the cup (ceramic/ceramic couple). In front of a table of early recurrent dislocations, a measurement of the version of the implants was performed (15° anteversion of the uption complete stem and a 0° anteversion cup). It was decided to use a double mobility cup. Intraoperatively, a mobility of the femoral implant motivated a change of the implant for a long cemented stem. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. Event Description It was reported that an uption complete stem was implanted for aseptic loosening of the femoral component without changing the cup (ceramic/ceramic couple). Due to early recurrent dislocations, a measurement of the version of the implants was performed (15° anteversion of the uption complete stem and a 0° anteversion cup). It was decided to use a double mobility cup. Intraoperatively, a mobility of the femoral implant motivated a change of the implant for a long cemented stem. | (b)(4). Report source, foreign - event occurred in (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported that an uption complete stem was implanted for aseptic loosening of the femoral component without changing the cup (ceramic/ceramic couple). In front of a table of early recurrent dislocations, a measurement of the version of the implants was performed (15° anteversion of the uption complete stem and a 0° anteversion cup). It was decided to use a double mobility cup. Intraoperatively, a mobility of the femoral implant motivated a change of the implant for a long cemented stem. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10604781&pc=JDI
3006946279-2020-00193 | - | 09/30/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UPTION COMPLETE FEMORAL STEM | COATED HIP FEMUR PROTHESIS, MODULAR | - | Pseudarthrosis of the femoral flap necessitated a decortication-graft with hook plate osteosynthesis without changing the stem at 15 months from the initial procedure. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. | (b)(4). Report source, foreign - event occurred in (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10604860&pc=LZO
3006946279-2020-00192 | - | 09/30/2020 | - | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | UPTION COMPLETE FEMORAL STEM | COATED HIP FEMUR PROTHESIS, MODULAR | - | It was reported that 3 patients suffered from deep infection. All three cases were resolved by lavage-synovectomy and change of modular implants without uption complete stem removal at 1 month for two patients and at 3 months for the last case where it was an acute infection with late onset. The current complaint comes from an article that has not been published yet and transmitted to zimmer biomet clinical department. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10604824&pc=LZO
3006946279-2020-00188 | 09/15/2020 | 09/24/2020 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that a hole was found in the sterile packaging and the power leaked. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that a hole was found in the sterile packaging and the power leaked. No adverse events have been reported as a result of the malfunction. Manufacturer Narrative (b)(4). The 2 products analysis shows that on the two inner pouches, the right sealing is opened on 1/6 of the length. The reported event has been confirmed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Seven complaints have been recorded on biomet bone cement r 40 -3, reference (b)(4) over the batch a848dh1404. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10578608&pc=MBB
3006946279-2020-00186 | 09/09/2020 | 09/24/2020 | 15 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that during the surgery, the inner sterile pouch of the cement was found opened and all the powder was dropped down to the floor. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10578937&pc=MBB
3006946279-2020-00187 | 09/10/2020 | 09/24/2020 | 14 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that the packaging was found not sealed correctly, the powder spread into the protective packaging. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10578603&pc=MBB
3006946279-2020-00185 | 09/08/2020 | 09/24/2020 | 16 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET BONE CEMENT R 40 -3 | BONE CEMENT | - | It has been reported that during the surgery, the packaging was not sealed correctly, the powder spread into the protective packaging. No adverse events have been reported as a result of the malfunction. | Cmp (b)(4). Report source, foreign - event occurred in (b)(6). The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No |  | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10578592&pc=MBB
3006946279-2020-00190 | 06/21/2020 | 09/23/2020 | 94 | Death | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | ORTHOPAEDIC CEMENT, MEDICATED | - | It was reported the patient death after cementing a femoral stem. To date, the cement brand is not confirmed. The current report is submitted as a conservative approach in waiting for the confirmation that a zimmer biomet cement was involved. | Cmp (b)(4). Concomitant medical products and therapy dates : avenir stem. Report source, foreign - event occurred in (b)(6). The device manufacturing quality record could not been reviewed as the reference and the lot number was not communicated. To date, the fact that the cement is a zimmer biomet cement is not confirmed. Investigation is ongoing. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported the patient death after cementing a femoral stem. To date, the cement brand is not confirmed. The current report is submitted as a conservative approach in waiting for the confirmation that a zimmer biomet cement was involved. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. D11 : avenir stem lateral, cemented 9 taper 12/14 - reference : (b)(4), - lot : 2836784. X-rays has been provided and shows a dislocated medical fracture of the neck of the left femur. There are no details on the x-rays to show the shock due to the cement. The surgical notes has been reviewed. Product was not returned. The reported event was confirmed by surgical notes review. The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. No other complaint on patient death was recorded on the batch since ever, and no other complaint on patient death was recorded on the reference since ever. Reserve sample from the same lot was tested under standardized conditions. The product did not show any unusual behavior during mixing, handling or setting. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10572077&pc=JDZ
3006946279-2020-00189 | 11/16/2018 | 09/21/2020 | 675 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | COBALT CHROME FEMORAL HEAD 542 32 / +4 / 12-14 / COL LONG | METALLIC FEMORAL HEAD PROTHESIS | - | It has been reported by a patient that he underwent a total hip arthroplasty in (b)(6) 2018 ((b)(6)) at the (b)(6). Subsequently, patient felt pain and has reported that a dislocation took place on (b)(6) 2018 where he was treated at the same hospital. A further dislocation took place and the patient was revised on (b)(6) 2018, again at the same hospital. No further information has been provided. | (b)(4). Report source, foreign - event occurred in (b)(6). Associated devices: ringloc-x e1 liner 10deg 52/32mm, reference ep-063252, batch 6002209; exceed ringloc-x shell porous coated d52mm, reference 131352, batch 3972375; unknown screw. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10559480&pc=LPH
3006946279-2020-00184 | 06/23/2020 | 09/18/2020 | 87 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 40 REFOBACIN BONE CEMENT R | BONE CEMENT, ANTIBIOTICS | - | It was reported a study patient underwent a right total knee arthroplasty on (b)(6) 2018. Subsequently, tibial baseplate loosening was noted on (b)(6) 2020 and the patient was revised on (b)(6) 2020. | (b)(4). List of associated devices: femoral component option size e right, reference (b)(4), batch 63957962; articular surface regular constraint, reference (b)(4), batch 63952346; poly patella standard size 35 mm diameter 9. 0 mm thickness cemented, reference (b)(4), batch 63984274; stemmed nonaugmentable tibial component option cr/ps/lps size 5 for cemented use only, reference (b)(4), batch 63905965; optipac 40 refobacin bone cement r, reference (b)(4), batch a720a04010. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10552488&pc=MBB
3006946279-2020-00180 | 09/08/2020 | 09/17/2020 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). (b)(6). The device will not be returned to the manufacturer. Therefore, it will not be analyzed. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event is confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 8 complaints have been recorded on refobacin bone cement r 1x40-3, reference (b)(4), batch a926ba2506. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545305&pc=MBB
3006946279-2020-00176 | 09/02/2020 | 09/17/2020 | 15 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device was not returned to the manufacturer. Therefore, it could not be analyzed. The event was confirmed through pictures. The device manufacturing quality record indicates that the released product met all requirements to perform as intended. 4 complaints have been recorded on refobacin bone cement r 1x40-3, reference (b)(4), batch a927ba1601. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. A customer letter will be send regarding the reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545290&pc=MBB
3006946279-2020-00162 | 08/12/2020 | 09/17/2020 | 36 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40 US | BONE CEMENT, ANTIBIOTIC | - | It was reported the patient underwent a second knee revision surgery due to loosening of the tibial and femoral component. Subsequently, the patient required a third knee revision surgery due to pain, instability, loosening of the tibial and femoral component. The patient fell, slipped shortly after the second knee revision. No additional patient consequences were reported. | (b)(4). Event occurred in united states. Devices investigated in (b)(4) : ref : (b)(4) name : prc agmt block dist sz e5mm lot : 61822489. Ref : (b)(4) name : stemmed tibial component lot : 64394387. Ref : (b)(4) name : stem extension sharp fluted 16mm dia x 75mmlength lot : 63580188. Ref : (b)(4) name :offset stem impl18mmx200mm lot : 62911816. Ref : (b)(4) name : lcck fem implant sz e-r lot : 64434949. Ref : (b)(4) name : lcck art surf ef 5-6/grn 20 lot : 63857712. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545757&pc=MBB
3006946279-2020-00183 | 09/09/2020 | 09/17/2020 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source: (b)(6). The device will not be returned to the manufacturer. Therefore, it will not be analyzed. If any further information is found, which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545310&pc=MBB
3006946279-2020-00182 | 09/09/2020 | 09/17/2020 | 8 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation. The doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545307&pc=MBB
3006946279-2020-00181 | 09/08/2020 | 09/17/2020 | 9 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. Event Description It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Foreign report source: (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that the inner sterile packaging was found opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545306&pc=MBB
3006946279-2020-00178 | 08/24/2020 | 09/17/2020 | 24 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was found that the inner sterile packaging was opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. Investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545296&pc=MBB
3006946279-2020-00177 | 08/20/2020 | 09/17/2020 | 28 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It was reported that the inner sterile packaging was opened during the operation, the doctor used a spare bone cement to complete the operation. No harm was done to the patient. | (b)(4). Report source, foreign - event occurred in (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. The investigation is in progress. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10545294&pc=MBB
3006946279-2020-00179 | 03/29/2016 | 09/14/2020 | 1630 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FMRL STEM SIZE 0 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was noted a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2016 due to unknown reason where stem was removed. This complaint has been received from siris report: "siris outlier report 2019 gts + g7 zimmer biomet, for uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary oa, status: confirmed outlier". | (b)(4). Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to unknown reason: - 5 complaints (5 products), this one included, have been recorded on gts standard femoral stem size 0, reference (b)(4) from (b)(6) 2017 to (b)(6) 2020. - 2 complaints (2 products), this one included, have been recorded on gts standard femoral stem size 0, reference (b)(4), batch 000740818. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10527820&pc=MEH
3006946279-2020-00175 | 02/22/2013 | 09/13/2020 | 2760 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FMRL STEM SIZE 0 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was noted a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2013. Due to unknown reason where stem was removed. This complaint has been received from siris report: "siris outlier report 2019 gts + g7 zimmer biomet, for uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary oa, status: confirmed outlier". | (b)(4). Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to unknown reason: - 5 complaints (5 products), this one included, have been recorded on gts standard femoral stem size 0, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 2 complaints (2 products), this one included, have been recorded on gts standard femoral stem size 0, reference (b)(4), batch 000740818. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526148&pc=LPH
3006946279-2020-00174 | 08/21/2019 | 09/13/2020 | 389 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FMRL STEM SIZE 0 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was noted a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2019 due to unknown reason where stem was removed. This complaint has been received from siris report: "siris outlier report 2019 gts + g7 zimmer biomet, for uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary oa, status: confirmed outlier". | (b)(6). Report source, foreign - event occurred in (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to unknown reason: 5 complaints (5 products), this one included, have been recorded on gts standard femoral stem size 0, reference ps129g00, from 1 january 2017 to 02 september 2020. 1 complaints (1 products), this one included, have been recorded on gts standard femoral stem size 0, reference ps129g00, batch 0000751774. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526147&pc=MEH
3006946279-2020-00173 | 01/10/2020 | 09/13/2020 | 247 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE 0 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was noted a patient underwent a primary procedure on (b)(6) 2016. Subsequently, that patient underwent a revision procedure on (b)(6) 2020 due to unknown reason where stem and cup were removed. This complaint has been received from siris report: "siris outlier report 2019 gts + g7 zimmer biomet, for uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary oa, status: confirmed outlier. ". | (b)(4) report source, foreign country event occurred in (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 5 complaints (5 products), this one included, have been recorded on gts standard femoral stem size 0, reference ps129g00, from 1 january 2017 to 2 september 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, batch 0001069765. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526146&pc=JDI
3006946279-2020-00169 | 11/06/2019 | 09/13/2020 | 312 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE -1 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2019 due to unknown reason where stem and cup were removed. No additional patient consequences were reported. This complaint has been received from siris report: "siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Concomitant medical devices ¿ g7 bispherical shell 52e, reference (b)(4), batch 6144757. Report source, foreign - event occurred in (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 3 complaints (3 products), this one included, have been recorded on gts standard femoral stem size -1, reference (b)(4), from 1 january 2017 to 3 september 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size -1, reference (b)(4), batch 0001067802. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526145&pc=LPH
3006946279-2020-00168 | 04/26/2018 | 09/13/2020 | 871 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE -4 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2018 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: - 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size -4, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size -4, reference (b)(4), batch 0000753248. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526144&pc=JDI
3006946279-2020-00167 | 05/08/2015 | 09/13/2020 | 1955 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS VARISED FEMORAL STEM SIZE -1 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2015 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: - 1 complaint (1 product), this one included, has been recorded on gts varised femoral stem size -5, reference pv129gm5, from (b)(6) 2017 to (b)(6) 2020. - 1 complaint (1 product), this one included, has been recorded on gts varised femoral stem size -5, reference pv129gm5, batch 0000615478. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526141&pc=LZO
3006946279-2020-00166 | 02/18/2014 | 09/13/2020 | 2399 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS VARISED FEMORAL STEM SIZE +1 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2014 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: - 1 complaint (1 product), this one included, has been recorded on gts varised femoral stem size +1, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 1 complaint (1 product), this one included, has been recorded on gts varised femoral stem size +1, reference (b)(4), batch 0000763707. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00165 | 03/07/2014 | 09/13/2020 | 2382 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS VARISED FEMORAL STEM SIZE -2 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2014 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: - 2 complaints (2 products), this one included, have been recorded on gts varised femoral stem size -2, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 2 complaints (2 products), this one included, have been recorded on gts varised femoral stem size -2, reference (b)(4), batch 0000625551. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526139&pc=LPH
3006946279-2020-00164 | 03/31/2017 | 09/13/2020 | 1262 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE -3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on (b)(6) 2016. Subsequently, that patient underwent a revision procedure on (b)(6) 2018 due to femoral loosening, infection and pain where all products were removed and component was reimplanted (after spacer or girdlestone). No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Concomitant medical products: g7 bispherical shell 54f, reference 110017333, batch 3333157, unknown liner and unknown head. Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. A complaint extract was done regarding revision, pain, infection and loosening: - 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size -3, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size -3, reference (b)(4), batch 00j3526477. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526138&pc=LPH
3006946279-2020-00161 | 08/22/2018 | 09/13/2020 | 753 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2018 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 6 complaints (6 products), this one included, have been recorded on gts standard femoral stem size +3, reference ps129gp3, from (b)(6) 2017 to (b)(6) 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference ps129gp3, batch 0001279969. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526137&pc=LPH
3006946279-2020-00160 | 01/24/2019 | 09/13/2020 | 598 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2019 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 6 complaints (6 products), this one included, have been recorded on gts standard femoral stem size +3, reference ps129gp3, from (b)(6) 2017 to (b)(6) 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference ps129gp3, batch 0001091615. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526136&pc=JDI
3006946279-2020-00159 | 05/22/2015 | 09/13/2020 | 1941 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2015 due to unknown reason where stem was removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 6 complaints (6 products), this one included, have been recorded on gts standard femoral stem size +3, reference ps129gp3, from (b)(6) 2017 to (b)(6) 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference ps129gp3, batch 00j3348891. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526135&pc=LPH
3006946279-2020-00172 | 09/02/2020 | 09/09/2020 | 7 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | GTS STANDARD FEMORAL STEM SIZE 0 | PROTHESIS, HIP | - | It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being filled to relay additional information. Received x-ray shows that the whole device seems correctly implanted. The review of the medical record showed that the surgeon respected the surgical technique as he perform trial before the final devices implantation. The product was returned and lab analysis was performed. The product analysis shows that the gts stem and the biolox head were stuck together after impaction. In order to unlock the products, it was necessary to cut a part of the stem. Once the products were unblocked, they were inspected and they seemed to be conform. The dimensional analysis of the component could not be performed as the stem was cut to unblock the head. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The instructions for use have been reviewed and it was found that the biolox head is compatible with the gts stem used. A complaint extract was done regarding revision due to implants stuck to each other: 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, from january 01, 2017 to november 24, 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, batch j6549101. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device has been returned and is currently being analyzed. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. | (b)(4). Report source, foreign - event occurred in (b)(6). Concomitant medical devices - list of associated devices: delta ceramic femoral head 036/0mm 12/14, reference 650-0837, batch 2020031367; g7 hi-wall e1 liner 36mm f, reference 010000936, batch 6804773; g7 acetabular shell 54f, reference 010000664, batch 6721122. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being filled to relay additional information. Received x-ray shows that the whole device seems correctly implanted. The review of the medical record showed that the surgeon respected the surgical technique as he perform trial before the final devices implantation. The product was returned and lab analysis was performed. The product analysis shows that the gts stem and the biolox head were stuck together after impaction. In order to unlock the products, it was necessary to cut a part of the stem. Once the products were unblocked, they were inspected and they seemed to be conform. The dimensional analysis of the component could not be performed as the stem was cut to unblock the head. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The instructions for use have been reviewed and it was found that the biolox head is compatible with the gts stem used. A complaint extract was done regarding revision due to implants stuck to each other: 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, from january 01, 2017 to november 24, 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, batch j6549101. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device has been returned and is currently being analyzed. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00163 | 03/20/2018 | 09/12/2020 | 907 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported the patient underwent a primary procedure on (b)(6) 2017. Subsequently, that patient underwent a revision procedure on (b)(6) 2018 due to acetabular loosening and pain where the acetabular components and femoral components were removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). G7 bispherical shell 52e, reference 110017332, batch 6034151, unknown liner and unknown head. (b)(6). The product analysis can't be performed as the product was not returned the device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision due to pain: 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference ps129gp3, from (b)(6) 2017 to (b)(6) 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference ps129gp3, batch 0001230610. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00157 | 05/05/2017 | 09/13/2020 | 1227 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +3 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on (b)(6) 2017. Subsequently, that patient underwent a revision procedure on (b)(6) 2017 due to unknown reason where stem and cup were removed. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Concomitant medical products: g7 bispherical shell 50d, reference (b)(4), batch 3979320. Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: - 6 complaints (6 products), this one included, have been recorded on gts standard femoral stem size +3, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. - 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +3, reference (b)(4), batch 0001190394. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00156 | 12/05/2019 | 09/11/2020 | 281 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | GTS STANDARD FEMORAL STEM SIZE +4 | PRESS-FIT HIP FEMUR PROSTHESIS, MODULAR | - | It was reported that a patient underwent a primary procedure on unknown date. Subsequently, that patient underwent a revision procedure on (b)(6) 2019 due to unknown reason where stem was removed. No additional patient consequences were reported. No additional patient consequences were reported. This complaint has been received from siris report: " siris outlier report 2019, gts + g7, zimmer biomet, uncemented stem-cup combination used in primary total hip arthroplasty, diagnosis primary osteoarthritis, status: confirmed outlier". | (b)(4). Foreign report source: (b)(6). The product analysis can't be performed as the product was not returned. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding revision: 2 complaints (2 products), this one included, have been recorded on gts standard femoral stem size +4, reference (b)(4), from (b)(6) 2017 to (b)(6) 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size +4, reference (b)(4), batch j6195861. According to available data, root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00171 | 08/05/2020 | 09/09/2020 | 35 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | BIOMET BONE CEMENT R 1X40 | BONE CEMENT | - | It was reported that the bone cement hardened faster than usual. | (b)(4). Report source, foreign - event occurred in (b)(4). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported that the bone cement hardened faster than usual. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The product analysis can't be performed as the product was not returned. A retain sample of batch 947cal0605 has been tested in the laboratory under standardized conditions. No unusual behavior during mixing, handling or setting. The reported behavior of the cement cannot be reproduced. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. A complaint extract was done regarding too short setting time: 2 complaints (3 products), this one included, have been recorded on biomet bone cment r 1x40, reference 110035372, batch 947cal0605. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00172 | 09/02/2020 | 09/09/2020 | 7 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | GTS STANDARD FEMORAL STEM SIZE 0 | PROTHESIS, HIP | - | It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being filled to relay additional information. Received x-ray shows that the whole device seems correctly implanted. The review of the medical record showed that the surgeon respected the surgical technique as he perform trial before the final devices implantation. The product was returned and lab analysis was performed. The product analysis shows that the gts stem and the biolox head were stuck together after impaction. In order to unlock the products, it was necessary to cut a part of the stem. Once the products were unblocked, they were inspected and they seemed to be conform. The dimensional analysis of the component could not be performed as the stem was cut to unblock the head. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The instructions for use have been reviewed and it was found that the biolox head is compatible with the gts stem used. A complaint extract was done regarding revision due to implants stuck to each other: 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, from january 01, 2017 to november 24, 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, batch j6549101. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device has been returned and is currently being analyzed. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. | (b)(4). Report source, foreign - event occurred in (b)(6). Concomitant medical devices - list of associated devices: delta ceramic femoral head 036/0mm 12/14, reference 650-0837, batch 2020031367; g7 hi-wall e1 liner 36mm f, reference 010000936, batch 6804773; g7 acetabular shell 54f, reference 010000664, batch 6721122. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being filled to relay additional information. Received x-ray shows that the whole device seems correctly implanted. The review of the medical record showed that the surgeon respected the surgical technique as he perform trial before the final devices implantation. The product was returned and lab analysis was performed. The product analysis shows that the gts stem and the biolox head were stuck together after impaction. In order to unlock the products, it was necessary to cut a part of the stem. Once the products were unblocked, they were inspected and they seemed to be conform. The dimensional analysis of the component could not be performed as the stem was cut to unblock the head. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The instructions for use have been reviewed and it was found that the biolox head is compatible with the gts stem used. A complaint extract was done regarding revision due to implants stuck to each other: 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, from january 01, 2017 to november 24, 2020. 1 complaint (1 product), this one included, has been recorded on gts standard femoral stem size 0, reference ps129g00, batch j6549101. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. The device has been returned and is currently being analyzed. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It has been reported that it was implanted gts stem and biolox head during an initial surgery. The patient was saturated, medicated, and placed in the supine position. The surgeon decides to reopen the patient and change the head because it does not suit/fit with the length of the limbs. He tried to remove the head but the whole stem comes out. It was not possible to separate the head and the stem. Biolox head and gts stem get stuck together. So they explanted both products. The procedure was finished using another head and stem. The surgery was extended by 15 minutes, no patient adverse consequences have been reported. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00170 | 06/23/2020 | 09/08/2020 | 77 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | OPTIPAC-S 40 REFOB BN CMT R | BONE CEMENT, ANTIBIOTICS | - | It was reported a study patient underwent a right total knee arthroplasty on (b)(6) 2018. Subsequently, tibial baseplate loosening was noted on (b)(6) 2020 and the pt is scheduled for reoperation on (b)(6) 2020. | (b)(4). List of associated devices: femoral component option size e right, reference (b)(4), batch 63957962. Articular surface regular constraint, reference (b)(4), batch 63952346. Poly patella standard size 35 mm diameter 9. 0 mm thickness cemented, reference (b)(4), batch 63984274. Stemmed nonaugmentable tibial component option cr/ps/lps size 5 for cemented use only, reference (b)(4), batch 63905965. Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported a study patient underwent a right total knee arthroplasty on (b)(6) 2018. Subsequently, tibial baseplate loosening was noted on (b)(6) 2020 and the pt is scheduled for reoperation on (b)(6) 2020. Manufacturer Narrative (b)(4). This follow-up report is being submitted to relay additional information. D11 - list of associated devices: femoral component option size e right, reference (b)(4), batch 63957962; articular surface regular constraint, reference 90597004010, batch 63952346; poly patella standard size 35 mm diameter 9. 0 mm thickness cemented, reference (b)(4), batch 63984274; stemmed nonaugmentable tibial component option cr/ps/lps size 5 for cemented use only, reference (b)(4), batch 63905965; optipac 40 refobacin bone cement r, reference (b)(4), batch a744b05300. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00155 | 08/19/2020 | 08/21/2020 | 2 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | VARIZED FEMORAL STEM EXCEPTION -CEMENTLESS- 542- 12/14 - LEFT S7 | PROTHESIS, HIP | - | It was reported a fracture of the hip prosthesis stem in the area of the prosthesis neck below the cone during the following event: the patient had been carrying a watering can in the garden and had experienced acute pain in the area of the right hip and then collapsed. Background: condition after first implantation of the cementless hip prosthesis on the right side on (b)(6) 2019 at the (b)(6). Condition after revision of the right hip joint on (b)(6) 2019 with head/adapter replacement after double dislocation in the (b)(6). Manufacturer Narrative (b)(4). This follow-up report is being submitted to correct information. No x-ray or pictures received and the product was not returned. Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The review of the device manufacturing record of the batch 0001208386 did not show any non-conformity during manufacturing that could be related to the event. A complaint extract was done regarding revision due to implant fracture and pain: 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, from january 01, 2017 to october 28, 2020. 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, batch 0001208386. The instruction for use have been reviewed and shows that the stem has a maximal load of 110 kg. It can be noticed that the stem fracture occurred when the patient carried a watering can in the garden, thus increasing the stem load. An internal non conformity nc 2019-03/19 has been found which could be linked to the event described in the complaint. Indeed, this non conformity relates to a non-conforming cone diameter after polishing at the subcontractor and involving potentially all batches manufactured before july 2019 and polished by this subcontractor. The patient risk related to this non-conformity was assessed through hhe-2019- 00105 with a healthcare professional and it was concluded that even if the neck diameter is under zimmer biomet specifications, it is capable of withstanding the load required in iso standard. Moreover, it was stated that the patient risk was very low, therefore, it was decided not to do a field action. This information can be confirmed by the occurrence rate of this type of issue: one exception stem, all references included, 2 complaints have been recorded since december 2007. During the same period of time, 127 885 exception stem, all references included have been sold. Therefore, the occurrence rate is 0. 002% which correspond to the occurrence rate provide in the risk management file. It can be also noticed that this complaint is a one-off on batch 0001208386. Finally, a capa (ca-05147) has been initiated in order to implement actions to avoid the recurrence of this type of issue. Indeed, since summer 2019, a dimensional inspection after polishing at the supplier has been set up which allow to detect non-conforming products. According to available data, root cause of the event was unable to be determined. Indeed, it was impossible to determine if the fracture of the stem was due to the internal non-conformity or to the patient weight which could has been exceed the maximal load of 110 kg mentioned in instructions for use. Without the patient information and the product return, the exact root cause of the incident cannot be determined. The complaint is closed but could be reopened if new information is received later. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It was reported a fracture of the hip prosthesis stem in the area of the prosthesis neck below the cone during the following event: the patient had been carrying a watering can in the garden and had experienced acute pain in the area of the right hip and then collapsed. Background: condition after first implantation of the cementless hip prosthesis on the right side on (b)(6) 2019 at the (b)(6) condition after revision of the right hip joint on (b)(6) 2019 with head/adapter replacement after double dislocation in the (b)(6). | (b)(4). Report source, foreign - event occurred in (b)(6). No non conformity or deviation was observed which could be linked to the event described in the complaint the review of the raw material certificate shows no non-conformity or deviation. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Event Description It was reported a fracture of the hip prosthesis stem in the area of the prosthesis neck below the cone during the following event: the patient had been carrying a watering can in the garden and had experienced acute pain in the area of the right hip and then collapsed. Background: condition after first implantation of the cementless hip prosthesis on the right side on (b)(6) 2019 at the (b)(6). Condition after revision of the right hip joint on (b)(6) 2019 with head/adapter replacement after double dislocation in the (b)(6). Manufacturer Narrative (b)(4). This follow-up report is being submitted to correct information. No x-ray or pictures received and the product was not returned. Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned. The review of the device manufacturing record of the batch 0001208386 did not show any non-conformity during manufacturing that could be related to the event. A complaint extract was done regarding revision due to implant fracture and pain: 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, from january 01, 2017 to october 28, 2020. 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, batch 0001208386. The instruction for use have been reviewed and shows that the stem has a maximal load of 110 kg. It can be noticed that the stem fracture occurred when the patient carried a watering can in the garden, thus increasing the stem load. An internal non conformity nc 2019-03/19 has been found which could be linked to the event described in the complaint. Indeed, this non conformity relates to a non-conforming cone diameter after polishing at the subcontractor and involving potentially all batches manufactured before july 2019 and polished by this subcontractor. The patient risk related to this non-conformity was assessed through hhe-2019- 00105 with a healthcare professional and it was concluded that even if the neck diameter is under zimmer biomet specifications, it is capable of withstanding the load required in iso standard. Moreover, it was stated that the patient risk was very low, therefore, it was decided not to do a field action. This information can be confirmed by the occurrence rate of this type of issue: one exception stem, all references included, 2 complaints have been recorded since december 2007. During the same period of time, 127 885 exception stem, all references included have been sold. Therefore, the occurrence rate is 0. 002% which correspond to the occurrence rate provide in the risk management file. It can be also noticed that this complaint is a one-off on batch 0001208386. Finally, a capa (ca-05147) has been initiated in order to implement actions to avoid the recurrence of this type of issue. Indeed, since summer 2019, a dimensional inspection after polishing at the supplier has been set up which allow to detect non-conforming products. According to available data, root cause of the event was unable to be determined. Indeed, it was impossible to determine if the fracture of the stem was due to the internal non-conformity or to the patient weight which could has been exceed the maximal load of 110 kg mentioned in instructions for use. Without the patient information and the product return, the exact root cause of the incident cannot be determined. The complaint is closed but could be reopened if new information is received later. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It was reported a fracture of the hip prosthesis stem in the area of the prosthesis neck below the cone during the following event: the patient had been carrying a watering can in the garden and had experienced acute pain in the area of the right hip and then collapsed. Background: condition after first implantation of the cementless hip prosthesis on the right side on (b)(6) 2019 at the (b)(6) condition after revision of the right hip joint on (b)(6) 2019 with head/adapter replacement after double dislocation in the (b)(6). Manufacturer Narrative (b)(4). No x-ray or pictures received and the product was not returned. Therefore, the reported event could not be confirmed. The product analysis can't be performed as the product was not returned the review of the device manufacturing record showed that nc 2019-03/19 on stem neck diameter out of specifications could be linked to the issue describe in the complaint. The instruction for use (f209_ifu exception_index 01_2018-10) have been reviewed and shows that the the stem has a maximal load of 110 kg. A complaint extract was done regarding revision due to implant fracture and pain: 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, from january 01, 2017 to october 28, 2020. 1 complaint (1 product), this one included, has been recorded on exception varized stem right size 7, reference pv126y07, batch 0001208386. Investigation results concluded that the most probable root cause of the event is a supplier issue: stem neck diameter out of specifications due to a lack of information submitted to the supplier (no diameter specifications indicated in the functional specifications) and a lack of incoming inspection at zimmer biomet valence. A capa (ca-05147) has been initiated in order to implement actions to avoid the recurrence of this type of issue. Indeed, since summer 2019, a dimensional inspection after polishing at the supplier has been set up which allow to detect non-conforming products. An hhed (hhe-2019-00105) has been previously initiated with no field action decision. Indeed, a mechanic forces test has been performed and showed that even if the neck diameter is under specifications, it is capable of withstanding the load required in iso standard. Moreover, it can be noticed that the stem has a maximal load of 110 kg which could have been exceeded by the patient when she carried the watering can. Finally, it can be noticed that this complaint is a one-off on batch 0001208386. According to all of these element, it was decided not to do further actions. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00154 | 08/04/2020 | 08/21/2020 | 17 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | OPTIPAC 80 REFOBACIN BONE CEMENT R-3 | BONE CEMENT, ANTIBIOTICS | - | It was reported that the cement did not mix well. The paste was watery after mixing. No known adverse event has been reported for now. | (b)(4). Report source, foreign - event occurred in (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00153 | 04/09/2020 | 08/17/2020 | 130 | Injury | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE CEMENTED ACETABULAR CUP SIZE 48 | PROTHESIS, HIP | - | Removal of avantage e1 insert, wash out and debridement and reinsertion of new insert. The revision of the insert was due to an unspecified infection. The original avantage shell and liner where implanted with a stryker exeter stem and a stryker 22 ball head with heraeus bone cement for the cup & the stem. This report handle the investigation made on the cup. | (b)(4). List of products implanted: the original avantage shell and liner where implanted with a stryker. Exeter stem and a stryker 22 ball head with heraeus bone cement for the cup & the stem. The product details for these items are unavailable. The original avantage shell that was not revised ref p0463048, lot 0001279524. This was supported by tm augment ref 00-4894-050-10, lot 64290950. Bone screw x 2 ref 00-6624-065-30, lot 64088045, ref 00-6624-065-40, lot 64434233. Foreign report source: (b)(6). No further information provided (x-rays, surgical report, photographs, lab test), no product returned, therefore, the reported event could not be confirmed. The device was not returned to the manufacturer. Therefore it could not be analyzed. The device manufacturing quality record indicate that the released product met all requirements to perform as intended. The review of the sterilization certificate indicates that the products were sterilized according to the specification. Within 3 years, no simir complaint has been registered regarding a revision due to infection on avantage cemented acetabular cup size 48, reference p0463048, batch 0001279524. The instructions for use have been reviewed and it was noticed that the devices used in association with the inlay are not compatible, especially the stem and the head that are not manufactured by zimmer biomet. Therefore, the compatibility of the whole device is not guaranteed. According to available data, root cause of the event was unable to be determined. However, there is no evidence that the event is related to the product. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00152 | 07/10/2020 | 08/12/2020 | 33 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | - | REFOBACIN BONE CEMENT R 1X40-3 | BONE CEMENT, ANTIBIOTIC | - | It has been reported that the inner sterile packaging was found already opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. Event Description It has been reported that the inner sterile packaging was found already opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | (b)(4). Report source, foreign - event occurred in (b)(6). The device will not be returned to the manufacturer. Therefore it will not be analyzed. Investigation is pending. According to available data, the most probable root cause is due to packaging issue (sealing process). Corrective action has been initiated to address reported issue. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. Event Description It has been reported that the inner sterile packaging was found already opened during the operation, the doctor used a spare bone cement to complete the operation. No adverse events have been reported as a result of the malfunction. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH
3006946279-2020-00136 | 07/14/2020 | 08/10/2020 | 27 | malfunction | Manufacturer | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | BIOMET FRANCE S.A.R.L. plateau de lautagne bp75 valence cedex 26903 FR 26903 | AVANTAGE IMPACTOR TIP 46MM | ORTHOPAEDIC PROTHESIS IMPLANTATION INSTRUMENT, REUSABLE | - | It was reported that upon impaction of the avantage cemented cup into the acetabulum, the 46mm black inserter broke. Item was swapped to the 44mm inserter and the case continued with only a slight interruption. No adverse health consequence has been reported as the result of the malfunction. | (b)(4). Foreign report source: (b)(6). The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=10526140&pc=LPH

## Inspections
City | State | Zip Code | Inspection End Date | Project Area | Classification
Valence Cedex 9 | - | - | 2017-02-03 | Compliance: Devices | No Action Indicated (NAI)
Valence Cedex 9 | - | - | 2017-02-03 | Postmarket Assurance: Devices | No Action Indicated (NAI)

## Devices
Device Summary | Device Detail | TPLC Information | # Reported Events
Device Name | ENDOBON XENOGRAFT GRANULES | TPLC Device Problem | MDR
Device Class | 2 | Failure to Osseointegrate33 | 16
Product Code | LYC | Adverse Event Without Identified Device or Use Problem34 | 14
Regulation No | 872.39 | Patient-Device Incompatibility35 | 8
Premarket Submission No | K110449 | Insufficient Information36 | 7
Establishment Operations | Manufacturer | Product Quality Problem37 | 4
Recalls | 1
Device Name | TEMPLATE | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions124 | 5563
Product Code | HWT | No Consequences Or Impact To Patient125 | 1771
Regulation No | 888.48 | Insufficient Information126 | 235
Premarket Submission No | - | No Code Available127 | 139
Establishment Operations | Manufacturer | No Patient Involvement128 | 115
Recalls | -
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HSB | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTY | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | JDW | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | KTT | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | LXT | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | Break29 | 10678
Product Code | LXH | Device-Device Incompatibility30 | 3479
Regulation No | 888.45 | Fracture31 | 2599
Premarket Submission No | - | Material Deformation32 | 1333
Establishment Operations | Manufacturer | Material Twisted/Bent33 | 1264
Recalls | -
Device Name | STAPLE, FIXATION BONE; SCREW FIXATION BONE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K931155 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | BIT, DRILL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 3549
Product Code | HTW | No Consequences Or Impact To Patient129 | 2172
Regulation No | 888.45 | Foreign Body In Patient130 | 881
Premarket Submission No | - | No Code Available131 | 342
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient132 | 308
Recalls | -
Device Name | REAMER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions123 | 10673
Product Code | HTO | No Consequences Or Impact To Patient124 | 4137
Regulation No | 888.45 | No Code Available125 | 459
Premarket Submission No | - | Insufficient Information126 | 396
Establishment Operations | Manufacturer | Foreign Body In Patient127 | 321
Recalls | -
Device Name | TAP, BONE | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions65 | 214
Product Code | HWX | No Consequences Or Impact To Patient66 | 127
Regulation No | 888.45 | No Patient Involvement67 | 78
Premarket Submission No | - | Foreign Body In Patient68 | 26
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient69 | 22
Recalls | -
Device Name | STARTER, BONE SCREW | TPLC Device Problem | MDR
Device Class | 1 | Unspecified Infection30 | 1
Product Code | HWD | No Patient Involvement31 | 1
Regulation No | 888.45 | Foreign Body In Patient32 | 1
Premarket Submission No | - | No Known Impact Or Consequence To Patient33 | 1
Establishment Operations | Manufacturer | No Information34 | 1
Recalls | -
Device Name | GUIDE, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 1714
Product Code | FZX | No Consequences Or Impact To Patient129 | 784
Regulation No | 878.48 | No Patient Involvement130 | 390
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 198
Establishment Operations | Manufacturer | No Code Available132 | 148
Recalls | -
Device Name | SCREWDRIVER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 4583
Product Code | HXX | No Consequences Or Impact To Patient129 | 2515
Regulation No | 888.45 | No Patient Involvement130 | 881
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 867
Establishment Operations | Manufacturer | No Code Available132 | 314
Recalls | -
Device Name | FORCEPS, WIRE HOLDING | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions27 | 2
Product Code | HYA | No Patient Involvement28 | 1
Regulation No | 878.48 | Foreign Body In Patient29 | 1
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | INSTRUMENT, BENDING OR CONTOURING | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions37 | 38
Product Code | HXP | No Consequences Or Impact To Patient38 | 8
Regulation No | 888.45 | No Patient Involvement39 | 4
Premarket Submission No | - | No Known Impact Or Consequence To Patient40 | 1
Establishment Operations | Manufacturer | Not Applicable41 | 1
Recalls | -
Device Name | PLIERS, SURGICAL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions48 | 49
Product Code | HTC | No Consequences Or Impact To Patient49 | 33
Regulation No | 878.48 | No Patient Involvement50 | 23
Premarket Submission No | - | No Known Impact Or Consequence To Patient51 | 5
Establishment Operations | Manufacturer | Foreign Body In Patient52 | 2
Recalls | -
Device Name | RETRACTOR | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions95 | 190
Product Code | GAD | No Consequences Or Impact To Patient96 | 76
Regulation No | 878.48 | No Known Impact Or Consequence To Patient97 | 41
Premarket Submission No | - | No Patient Involvement98 | 33
Establishment Operations | Manufacturer | Foreign Body In Patient99 | 14
Recalls | -
Device Name | GAUGE, DEPTH | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions75 | 933
Product Code | HTJ | No Consequences Or Impact To Patient76 | 308
Regulation No | 888.43 | No Patient Involvement77 | 248
Premarket Submission No | - | No Known Impact Or Consequence To Patient78 | 24
Establishment Operations | Manufacturer | Insufficient Information79 | 18
Recalls | -
Device Name | TRAY, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions119 | 396
Product Code | FSM | No Consequences Or Impact To Patient120 | 297
Regulation No | 878.48 | No Known Impact Or Consequence To Patient121 | 38
Premarket Submission No | - | No Patient Involvement122 | 24
Establishment Operations | Manufacturer | No Code Available123 | 22
Recalls | -
Device Name | ACE HUMERAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K9346439 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | AIM SUPRACONDYLAR NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K934962 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K882831 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY AE TK2 HIP SCREW SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem56 | 1765
Product Code | KTT | Break57 | 764
Regulation No | 888.3 | Device-Device Incompatibility58 | 380
Premarket Submission No | K972629 | Migration59 | 354
Establishment Operations | Manufacturer | Fracture60 | 129
Recalls | 2
Device Name | ACE RECONSTRUCTIVE PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K905595 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ACE CORTICAL BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K912598 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE A.I.M. TITANIUM TIBIAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K972183 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 2
Device Name | ACE AIM TTC FUSION NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K003797 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE BROOKER BONE PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K833865 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ZINDRICK IM ROD | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K871539 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 1
Device Name | REVERSE CUTTING CANNULATED CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K892214 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE CANNULATED SELF TAPPING CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K903810 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | 6.5MM SOLID CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K954485 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ZINDRICK IM ROD INSTRUMENTS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K882443 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | SMALL BONE LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K081546 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 1
Device Name | ACE 3.5MM CORTICAL BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K905048 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE HUMERUS & RADIUS PLATES | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K955472 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | THACKRAY TITANIUM MINI-PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K905774 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | DEPUY ACE TIMAX META PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K983853 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ANTEROLATERAL AND MEDIAL LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072832 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | POLYAX LOCKED PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K060969 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 10
Device Name | PE.R.I. II KNEE FRACTURE SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K003235 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | 4.5MM LOCKING BROAD & NARROW COMPRESSION PLATES, 4.5MM & 5.5MM CORTICAL & CANCELLOUS LOCKING SCREWS, RESPECTIVELY. | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072423 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | LOCKING ANATOMIC & COMPOSITE PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K083843 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | CLOSED FUSION PLATE, MTP PLATE, TALUS PLATE, MODEL 824077021, 22; 824677021, 22, 824071001, 824071101, 824671001, 82467 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101240 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | SCARF THREAD-HEAD SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K962236 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE UNIVERSAL AND TROCH ENTRY FEMORAL NAIL SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K033329 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | HFN 125 DEGREE 9-13 MM, 180MM, HFN 130 DEGREE 9-13 MM, HFN RH 125 DEGREE 9-13MM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K100238 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 100
Device Name | UNIVERSAL AND TROCH ENTRY FEMORAL NAILING SYSTEMS LINE EXTENSION | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K111444 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE BIO WIN TIBIAL NAILING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K013430 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | DISPENSER, CEMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions63 | 53
Product Code | KIH | No Consequences Or Impact To Patient64 | 41
Regulation No | 888.42 | Foreign Body In Patient65 | 9
Premarket Submission No | - | No Known Impact Or Consequence To Patient66 | 7
Establishment Operations | Manufacturer | Insufficient Information67 | 7
Recalls | -
Device Name | MIXER, CEMENT, FOR CLINICAL USE | TPLC Device Problem | MDR
Device Class | 1 | No Patient Involvement70 | 151
Product Code | JDZ | No Clinical Signs, Symptoms or Conditions71 | 135
Regulation No | 888.42 | No Consequences Or Impact To Patient72 | 35
Premarket Submission No | - | No Known Impact Or Consequence To Patient73 | 30
Establishment Operations | Manufacturer | No Information74 | 6
Recalls | -
Device Name | TROCHANTERIC NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K010780 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 84
Device Name | ACE VERSANAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K0231159 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | DEPUY ACE VERSANAIL TIBIAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K032097 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE NANCY NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem47 | 354
Product Code | HTY | Break48 | 196
Regulation No | 888.3 | Fracture49 | 71
Premarket Submission No | K032687 | Mechanical Jam50 | 44
Establishment Operations | Manufacturer | Material Twisted/Bent51 | 19
Recalls | -
Device Name | ACE ANTEGRADE RETROGRADE HUMERAL NAIL SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K033878 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | TROCHANTERIC NAIL SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K042325 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 11
Device Name | MEMORY STAPLE | TPLC Device Problem | MDR
Device Class | 2 | Break52 | 394
Product Code | JDR | Adverse Event Without Identified Device or Use Problem53 | 229
Regulation No | 888.3 | Detachment of Device or Device Component54 | 122
Premarket Submission No | K060746 | Fracture55 | 74
Establishment Operations | Manufacturer | Material Frayed56 | 37
Recalls | -
Device Name | FRS SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K062352 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | SMALL FRAGMENT LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072083 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | MULTIPLICITY OF PRODUCTS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K782123 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ACE CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K872859 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE OPEN-BACK HALO RING | TPLC Device Problem | MDR
Device Class | 2 | Device-Device Incompatibility33 | 25
Product Code | JEC | Detachment of Device or Device Component34 | 3
Regulation No | 888.3 | Break35 | 3
Premarket Submission No | K914858 | Mechanical Jam36 | 2
Establishment Operations | Manufacturer | Adverse Event Without Identified Device or Use Problem37 | 2
Recalls | -
Device Name | NANCY NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem47 | 354
Product Code | HTY | Break48 | 196
Regulation No | 888.3 | Fracture49 | 71
Premarket Submission No | K960642 | Mechanical Jam50 | 44
Establishment Operations | Manufacturer | Material Twisted/Bent51 | 19
Recalls | -
Device Name | TWIST-OFF SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K962233 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | MEMORY STAPLE (20 SIZES AVAILABLE) | TPLC Device Problem | MDR
Device Class | 2 | Break52 | 394
Product Code | JDR | Adverse Event Without Identified Device or Use Problem53 | 229
Regulation No | 888.3 | Detachment of Device or Device Component54 | 122
Premarket Submission No | K964226 | Fracture55 | 74
Establishment Operations | Manufacturer | Material Frayed56 | 37
Recalls | -
Device Name | TWIST-OFF (TM) SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K971069 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE AIM(R) TITANIUM SUPRACONDYLAR NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K974781 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | MODIFICATION OF DEPUY ACE TIMAX MEDIAL PILON PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K990120 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | Refobacin Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | LOD | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K171540 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | Refobacin Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | MBB | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K171540 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | Biomet Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | LOD | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K172408 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | MODIFICATION TO: LIGAFIX INTERFERENCE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K070507 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | 1
Device Name | LIGAFIX INTERFACE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K090994 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | A. L. P. S. Clavicle Plating System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K173767 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | Biomet Proximal Humerus Plating System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K143697 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 18
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | POSTERIOR LATERAL PLATES | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K103408 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | PROXIMAL TIBIA LOCKING PLATE STANDARD AND LARGE, 3.5MM LOW PROFILE AND MULTIDIRECTIONAL SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090877 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | PROXIMAL TIBIA LOCKING PLATE STANDARD AND LARGE, 3.5MM LOW PROFILE AND MULTIDIRECTIONAL SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090877 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | LAPIDUS PLATE, 3.5MM INLINE FUSION PLATE, 3.5MM MTP PLATE (LEFT AND RIGHT) | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K131670 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | 1.5MM SMALL LOCKING PLATE, 1.5MM NON-LOCKING AND LOCKING SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090492 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -

## Info
Establishment Name | Error | Timestamp
BIOMET FRANCE | 510K Table Not Found | 2023-08-07 13:39
BIOMET FRANCE | Warning Letters Not Found | 2023-08-07 14:16