## Entity Summary
Field | Extracted Data | Remarks
Establishment Name | BIOMET UK LTD
Registration No | 3002806535
Registration Year | 2023
Total No Of Devices | 153
Class 3 Devices | 9
No of Inspections | 4
No Of Recalls | 5
Warning Letters | -
1.Response Letter | N/A
2.Closeout Letter | N/A
Adverse Events | 500
Is This an adverse Report? | 429
Establishment | BIOMET UK LTD. WATERTON INDUSTRIAL ESTATE BRIDGEND, GB CF31 3XA Registration Number: 3002806535 FEI Number*: 3002806535 Status: Active Date of Registration Status: 2023
Owner/Operator | Biomet, Inc. 56 East Bell Dr. PO Box 587 Warsaw, IN US 46582 Owner/Operator Number: 1825034
Official Correspondent | phillip evans Waterton Industrial Estate Bridgend, GB CF31 3XA Phone: 44-1656-655221 US Agent: Holly Seppanen Zimmer Biomet 1800 West Center Street Warsaw, IN US 46580 Phone: 952 8470494 ext Email: holly.seppanen@zimmerbiomet.com
510(k) Device Name | Applicant | 510(k)Number | Decision Date
biolox® delta ceramic heads, biolox® delta option ceramic heads mr labeling | Biomet UK Ltd. | K200959 | 07/29/2020
biolox delta ceramic heads, biolox delta option ceramic heads | Biomet UK Ltd. | K192683 | 02/27/2020
oxford fixed lateral bearing partial knee replacement (bearings / trays) | BIOMET UK LTD. | K133940 | 03/28/2014

## Recalls
Recall Link | Trade Name | Recall Class | Posted internet Date | Termination Date | FDA Determine Cause | Quantity in Commerce | Manufacturer Reason For Recall | Action | Distribution
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=166946 | Device Recall The HipLOC Compression Hip Screw | Class 2 | - | May 12, 2020 | Process design | 1 unit | HipLOC System Plate Barrel due to an undersized bore on the HipLOC plate barrel, can result in the lag screw not fitting through the barrel | Zimmer Biomet issued US recall notices issued on August 9, 2018 via email and FedEx to Distributors and Hospitals advising of the problem and instruction to remove and return.Your Zimmer Biomet sales representative will remove the affected product from your facility. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.comnistructions.mailto:CorporateQQuality.PostMarket@zimmerbiomet.com. | US Nationwide in the states of AZ, MO, NJ and NM
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=166972 | Device Recall HipLOC Compression Hip Screw | Class 2 | - | May 12, 2020 | Process design | 6 units | HipLOC System Plate Barrel due to an undersized bore on the HipLOC plate barrel, can result in the lag screw not fitting through the barrel | Zimmer Biomet issued US recall notices issued on August 9, 2018 via email and FedEx to Distributors and Hospitals advising of the problem and instruction to remove and return.Your Zimmer Biomet sales representative will remove the affected product from your facility. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.comnistructions.mailto:CorporateQQuality.PostMarket@zimmerbiomet.com. | US Nationwide in the states of AZ, MO, NJ and NM
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=166971 | Device Recall HipLOC Compression Hip Screw | Class 2 | - | May 12, 2020 | Process design | 44 units | HipLOC System Plate Barrel due to an undersized bore on the HipLOC plate barrel, can result in the lag screw not fitting through the barrel | Zimmer Biomet issued US recall notices issued on August 9, 2018 via email and FedEx to Distributors and Hospitals advising of the problem and instruction to remove and return.Your Zimmer Biomet sales representative will remove the affected product from your facility. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.comnistructions.mailto:CorporateQQuality.PostMarket@zimmerbiomet.com. | US Nationwide in the states of AZ, MO, NJ and NM
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=166947 | Device Recall HipLOC Compression Hip Screw | Class 2 | - | May 12, 2020 | Process design | 10 units | HipLOC System Plate Barrel due to an undersized bore on the HipLOC plate barrel, can result in the lag screw not fitting through the barrel | Zimmer Biomet issued US recall notices issued on August 9, 2018 via email and FedEx to Distributors and Hospitals advising of the problem and instruction to remove and return.Your Zimmer Biomet sales representative will remove the affected product from your facility. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.comnistructions.mailto:CorporateQQuality.PostMarket@zimmerbiomet.com. | US Nationwide in the states of AZ, MO, NJ and NM

## MAUDE
Report Number | Event Date | FDA Received Date | Report Lag | Event Type | Report Source | Manufacturer-D | Manufacturer-G | Brand Name | Type Of Device | Manufacturer Review | Event Text | Manufacturer Narrative | Was Device Available For Evaluation ? | Is This an Adverse Event Report ? | MAUDE Link
3002806535-2023-00229 | Result Not Found | 07/31/2023 | 992 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 32MM +3 NK | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that the patient underwent and initial right total hip arthroplasty. Subsequently, approximately 8 months post-op the patient developed persistent lateral thigh and groin pain. Approximately 9 months post-op, an injection of steroids and local aesthetic was given under fluoroscopic guidance. The patient has since completed the study with improvement in pain and all implants remaining in place. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: medical product: g7 pps ltd acet shell 46b; catalog no. : 010000660; lot no. : 6575360. Medical product: g7 hi-wall e1 liner 32mm b; catalog no. : 010000924; lot no. : 6351755. Medical product: tprlc 133 fp type1 pps so 8. 0; catalog no. : 51100080; lot no. : 6298802. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17423161&pc=LZO&device_sequence_no=1
3002806535-2023-00259 | 01/14/2020 | 07/28/2023 | 1291 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN BIOMET OXFORD TIBIAL TRAY | KNEE PROSTHESIS | - | It was reported that a patient was revised approximately one year and eight months post initial left partial knee replacement due to pain, subsidence, and tibial implant loosening. Attempts to obtain additional information have been made; however, no more information is available. | (b)(4). D-10 - unknown oxford femoral component, item# unknown, lot# unknown. Unknown oxford bearing, item# unknown, lot# unknown. H3 ¿ other: device evaluation could not be performed as part# and lot# unknown. Also, device location is unknown. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without product identification. Device is used for treatment. Medical records were provided and reviewed. Third party review states medial compartment arthroplasty in normal alignment, no lucency, no effusion, no osseous abnormalities present. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. H3 other text : see h10. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17411889&pc=NRA&device_sequence_no=1
3002806535-2023-00261 | 06/27/2023 | 07/27/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL IMPLANT | KNEE PROSTHESIS | - | It was reported that the patient underwent a knee arthroplasty and approximately 17 years later a revision surgery was performed due to the fracture of the tibial component. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: foreign: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : product is unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17405832&pc=NRA&device_sequence_no=1
3002806535-2023-00260 | 07/18/2023 | 07/27/2023 | 9 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 36MM F | G7 CERAMIC LINER | - | It was reported that a patient underwent left hip arthroplasty. Approximately 2 years later, a revision procedure was performed due to dislocation. During the revision procedure impingement and scar tissue were also noted. The cup, head, and liner were exchanged without any reported complications and no further information is known at this time. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: g7 bonemaster ltd acet shl; item# 010000705; lot# 6813056. Biolox option, head, xl, 36/+7, taper 12/14; item# 00-8777-036-04; lot# 3099768. G2 - foreign: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17405783&pc=LPH&device_sequence_no=1
3002806535-2023-00258 | 06/27/2023 | 07/27/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 7 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent knee revision surgery due to poly dislocation about 18 months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4), g2 ¿ foreign ¿ australia, the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17404143&pc=NRA&device_sequence_no=1
3002806535-2023-00257 | 05/08/2023 | 07/26/2023 | 79 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported the patient was implanted with non-zimmer biomet products. The non-zimmer biomet products were revised for an unknown reason. During head implantation, the head exploded inside the patient's body when it was being impacted onto the trunnion of the proximal body as per standard surgical technique. All the pieces of the fractured implant were collected. The surgery was completed with another head. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: foreign - australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17396360&pc=LZO&device_sequence_no=1
3002806535-2023-00254 | 06/26/2023 | 07/24/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROTHESIS | - | It was reported that the patient underwent revision surgery due to infection, approximately fifteen (15) years from initial knee surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ australia. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : device location is unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17378258&pc=NRA&device_sequence_no=1
3002806535-2023-00256 | 06/26/2023 | 07/24/2023 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI I/M ROD LINK | KNEE INSTRUMENT | - | It was reported that during the surgical procedure, part of the oxford i/m rod link fractured and broke off. The broken piece was retrieved, and no consequences or impact to the patient has been reported. Due diligence is in progress for this complaint; to date no additional information has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17377199&pc=HWA&device_sequence_no=1
3002806535-2023-00255 | 02/24/2023 | 07/21/2023 | 147 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF FEMORAL SLAP HAMMER | KNEE INSTRUMENT | - | It was reported that the product is broken. No patient involvement was reported. No further event information available at the time of this report. | (b)(4). G2 ¿ foreign ¿ spain. Visual examination of the returned product identified item and lot numbers match the complaint. Product was not returned in its original packaging. Device shows wear and damage and is missing part of the fractured spring. A review of the device manufacturing records confirmed no abnormalities or deviations. Device is used for treatment. Reported event did not occur in an operating room or as part of a medical procedure; therefore medical records are not required for review. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17369740&pc=HRY&device_sequence_no=1
3002806535-2023-00253 | 06/23/2023 | 07/19/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/0MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported a patient underwent a hip revision approximately two months post implantation due to instability and an infection. During the procedure, while removing the head and liner, the stem fell out of the patient¿s femur. The head, liner and the stem were removed and replaced. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D ¿10: continuum tm shell clust 52 ii; item#: 00-8757-052-01, lot#: 65792472; hxpe liner elevated ii 32; item#: 00-8752-010-32, lot#: 65780369; tprlc 133 fp type1 pps ho 8. 0; item#: 51-101080, lot#: 3622261. G2 ¿ foreign ¿ new zealand. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17349242&pc=LZO&device_sequence_no=1
3002806535-2023-00252 | 07/05/2023 | 07/18/2023 | 13 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TAPERLC BMPC LAT 6.0X132 12/14 | TAPERLOC PPS TYPE 1 | - | It was reported that during the surgery, it was detected that the distal part of the stem had broken through the sterile packaging. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: spain. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17341379&pc=JDI&device_sequence_no=1
3002806535-2023-00251 | 06/13/2023 | 07/18/2023 | 35 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/+6MM T1 | HIP PROSTHESIS | - | It was reported that patient underwent revision of total hip prosthesis due to ceramic head fracture approximately two weeks after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign- netherlands. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17341368&pc=LZO&device_sequence_no=1
3002806535-2023-00250 | 06/22/2023 | 07/18/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that approximately seven years post initial implantation, the patient underwent a revision surgery for a knee bearing dislocation after a fall. Attempts have been performed and no further event information available at the time of this report. | (b)(4). Associated products: oxf uni tib tray sz c rm pma; item# 154723; lot# 3645306 oxford uni twin-peg femoral md cocr med; item# 166942; lot# j3645657 foreign: japan. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records confirmed no abnormalities or deviations. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17340653&pc=NRA&device_sequence_no=1
3002806535-2023-00247 | 06/19/2023 | 07/17/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision due to the poly wear and instability in unknown timeframe after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17332377&pc=NRA&device_sequence_no=1
3002806535-2023-00249 | 05/09/2012 | 07/17/2023 | 4086 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN CUP | HIP PROSTHESIS | - | It was reported that computed tomography revealed slight protrusion of the hip cup into the pelvis minor (acetabulum defect). Ct-imaging approximately eleven (11) years post implantation showed no further protrusion compared to immediately postop. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products: item #650-0791 item name #biolox delta lnr 32mm 50-52mm lot #unknown; item #650-0833 item name #delta cer fm hd 032/-4. 0 12/14 lot #unknown; item #650-0953 item name #micro taperlc std pc 9mm 12/14 lot #unknown. Foreign ¿ austria. Unknown item and lot number. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Unknown item and lot number. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17331785&pc=LPH&device_sequence_no=1
3002806535-2023-00243 | 08/18/2011 | 07/17/2023 | 4351 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | MICRO TAPERLC LAT PC 9MM 12/14 | COATED HIP FEMUR PROSTHESIS, MODULAR | - | It was reported through a clinical study that a patient underwent implantation with hip prostheses approximately twelve (12) years ago. Postoperatively, one (1) day after implantation, patient has experienced hip luxation. The injury was treated by repositioning under anesthesia. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: foreign - austria. D10: biolox delta lnr 32mm 50-52mm; item# 650-0791; lot# unknown; delta cer fm hd 032/-4. 0 12/14; item# 650-0833; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00240; 3002806535 -2023 -00242. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17331401&pc=KWA&device_sequence_no=1
3002806535-2023-00240 | 08/18/2011 | 07/17/2023 | 4351 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA LNR 32MM 50-52MM | CERAMIC ACETABULAR LINER | - | It was reported through a clinical study that a patient underwent implantation with hip prostheses approximately twelve (12) years ago. Postoperatively, one (1) day after implantation, patient has experienced hip luxation. The injury was treated by repositioning under anesthesia. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: foreign - austria. D10: delta cer fm hd 032/-4. 0 12/14; item# 650-0833; lot# unknown; micro taperlc lat pc 9mm 12/14; item# 650-0973; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00242; 3002806535 -2023 -00243. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17331400&pc=LZO&device_sequence_no=1
3002806535-2023-00242 | 08/18/2011 | 07/17/2023 | 4351 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FM HD 032/-4.0 12/14 | CERAMIC FEMORAL HEAD PROSTHESIS | - | It was reported through a clinical study that a patient underwent implantation with hip prostheses approximately twelve (12) years ago. Postoperatively, one (1) day after implantation, patient has experienced hip luxation. The injury was treated by repositioning under anesthesia. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign - austria. Biolox delta lnr 32mm 50-52mm; item# 650-0791; lot# unknown; micro taperlc lat pc 9mm 12/14; item# 650-0973; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00240; 3002806535 -2023 -00243. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17331399&pc=LZO&device_sequence_no=1
3002806535-2023-00248 | 04/25/2023 | 07/17/2023 | 83 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 4 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision surgery due mobile bearing dislocation. During the revision valgus laxity was noted in extension and 20 degrees flexion. The mobile bearing was replaced without complications. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ oxf uni tib tray sz c lm pma item#154722 lot#unknown. Oxf twin-peg cmntd fem sm pma item#161468 lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17331394&pc=NRA&device_sequence_no=1
3002806535-2023-00245 | - | 07/12/2023 | - | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNCEMENTED TIBIAL IMPACTOR | ORTHOPAEDIC IMPLANT IMPACTOR | - | It was reported that tip broke off the instrument during surgery. No impact on a patient. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Denmark. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17305104&pc=LXH&device_sequence_no=1
3002806535-2023-00244 | - | 07/11/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that a surgeon suspects that an existing poly bearing implant has rotated 180 degrees, causing complications. No surgical procedure or revision has taken place at this point. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Unknown oxford femoral component; item number# unknown; lot# unknown. Unknown oxford tibial component; item number# unknown; lot# unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Item number and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17297243&pc=NRA&device_sequence_no=1
3002806535-2023-00218 | 06/06/2023 | 07/11/2023 | 35 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN HEAD | HIP PROSTHESIS | - | It was reported that patient underwent a total hip arthroplasty and approximately 3 years later a revision surgery was performed due to dislocation and loosening of the liner. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: unknown liner; item# unknown; lot# unknown. G2 - foreign: poland. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00217. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : product is unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296955&pc=LZO&device_sequence_no=1
3002806535-2023-00217 | 06/06/2023 | 07/11/2023 | 35 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | E1 TAPERED STD +3MM LNR 32/39 | E1 TAPERED LINERS | - | It was reported that patient underwent a total hip arthroplasty and approximately 3 years later a revision surgery was performed due to dislocation and loosening of the liner. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: unknown head; item# unknown; lot# unknown. Foreign: poland. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00217; 3002806535-2023-00218. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Product is unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296953&pc=OQG&device_sequence_no=1
3002806535-2023-00219 | 06/06/2023 | 07/11/2023 | 35 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | EXC ABT STD SHELL HA/PC 050MM | EXCEED ABT CMTLESS CUPS HA/PC | - | It was reported that a revision surgery was performed due to dislocation and loosening of the liner. However, while the polyethylene part came out as planned, the metal part remained and could not be removed. The procedure ended with the removal of the acetabular cup and its replacement with a revision one. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: unknown liner; item# unknown; lot# unknown. G2 - foreign: poland. The original revision surgery has been reported in a separate complaint. The associated medwatch# are the following: 3002806535 - 2023 - 00217, 3002806535 - 2023 - 00218. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296952&pc=JDI&device_sequence_no=1
3002806535-2023-00236 | 06/12/2023 | 07/11/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial right unilateral (medial side) partial knee replacement. Subsequently, was revised to a total knee six months later due to pain, difficulty with ambulation, and extensive radiolucency along the tibial implant which appears loose and subsided laterally. There is also suspected lucency along the femoral implant without frank loosening. Bone quality is markedly osteopenic. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ oxf twin-peg cmntd fem sm pma, item# 161468 , lot# 635200. Oxf uni tib tray sz c rm pma, item# 154723, lot# 65717233. Unknown cement, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00234; 3002806535 - 2023 - 00235. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296145&pc=NRA&device_sequence_no=1
3002806535-2023-00235 | 06/12/2023 | 07/11/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial right unilateral (medial side) partial knee replacement. Subsequently, was revised to a total knee six months later due to pain, difficulty with ambulation, and extensive radiolucency along the tibial implant which appears loose and subsided laterally. There is also suspected lucency along the femoral implant without frank loosening. Bone quality is markedly osteopenic. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxf twin-peg cmntd fem sm pma, item# 161468 , lot# 635200. Oxf anat brg rt sm size 3 pma, item# 159568, lot# 164880. Unknown cement, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00234. 3002806535 - 2023 - 00236. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296140&pc=NRA&device_sequence_no=1
3002806535-2023-00231 | 06/13/2023 | 07/11/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent left knee revision surgery due to unknown reason approximately six (6) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxford pks anatomic arcom meniscal brg - small left size 5, item# 159542, lot# 735940. Oxford pks twin peg femur - small, item# 161468, lot# 913310. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00232 and 3002806535-2023-00233. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17296005&pc=NRA&device_sequence_no=1
3002806535-2023-00234 | 06/12/2023 | 07/11/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial right unilateral (medial side) partial knee replacement. Subsequently, was revised to a total knee six months later due to pain, difficulty with ambulation, and extensive radiolucency along the tibial implant which appears loose and subsided laterally. There is also suspected lucency along the femoral implant without frank loosening. Bone quality is markedly osteopenic. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxf uni tib tray sz c rm pma, item# 154723, lot# 65717233; oxf anat brg rt sm size 3 pma, item# 159568, lot# 164880; unknown cement, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00235; 3002806535-2023-00236. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17295985&pc=NRA&device_sequence_no=1
3002806535-2023-00238 | 05/05/2023 | 07/10/2023 | 66 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | MICRO TAPRLC STD PC 12.5 12/14 | HIP PROSTHESIS | - | It was reported that patient experienced groin pain about twelve (12) years after initial hip surgery. Mri done showed healed stress fracture at the tip of the hip stem. The mri results were discussed with the patient and at that time patient was pain free. The adverse event is considered resolved. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ associated products: item#: 650-0830, item name#: delta cer fm hd 028/-3. 5 12/14, lot #: unknown; item#: 650-0790, item name#: biolox delta lnr 28mm 46-48mm, lot #: unknown. G2 ¿ foreign ¿ austria. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17286919&pc=KWA&device_sequence_no=1
3002806535-2023-00239 | 12/18/2020 | 07/10/2023 | 934 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported the patient underwent an initial right total hip arthroplasty due to hip fracture. Subsequently, the patient experienced a dislocation due to unknown reasons and was reduced in the emergency department. Later on, the patient underwent a head and liner exchange due to an intra-prosthetic dislocation (disassociation) while bending. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: g7 dual mobility liner 40mm d; item# 110024462; lot# unknown. Act artic e1 hip brg 28x40mm; item# ep-200146; lot# unknown. G2 - foreing: south korea. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without product identification. Device is used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17286763&pc=LZO&device_sequence_no=1
3002806535-2023-00237 | 09/28/2020 | 07/10/2023 | 1015 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/+3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient underwent an initial right total hip arthroplasty. Subsequently the patient experienced pain after showering. Intra-prosthetic dislocation (disassociation) was noted, resulting in a head and liner exchange. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: g7 dual mobility liner 40mm d; item# 110024462; lot# unknown. Act artic e1 hip brg 28x40mm; item# ep-200146; lot# unknown. G2 - foreign: south korea. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without product identification. Device is used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17286760&pc=LZO&device_sequence_no=1
3002806535-2023-00233 | 06/13/2023 | 07/10/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent left knee revision surgery due to unknown reason approximately six (6) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxford pks cemented tibia, alpha, standard - lm size b, item# 154720, lot# 777150; oxford pks anatomic arcom meniscal brg - small left size 5, item# 159542, lot# 735940. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00231, 3002806535-2023-00232. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17286100&pc=NRA&device_sequence_no=1
3002806535-2023-00232 | 06/13/2023 | 07/10/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent left knee revision surgery due to unknown reason approximately six (6) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxford pks cemented tibia, alpha, standard - lm size b, item# (b)(4), lot# 777150; oxford pks twin peg femur - small, item# (b)(4), lot# 913310. Multiple mdr reports were filed for this event, associated reports: 3002806535-2023-00231; 3002806535-2023-00233. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17286094&pc=NRA&device_sequence_no=1
3002806535-2023-00220 | 06/09/2023 | 07/07/2023 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that t-handle broke while trying to remove im rod during surgery. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign ¿ canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17276883&pc=LXH&device_sequence_no=1
3002806535-2023-00221 | 06/12/2023 | 07/07/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty. Approximately 7 years later the patient was revised to a total knee arthroplasty. During the revision, medial laxity was noted, significant damage to the medial capsular tissue and portions of the mcl, as well as bone quality to be soft and tissue looked to be atrophied. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical product ¿ oxf twin-peg cmntd fem sm pma item#161468 lot#unknown oxf anat brg lt sm size 7 pma item#159544 lot#865430. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00222, 3002806535-2023-00223. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17276821&pc=NRA&device_sequence_no=1
3002806535-2023-00222 | 06/12/2023 | 07/07/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty. Approximately 7 years later the patient was revised to a total knee arthroplasty. During the revision, medial laxity was noted, significant damage to the medial capsular tissue and portions of the mcl, as well as bone quality to be soft and tissue looked to be atrophied. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ oxf uni tib tray sz c lm pma item#154722 lot#unknown. Oxf anat brg lt sm size 7 pma item#159544 lot#865430. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00221, 3002806535-2023-00223. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17276819&pc=NRA&device_sequence_no=1
3002806535-2023-00223 | 06/12/2023 | 07/07/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 7 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty. Approximately 7 years later the patient was revised to a total knee arthroplasty. During the revision, medial laxity was noted, significant damage to the medial capsular tissue and portions of the mcl, as well as bone quality to be soft and tissue looked to be atrophied. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ oxf uni tib tray sz c lm pma item#154722 lot#unknown. Oxf twin-peg cmntd fem sm pma item#161468 lot#unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00221, 3002806535-2023-00222. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17276798&pc=NRA&device_sequence_no=1
3002806535-2023-00225 | 06/14/2023 | 07/05/2023 | 21 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | BIOLOX DELTA CER FEM HD LG DIA | - | It was reported that the patient underwent a total hip arthroplasty and subsequently a revision surgery was performed due to dislocation. Attempts have been made and no further information has been provided. | Cmp-(b)(4). D10 - medical devices: tprlc 133 mp type1 bm so 15. 0; item# 51-116150; lot# 6222310. G2 - foreign: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17259004&pc=LZO&device_sequence_no=1
3002806535-2023-00228 | 06/07/2023 | 07/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 4 PMA | KNEE PROTHESIS | - | It was reported that patient underwent a revision surgery due to tibial fracture, approximately six (6) days after initial surgery. Attempts have been made and no further information has been provided. | (b)(4). D10 - oxf uni cmntls tib sz b lm item#166572 lot#7165186. G2 ¿ foreign ¿ new zealand. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00227. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17255638&pc=NRA&device_sequence_no=1
3002806535-2023-00226 | 06/07/2023 | 07/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 28MM | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure three (3) years and four (4) months post implantation due to a dislocation of the patient's dual mobility bearing and disassociated head from the active articulation head. Attempts have been made and no further information has been provided. | (b)(4). D10: associated products: item reference(b)(4), item name: cer option type 1 tpr sleve -6, lot number 2971089. Item reference:(b)(4), item name: act artic e1 hip brg 28x44mm, lot number 885030. Item reference:(b)(4), item name: g7 dual mobility liner 44mm f, lot number 870660. Item reference: (b)(4), item name: bone scr 6. 5x30 self-tap, lot number 64531906. Item reference: (b)(4), item name: g7 osseoti multihole 54mm f, lot number 6575098. Item reference: (b)(4), item name: tprlc 133 t1 pps so 14x148mm, lot number 6537092. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17255076&pc=LZO&device_sequence_no=1
3002806535-2023-00227 | 06/07/2023 | 07/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI CMNTLS TIB SZ B LM | KNEE PROTHESIS | - | It was reported that patient underwent a revision surgery due to tibial fracture, approximately six (6) days after initial surgery. Attempts have been made and no further information has been provided. | (b)(4). D10 - oxf anat brg lt sm size 4 pma item#159541, lot#6722311. G2 ¿ foreign ¿ new zealand. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00228. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17255009&pc=NRA&device_sequence_no=1
3002806535-2023-00224 | 06/15/2023 | 07/05/2023 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 6 PMA | OXFORD HXLPE BEARINGS | - | It was reported that a patient underwent an initial right knee arthroplasty and approximately 7 years later a revision surgery was performed due to bearing dislocation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: saw blades stryker power systems; item# 11-4844; lot# 63009; oxf uni tib tray sz c rm pma; item# 154723; lot# 3819019; oxford uni twin-peg femoral md; item# 166942; lot# j3608906; oxford cless res 3pk stry 6; item# 506273; lot# p398138. G2 - foreign: australia. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17254556&pc=NRA&device_sequence_no=1
3002806535-2023-00198 | 05/25/2023 | 07/04/2023 | 40 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf uni tib tray sz d rm pma, item #(b)(4), lot# 999960. Oxf anat brg rt lg size 3 pma item#(b)(4), lot#385560. Biomet bone cement r 1x40 us item# (b)(4), lot# 745aah1407. Multiple mdr reports were filed for this event, see associated reports: 3002806535-2023-00197; 3002806535-2023-00200. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17253689&pc=NRA&device_sequence_no=1
3002806535-2023-00197 | 05/25/2023 | 07/03/2023 | 39 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT LG SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf uni tib tray sz d rm pma, item #154725, lot# 999960. Oxf twin-peg cmntd fem lg pma, item#161470, lot# 016890. Biomet bone cement r 1x40 us item# 110035368, lot# 745aah1407. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00198, 3002806535-2023-00200. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17247467&pc=NRA&device_sequence_no=1
3002806535-2023-00200 | 05/25/2023 | 07/03/2023 | 39 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf twin-peg cmntd fem lg pma, item161470, lot# 016890. Oxf anat brg rt lg size 3 pma item#159582, lot#385560. Biomet bone cement r 1x40 us item# 110035368, lot# 745aah1407. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00197, 3002806535-2023-00198. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17246875&pc=NRA&device_sequence_no=1
3002806535-2023-00216 | 06/03/2023 | 07/03/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN BEARING COMPONENT | KNEE IMPLANT | - | It was reported that patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10 ¿ associated products. Item #unknown femoral component lot #unknown. Item #unknown tibial component lot #unknown. G2- foreign - (b)(6). H3 - item and lot number unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00214; 3002806535-2023-00215. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17246634&pc=NRA&device_sequence_no=1
3002806535-2023-00215 | 06/03/2023 | 07/03/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN TIBIAL COMPONENT | KNEE IMPLANT | - | It was reported that patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10 ¿ associated products: item #unknown femoral component lot #unknown. Item #unknown bearing component lot #unknown. G2 ¿ foreign ¿ poland. H3 - item and lot number unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00214, 3002806535-2023-00215, 3002806535-2023-00216. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17246605&pc=NRA&device_sequence_no=1
3002806535-2023-00214 | 06/03/2023 | 07/03/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN FEMORAL COMPONENT | KNEE IMPLANT | - | It was reported that patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10 ¿ associated products: item #unknown tibial component lot #unknown. Item #unknown bearing component lot #unknown. G2 ¿ foreign ¿ poland. H3 - item and lot number unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00214, 3002806535-2023-00215, 3002806535-2023-00216. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17246560&pc=NRA&device_sequence_no=1
3002806535-2023-00213 | 06/05/2023 | 06/30/2023 | 25 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TRL BRG W SLOTS MED 3MM | KNEE INSTRUMENT | - | It was reported that instrument broke into two pieces, which were retrieved from patient. Surgery was completed without any further complications nor delay. Attempts have been made and no further information has been provided. | (b)(4). G2 ¿ foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17236976&pc=NRA&device_sequence_no=1
3002806535-2023-00212 | 05/15/2023 | 06/28/2023 | 44 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROSTHESIS | - | It was reported that patient underwent a revision approximately four (4) years after initial left total hip arthroplasty. During the revision some metallosis was noted within the soft tissues, acetabular liner was fragmented, the head had significant metal like wear, and one screw was prominent of 0. 5mm within the acetabulum. The head, liner, taper, and ringloc were replaced and the prominent screw was removed. The joint was irrigated without debridement and the surgery was completed without complications. Attempts have been made and no further information has been provided. | (b)(4). D10: associated products: item reference: ep-105790, item name: epoly 36mm +5 rglc lnr hw sz23, item lot:315070. Item reference: 103533, item name: ti low profile screw 6. 5x30mm, item lot:590690. Item reference: 650-1066, item name: cer opt type 1 tpr sleve 0mm, item lot:2946658. Item reference: pt-116054, item name: regen/rnglc+ ltd 54mm sz 23, item lot:497840. Item reference: 103533, item name: ti low profile screw 6. 5x30mm, item lot:590690. Item reference: 51-103090, item name: tprlc 133 type1 pps so 9x137mm, item lot:6137664. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17215615&pc=LZO&device_sequence_no=1
3002806535-2023-00211 | 11/17/2022 | 06/27/2023 | 222 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNK OXFORD BEARING | PROSTHESIS, KNEE, FEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, METAL/POLYMER | - | It was reported by legal that a right initial unicompartmental knee arthroplasty was performed. Subsequently, approximately six years and four months post-procedure the patient was revised due to pain, noise, difficulty ambulating, and dislocation. During the revision, it was noted the polyethylene was fractured. Tibia and femur remained implanted. Polyethylene exchanged without complications. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). D10: unknown medium femur; item# unknown; lot# unknown. Unknown size c tibia; item# unknown; lot# unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item / lot unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17208629&pc=HRY&device_sequence_no=1
3002806535-2023-00210 | 12/06/2022 | 06/27/2023 | 203 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that a patient underwent an initial left total hip arthroplasty. Subsequently, two weeks later the patient was revised due to fractured femoral head with poly damage. The poly liner and femoral head were revised without complications. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: 36mm i. D. Size e neutral; item# 30103605; lot# 65516876. G7 pps ltd acet shell 52e; item# 010000663; lot# 7311439. Tprlc 133 t1 pps ho 12x144mm; item# 51-104120; lot# 6992221. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17207345&pc=LZO&device_sequence_no=1
3002806535-2023-00172 | 05/28/2023 | 06/27/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that the patient underwent a revision procedure due to ceramic head fracture, 3 days after the initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source south korea. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17206980&pc=LZO&device_sequence_no=1
3002806535-2023-00209 | 06/08/2023 | 06/27/2023 | 19 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 5 PMA | OXFORD HXLPE BEARINGS | - | It was reported that the patient underwent initial uni knee arthroplasty. Approximately 10 years later, a revision surgery was performed due to the bearing dislocating, likely caused by bearing wear. | (b)(4). Report source foreign: japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17206972&pc=NRA&device_sequence_no=1
3002806535-2023-00208 | 09/18/2019 | 06/26/2023 | 1377 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that approximately four (4) years after initial left unicompartmental knee arthroplasty, the patient underwent revision surgery due to polyethylene dislocation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item / lot unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17199226&pc=HRY&device_sequence_no=1
3002806535-2023-00207 | 06/05/2023 | 06/26/2023 | 21 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | OXFORD HXLPE BEARINGS | - | It was reported that a patient underwent an initial knee arthroplasty. Approximately 19 years later a revision procedure was performed due to articular surface wear and fracture. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17198176&pc=NRA&device_sequence_no=1
3002806535-2023-00206 | 01/31/2022 | 06/26/2023 | 511 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 36MM STD NK | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that patient underwent initial right hip arthroplasty. Approximately 8 years later, the patient began having progressively worse right hip pain. The source of the pain could not be confirmed. Subsequently, patient was revised 13 year later due to metallosis. Attempts have been made and no further information has been provided. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17197786&pc=LZO&device_sequence_no=1
3002806535-2023-00205 | 05/19/2023 | 06/23/2023 | 35 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported the patient was implanted with non-zimmer biomet products. The non-zimmer biomet products were revised for an unknown reason, and the patient was implanted with the arcos stem. The stem was undersized due to an iatrogenic error on the surgeon¿s behalf during the first revision. Eleven days post implantation the patient was revised due to the stem under sizing and dislocation that occurred due to the stem under sizing. There was no need to adjust the acetabular side as it remained stable. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - 11-300916 arcos 16x190mm spl tpr dist 943100. 11-301353 arcos con sz c hi 80mm 590390. 00223200204 integral short gtr w/ 2 cables 56705650. 00625006550 bone screw self-tapping 6. 5 mm dia. 50 mm length 64084511. 00625006550 bone screw self-tapping 6. 5 mm dia. 50 mm length j7307819. 00875201336 36mm i. D. Size ll elevated rim liner use with 58mm o. D. Size ll shell 64239463. 00875705802 58mm o. D. Size ll porous uncemented with multi-holes shell use with ll liners 63290514. 31-323240 3. 2mmx40mm rnglc+ acet drl bit 233750. 31-323260 3. 2mmx60mm rnglc+ acet drl bit 597010. 310. 250s drill bit 2. 5mm, l 110/85mm 2-flute, f/quick coupling 9l85899 depuy. G2: foreign - australia. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17188892&pc=LZO&device_sequence_no=1
3002806535-2023-00204 | 03/27/2023 | 06/22/2023 | 87 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent a right knee revision surgery due to unknown reason approximately seven years nine months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxf uni tib tray sza rm, item# 154719,lot# 125040. Oxf twin-peg cmntd fem sm pma, item# 161468, lot# 924810. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00202. 3002806535 - 2023 - 00203. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17180200&pc=NRA&device_sequence_no=1
3002806535-2023-00203 | 03/27/2023 | 06/22/2023 | 87 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZA RM | KNEE PROSTHESIS | - | It was reported that patient underwent a right knee revision surgery due to unknown reason approximately seven years nine months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxf twin-peg cmntd fem sm pma, item# 161468, lot# 924810; oxf anat brg rt sm size 5 pma, item# 159570, lot# 439320. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00202; 3002806535 - 2023 - 00204. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17180135&pc=NRA&device_sequence_no=1
3002806535-2023-00202 | 03/27/2023 | 06/22/2023 | 87 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent a right knee revision surgery due to unknown reason approximately seven years nine months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxf uni tib tray sza rm, item# 154719,lot# 125040. Oxf anat brg rt sm size 5 pma, item# 159570, lot# 439320. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00203. 3002806535 - 2023 - 00204. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17180122&pc=NRA&device_sequence_no=1
3002806535-2023-00201 | 05/23/2022 | 06/20/2023 | 393 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | HIP PROSTHESIS | - | It was reported that a clinical study patient experienced a myocardial infarction fourteen (14) days after initial left total hip arthroplasty. No further information has been provided to date. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - g7 pps ltd acet shell 58g, item# 010000666, lot# 7128460; g7 vit e neutral lnr 36mm g, item# 30103607, lot# 65224087; tprlc 133 t1 pps ho 18x156mm, item# 51-104180, lot# 6966785. G2 ¿ foreign ¿ the netherlands. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17161118&pc=LZO&device_sequence_no=1
3002806535-2023-00199 | 05/16/2023 | 06/19/2023 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIA | UNICONDYLAR KNEE PROSTHESIS | - | It was reported that 10 years after implantation of the unicondylar knee prosthesis on the right side, the patient is experiencing significant osteolysis in the tibial head. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2- germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item number and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17152006&pc=NRA&device_sequence_no=1
3002806535-2023-00194 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford bearing component ; item#:unknown ;lot#:unknown ;item name: unknown oxford femoral component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00192, 3002806535 - 2023 - 00193. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149104&pc=HRY&device_sequence_no=1
3002806535-2023-00194 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford bearing component ; item#:unknown ;lot#:unknown ;item name: unknown oxford femoral component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00192, 3002806535 - 2023 - 00193. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149104&pc=NRA&device_sequence_no=1
3002806535-2023-00193 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford tibial component ; item#:unknown ;lot#:unknown ;item name: unknown oxford femoral component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00192. 3002806535 - 2023 - 00194. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149101&pc=HRY&device_sequence_no=1
3002806535-2023-00193 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford tibial component ; item#:unknown ;lot#:unknown ;item name: unknown oxford femoral component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00192. 3002806535 - 2023 - 00194. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149101&pc=NRA&device_sequence_no=1
3002806535-2023-00192 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford tibial component ; item#:unknown ;lot#:unknown ;item name: unknown oxford bearing component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00193. 3002806535 - 2023 - 00194. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149100&pc=HRY&device_sequence_no=1
3002806535-2023-00192 | 05/25/2023 | 06/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent partial knee arthroplasty. Subsequently, the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10. Item#:unknown ;lot#:unknown ;item name: unknown oxford tibial component ; item#:unknown ;lot#:unknown ;item name: unknown oxford bearing component ; multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00193. 3002806535 - 2023 - 00194. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17149100&pc=NRA&device_sequence_no=1
3002806535-2023-00196 | 05/26/2023 | 06/16/2023 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF FEMORAL SLAP HAMMER | BMET GENERAL KNEE INSTRUMENTS | - | It was reported that the spring of the slap hammer fractured. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17142734&pc=HWA&device_sequence_no=1
3002806535-2023-00195 | 05/22/2023 | 06/16/2023 | 25 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN TRIAL BEARING | KNEE INSTRUMENT | - | It was reported that trial insert fractured during surgery. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ poland. H-3 - unknown product. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : unknown product. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17142710&pc=HRY&device_sequence_no=1
3002806535-2023-00190 | 05/25/2023 | 06/15/2023 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI I/M ROD LINK | BMET GENERAL KNEE INSTRUMENTS | - | It was reported that during the surgery the rod link broke off. However, the broken piece was successfully retrieved, and the surgery proceeded as planned. No consequences or impact to the patient. Attempts have been made and no further information has been provided. | (b)(4). G2 - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17135349&pc=HWT&device_sequence_no=1
3002806535-2023-00191 | 06/22/2020 | 06/15/2023 | 1088 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN CERAMIC HEAD | HIP PROSTHESIS | - | It was reported that patient underwent a left hip total arthroplasty. Approximately a year and a half after initial surgery, has been diagnosed with bursitis and treated with formal physical therapy and over the counter topical cream for ongoing pain. | (b)(4). D10: associated products: item number: 193109, item name: echo b-mtrc mp fp ho 9 , item lot: 089870. Item number: 110010243, item name:g7 osseoti 3 hole shell 50mm d, item lot: 6378467. Item number: 010000856, item name:g7 neutral e1 liner 36mm d, item lot: 6383738. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without product identification. Device is used for treatment. Medical records were provided and reviewed by a health care professional. Patient presented with greater trochanteric bursitis of the left hip was proscribed stretching exercises, contributing factors of event: recurrent falls due to alcoholism and recent treatment for cellulitis/superficial infection of the buttock which can increase the inflammation in the joint. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. H3 other text : product and product data not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17133566&pc=LZO&device_sequence_no=1
3002806535-2023-00189 | 05/16/2023 | 06/13/2023 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FM HD 036/+4MM 12/14 M 12/14 | HIP PROSTHESIS | - | It was reported that the product had debris inside the sterile packaging. Product was not used. No patient involvement. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17115434&pc=LZO&device_sequence_no=1
3002806535-2023-00186 | 05/16/2023 | 06/12/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to instability about three (3) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxf anat brg lt md size 5 pma, item# 159549, lot# 6376325. Oxford uni twin-peg femoral md, item# 166942, lot# j6561560. G2 ¿ foreign ¿ australia. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00187, 3002806535-2023-00188. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17105864&pc=NRA&device_sequence_no=1
3002806535-2023-00187 | 05/16/2023 | 06/12/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to instability about three (3) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxf uni tib tray sz b lm pma, item# 154720, lot# 6499378; oxford uni twin-peg femoral md, item# 166942, lot# j6561560. Foreign ¿ australia. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00186. 3002806535-2023-00188. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17105859&pc=NRA&device_sequence_no=1
3002806535-2023-00188 | 05/16/2023 | 06/12/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNI TWIN-PEG FEMORAL MD | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to instability about three (3) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxf uni tib tray sz b lm pma, item# 154720, lot# 6499378; oxf anat brg lt md size 5 pma, item# 159549, lot# 6376325. Foreign ¿ australia. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00186. 3002806535-2023-00187. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17105856&pc=NRA&device_sequence_no=1
3002806535-2023-00185 | 05/16/2023 | 06/12/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 36MM +3 NK | HIP PROSTHESIS | - | It was reported that approximately 4 years post implantation of left hip arthroplasty, the patient experienced a work related fall and suffered a clavicle fracture as a result of the fall. After the fall, the patient began experiencing squeaking during ambulation, along with slight discomfort. The patient was subsequently revised. During the revision, it was noted that the femoral head was wearing against the shell. The head and liner were exchanged, with the surgeon exchanging the liner from a size 36 to 40. Due diligence is in progress for this complaint; to date no additional information has been received. | (b)(4). D10: item#: 010000742; item name: g7 neutral arcomxl lnr 36mm g; lot#: 6109134. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17105637&pc=LZO&device_sequence_no=1
3002806535-2023-00179 | 05/15/2023 | 06/07/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery due to unknown reason approximately nine (9) years post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products: item #159576, item name #oxf anat brg rt md size 4 pma, lot #697840. Item #154723, item name #oxf uni tib tray sz c rm pma, lot #146370. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00177. 3002806535 -2023 -00178. 3002806535 -2023 -00179. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17078820&pc=NRA&device_sequence_no=1
3002806535-2023-00178 | 05/15/2023 | 06/07/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery due to unknown reason approximately nine (9) years post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 ¿ associated products: item #159576, item name #oxf anat brg rt md size 4 pma, lot #697840; item #161469, item name #oxf twin-peg cmntd fem md pma, lot #413870. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00177, 3002806535 -2023 -00179. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17078798&pc=NRA&device_sequence_no=1
3002806535-2023-00177 | 05/15/2023 | 06/07/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery due to unknown reason approximately nine (9) years post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products. Item #154723, item name #oxf uni tib tray sz c rm pma, lot #146370 item #161469, item name #oxf twin-peg cmntd fem md pma, lot #413870 multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00178; 3002806535 -2023 -00179. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17078772&pc=NRA&device_sequence_no=1
3002806535-2023-00184 | 05/15/2023 | 06/07/2023 | 23 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL, SEMI-CONSTRAINED | - | It was reported that during hip arthroplasty, the surgeon in error implanted a medium-size femoral implant instead of the intended small-size femoral implant. The tibia and bearing sizes were correct. No revision surgery is currently planned, the patient will be closely monitored by the surgeon. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). D10: oxf uni tib tray sz c lm pma; item#154722; lot#783650; oxf anat brg lt sm size 3 pma; item#159540; lot#765570; the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17077468&pc=NRA&device_sequence_no=1
3002806535-2023-00183 | 05/12/2023 | 06/07/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF MP SLOT POST CUT GD MED | BMET GENERAL KNEE INSTRUMENTS | - | It was reported that the blue plug disassembled from the instrument. The blue plug was found and did not fall into the patient. However, post op x-rays confirm the metal pin that secures the blue plug was retained in the patients body. The surgeon tried to remove it by scope with no result. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17077178&pc=LXH&device_sequence_no=1
3002806535-2023-00182 | 05/11/2023 | 06/06/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | OXFORD UNCEMENTED KNEE SYSTEM INSTRUMENTS | - | It was reported that during an initial knee arthroplasty, the posterior part of the instrument broke off and was retained in the patient. A post-operative x-ray confirmed the presence of a small metal thread. However, it was reported that this incident did not cause any harm to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: item#:154926; lot#:7406233; item name: oxford ph3 cementless fem sz m. Item#:166576; lot#: 7290104; item name: oxf uni cmntls tib sz d lm. Item#: 159547; lot#: 7476363; item name: oxf anat brg lt md size 3 pma. G2- norway. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17070038&pc=LXH&device_sequence_no=1
3002806535-2023-00181 | 05/11/2023 | 06/06/2023 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENT | - | It was reported that during the partial left knee arthroplasty that the instrument fractured. No piece fell nor remain into the patient body. All pieces were found. No impact or injury to patient. Attempts have been made and no further information has been provided. | (b)(4). Associated products: item reference: (b)(4), item name: oxf uni ph3 tib impactor, item lot: zb7117655. Foreign- canada. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00180. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17069671&pc=LXH&device_sequence_no=1
3002806535-2023-00180 | 05/11/2023 | 06/06/2023 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI PH3 TIB IMPACTOR | KNEE INSTRUMENT | - | It was reported that during the partial left knee arthroplasty that the instrument fractured. No piece fell nor remain into the patient body. All pieces were found. No impact or injury to patient. Attempts have been made and no further information has been provided. | (b)(4). Item reference: 32-422097, item name: oxf cmntls implant insert, item lot: zb130404. Report source foreign- canada. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00181. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17069670&pc=LXH&device_sequence_no=1
3002806535-2023-00175 | 05/10/2023 | 06/05/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent right knee revision due to loosening of tibial component approximately twenty (20) months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - oxf anat brg rt md size 5 pma, item# 159577, lot# 769520; oxf twin-peg cmntd fem md pma, item# 161469, lot# 586420. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17058406&pc=NRA&device_sequence_no=1
3002806535-2023-00173 | 05/11/2023 | 06/01/2023 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI CMNTLS TIB SZ C LM | UNICONDYLAR KNEE PROSTHESIS | - | It was reported that the inner packaging had no vacuum, and it was unsure if sterilization was intact. Another implant was used during procedure. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17037241&pc=NRA&device_sequence_no=1
3002806535-2023-00171 | 05/08/2023 | 05/31/2023 | 23 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that in the process of a hip revision procedure, the femoral head fractured while being impacted unto the proximal stem body. This subsequently caused an injury to the patient. The surgeon took some time to clean the wound and remove the shards of the ceramic head. All pieces were taken out by the surgeon. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign : australia. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17029088&pc=LZO&device_sequence_no=1
3002806535-2023-00170 | 03/24/2023 | 05/30/2023 | 67 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/-3MM T1 | HIP PROSTHESIS | - | It was reported the patient was implanted with the g7 dual mobility system. The head came off the bearing three (3) weeks postoperatively. The patient was revised due to head and bearing disassembly. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). 28mm i. D. 44mm o. D. Size f bearing, item# 110031012, lot# 65461996. Foreign ¿ japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17018155&pc=LZO&device_sequence_no=1
3002806535-2023-00169 | - | 05/30/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to unknown reason about two (2) years after initial surgery. No further event information available at the time of this report. | (b)(4). D10: unknown oxford bearing, item# unknown, lot# unknown; unknown oxford femoral component, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00167, 3002806535-2023-00168. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records cannot be performed without product identification. Devices are used for treatment. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. H3 other text : device location unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17018033&pc=NRA&device_sequence_no=1
3002806535-2023-00168 | - | 05/30/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to unknown reason about two (2) years after initial surgery. No further event information available at the time of this report. | (b)(4). D10 - unknown oxford femoral component, item# unknown, lot# unknown; unknown oxford tibial component, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00167. 3002806535-2023-00169. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records cannot be performed without product identification. Devices are used for treatment. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. Device location unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17018028&pc=NRA&device_sequence_no=1
3002806535-2023-00167 | - | 05/30/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to unknown reason about two (2) years after initial surgery. No further event information available at the time of this report. | (b)(4). D10 - unknown oxford bearing, item# unknown, lot# unknown; unknown oxford tibial component, item# unknown, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00168, 3002806535-2023-00169. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records cannot be performed without product identification. Devices are used for treatment. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. H3 other text : device location unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17018026&pc=NRA&device_sequence_no=1
3002806535-2023-00166 | 09/10/2020 | 05/30/2023 | 992 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | According to legal allegations, a patient who underwent an initial left total hip arthroplasty experienced acute onset of pain, limited mobility, and radiographic evidence revealed acetabular dissociation. During the subsequent revision surgery, the presence of metallosis was detected between the femoral head and the metal acetabular shell, and a broken poly liner consisting of five pieces with noticeable wear was discovered. All components, excluding the stem, were revised without any reported intraoperative complications. | (b)(4). D10: item#:15-103203 ;lot#:565060 ;item name: taperloc microp lat fmrl 9. 0mm. Item#:12-104152 ;lot#:447270 ;item name: m/h radial 3-hole shell 52mm. Item#:ep-108223 ;lot#:279030 ;item name: e-poly 36mm +3 maxrom lnr sz23. Item#:103530 ;lot#:939650 ;item name: ti low profile screw 6. 5x15mm. Item#:103531 ;lot#:921530 ;item name: ti low profile screw 6. 5x20mm. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records confirmed no abnormalities or deviations. Device is used for treatment. Medical records were provided and reviewed by a health care professional. Evidence of metallosis between the femoral head articulating against the metal acetabular shell. Contributing factors of event: osteopenic bone, multiple prior surgeries, hx crowe iii hip dysplasia, acetabular component noted to have excessive anteversion. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17017987&pc=LZO&device_sequence_no=1
3002806535-2023-00165 | 03/10/2023 | 05/26/2023 | 77 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 36MM STD NK | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient developed severe pain requiring medication approximately two years post initial right total hip arthroplasty. Upon receipt of the medication, the issue was resolved approximately one month later. Attempts have been made and no further information is available. | (b)(4). D10 - item#: 110010244 / g7 osseoti 3 hole shell 52mm e; lot 6424109. Item#: 010000935 / g7 hi-wall e1 liner 36mm e; lot #: 6374859. Item#: 193111 / echo b-mtrc mp fp ho 11; lot #: 163920. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17009344&pc=LZO&device_sequence_no=1
3002806535-2023-00129 | 04/12/2023 | 05/26/2023 | 44 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI CMNTLS TIB SZ C LM | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial unicompartmental knee arthroplasty, and approximately 1. 5 months later experienced pain. Diagnosis confirmed a medial tibia bone fracture. Revision surgery was performed. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf anat brg lt sm size 3 pma, item # 159540, lot# 7280185. Oxf pks c/less fem cocr sz sml pc/ha # 154925, lot# 7268399. G2 ¿ foreign ¿ japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17008754&pc=NRA&device_sequence_no=1
3002806535-2023-00128 | 04/26/2023 | 05/25/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient was revised from a oxford pks to a total knee for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf uni tib tray sz c rm pma, item# 154723, lot# unknown. Oxf twin-peg cmntd fem sm pma, item# 161468 , lot# unknown. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00164, 3002806535-2023-00127. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000982&pc=NRA&device_sequence_no=1
3002806535-2023-00164 | 04/26/2023 | 05/25/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that the patient was revised from a oxford pks to a total knee for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf twin-peg cmntd fem sm pma item# 161468, lot# unknown. Oxf anat brg rt sm size 3 pma item# 159568, lot# unknown. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00127, 3002806535-2023-00128. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000866&pc=NRA&device_sequence_no=1
3002806535-2023-00163 | 04/27/2023 | 05/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 6 PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL, SEMI-CONSTRAINED | - | It was reported that the patient underwent revision surgery due to unknown reasons. Patient revised to total knee with medial augment. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: : oxf twin-peg cmntd fem md pma, item: 161469, lot: 817640 oxf uni tib tray sz d rm pma, item: 154725, lot: 999690 multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2023 -00161, 3002806535 -2023 -00162. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000815&pc=NRA&device_sequence_no=1
3002806535-2023-00162 | 04/27/2023 | 05/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL, SEMI-CONSTRAINED | - | It was reported that the patient underwent revision surgery due to unknown reasons. Patient revised to total knee with medial augment. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: oxf anat brg rt md size 6 pma, item: 159578, lot: 056640. Oxf uni tib tray sz d rm pma, item: 154725, lot: 999690. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2023 -00161, 3002806535 -2023 -00163. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000813&pc=NRA&device_sequence_no=1
3002806535-2023-00161 | 04/27/2023 | 05/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D RM PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL, SEMI-CONSTRAINED | - | It was reported that the patient underwent revision surgery due to unknown reasons. Patient revised to total knee with medial augment. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: oxf twin-peg cmntd fem md pma, item: 161469, lot: 817640. Oxf anat brg rt md size 6 pma, item: 159578, lot: 056640. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2023 -00162, 3002806535 -2023 -00163. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000812&pc=NRA&device_sequence_no=1
3002806535-2023-00127 | 04/26/2023 | 05/25/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that the patient was revised from a oxford pks to a total knee for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10- oxf uni tib tray sz c rm pma item# 154723, lot# unknown. Oxf anat brg rt sm size 3 pma item# 159568, lot# unknown. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00164, 3002806535-2023-00128. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=17000654&pc=NRA&device_sequence_no=1
3002806535-2023-00159 | 04/25/2023 | 05/24/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEM | KNEE PROSTHESIS | - | It was reported that a patient underwent a revision from an oxford pks to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: united kingdom. D10: medical product: unknown oxford bearing; catalog no. : unknown; lot no. : unknown. Medical product: unknown oxford tib; catalog no. : unknown; lot no. : unknown. Multiple mdr reports were filed for this event, please see associated reports (add this for a multiple item complaint): 3002806535-2023-00142, 3002806535-2023-00160. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item/lot unknown. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16990340&pc=NRA&device_sequence_no=1
3002806535-2023-00160 | 04/25/2023 | 05/24/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL | KNEE PROSTHESIS | - | It was reported that a patient underwent a revision from an oxford pks to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: united kingdom. D10: medical product: unknown oxford bearing; catalog no. : unknown; lot no. : unknown. Medical product: unknown oxford femoral; catalog no. : unknown; lot no. : unknown. Multiple mdr reports were filed for this event, please see associated reports (add this for a multiple item complaint): 3002806535-2023-00159, 3002806535-2023-00142. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16990329&pc=NRA&device_sequence_no=1
3002806535-2023-00156 | 05/10/2023 | 05/24/2023 | 14 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF GAP GAUGE MED 5/6MM | KNEE MICROPLASTY INSTRUMENTS | - | It was reported that during surgery while using the device, plastic parts present on the site were noticed. The surgery was delayed due to the need of removing plastic from the site. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ germany the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16990200&pc=HTJ&device_sequence_no=1
3002806535-2023-00158 | 05/10/2023 | 05/24/2023 | 14 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF GAP GAUGE MED 1/2MM | KNEE MICROPLASTY INSTRUMENTS | - | It was reported that during surgery while using the device, plastic parts present on the site were noticed. The surgery was delayed due to the need of removing plastic from the site. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ germany the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16990199&pc=HTJ&device_sequence_no=1
3002806535-2023-00142 | 04/25/2023 | 05/23/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that a patient underwent a revision from an oxford pks to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). United kingdom. Medical product: unknown oxford fem; catalog no. : unknown; lot no. : unknown. Medical product: unknown oxford tib; catalog no. : unknown; lot no. : unknown. Multiple mdr reports were filed for this event, please see associated reports (add this for a multiple item complaint): 3002806535-2023-00159. 3002806535-2023-00160. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Item/lot unknown. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16983538&pc=NRA&device_sequence_no=1
3002806535-2023-00157 | 04/15/2023 | 05/23/2023 | 38 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure approximately 6 months post implantation due to a fall. The patient was no longer stable due to tissue damage. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: cer option type 1 tpr sleve -3; item# 650-1065; lot# 3130574. G7 vit e neutral lnr 36mm f; item# 30103606; lot# 65596487. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00152. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16983318&pc=LZO&device_sequence_no=1
3002806535-2023-00153 | 05/10/2023 | 05/23/2023 | 13 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF GAP GAUGE MED 3/4MM | KNEE MICROPLASTY INSTRUMENTS | - | It was reported that during surgery while using the device, plastic parts present on the site were noticed. The surgery was delayed due to the need of removing plastic from the site. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16983305&pc=HTJ&device_sequence_no=1
3002806535-2023-00152 | 04/15/2023 | 05/23/2023 | 38 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPTION TYPE 1 TPR SLEVE -3 | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure approximately 6 months post implantation due to a fall. The patient was no longer stable due to tissue damage. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: cer bioloxd option hd 36mm; item# 650-1057; lot# 3102327. G7 vit e neutral lnr 36mm f; item# 30103606; lot# 65596487. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00157. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16983269&pc=LZO&device_sequence_no=1
3002806535-2023-00154 | 05/01/2023 | 05/23/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 5 PMA | KNEE PROTHESIS | - | It was reported that patient underwent a knee revision surgery due to unknown reason. Approximately four (4) years ten (10) months from initial surgery. Attempts have been made and no further information has been provided. | (b)(4). D10 - oxf twin-peg cmntd fem sm pma, item#: 161468, lot#: 124440; oxf uni tib tray sz c lm pma, item#: 154722, lot#: 163640. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00151; 3002806535-2023-00155. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16982353&pc=NRA&device_sequence_no=1
3002806535-2023-00155 | 05/01/2023 | 05/23/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROTHESIS | - | It was reported that patient underwent a knee revision surgery due to unknown reason. Approximately four (4) years ten (10) months from initial surgery. Attempts have been made and no further information has been provided. | (b)(4). D10 - oxf twin-peg cmntd fem sm pma item#161468 lot#: 124440; oxf anat brg lt sm size 5 pma item#: 159542 lot#: 825790. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00151; 3002806535-2023-00154. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16982352&pc=NRA&device_sequence_no=1
3002806535-2023-00151 | 05/01/2023 | 05/23/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROTHESIS | - | It was reported that patient underwent a knee revision surgery due to unknown reason. Approximately four (4) years ten (10) months from initial surgery. Attempts have been made and no further information has been provided. | (b)(4). D10 - oxf anat brg lt sm size 5 pma item# 159542 lot#825790; oxf uni tib tray sz c lm pma item#154722 lot#163640. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00154, 3002806535-2023-00155. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16982350&pc=NRA&device_sequence_no=1
3002806535-2023-00150 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D RM PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial right knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: associated products: item reference: unknown, item name: unknown oxford femoral component right, item lot: unknown. Item reference: unknown, item name: unknown oxford bearing right, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00148, 3002806535-2023-00149. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960891&pc=NRA&device_sequence_no=1
3002806535-2023-00149 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial right knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: concomitant medical products: item reference: unknown, item name: unknown oxford tibial component right, item lot: unknown. Item reference: unknown, item name: unknown oxford bearing right, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00148 and 3002806535-2023-00150. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text: product data and product not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960689&pc=NRA&device_sequence_no=1
3002806535-2023-00148 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial right knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: concomitant medical products: item reference: unknown, item name: unknown oxford tibial component right, item lot: unknown. Item reference: unknown, item name: unknown oxford femoral component right, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00149 and 3002806535-2023-00150. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text: product data and product not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960686&pc=NRA&device_sequence_no=1
3002806535-2023-00144 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial left knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: associated products: item reference: unknown, item name: unknown oxford tibial component left, item lot: unknown. Item reference: unknown, item name: unknown oxford femoral component left, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00145, 3002806535-2023-00146. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : product and product data not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960680&pc=NRA&device_sequence_no=1
3002806535-2023-00146 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial left knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: associated products: item reference: unknown, item name: unknown oxford bearing left, item lot: unknown. Item reference: unknown, item name: unknown oxford femoral component left, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00144, 3002806535-2023-00145. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : product data and product not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960677&pc=NRA&device_sequence_no=1
3002806535-2023-00145 | 01/20/2023 | 05/19/2023 | 119 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that approximately year after partial left knee arthroplasty patient experienced cracking, popping and pain. Bone scan showed loosening. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: concomitant medical products: item reference: unknown, item name: unknown oxford bearing left, item lot: unknown item reference: unknown, item name: unknown oxford tibial component left, item lot: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00144 and 3002806535-2023-00146. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text: product data and product not available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960674&pc=NRA&device_sequence_no=1
3002806535-2023-00141 | 04/24/2023 | 05/19/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROSTHESIS | - | It was reported that patient underwent spinal fusion surgery since the original total hip replacement was performed approximately (7) years ago. Recently, a revision has been performed due to dislocation. Revised to head and liner. No further outcome. No further event information available at the time of this report. | (b)(4). D10 ¿ cer option type 1 tpr sleve -6, item# 650-1064, lot #753270. G7 neutral e1 liner 36mm d, item # 010000856, lot #3588459. Visual examination of the returned provided pictures identified metal transfer marks on the spherical radius on the ceramic head, no further information can be gleaned from the provided photographs. A review of the device manufacturing records confirmed no abnormalities or deviations. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16960074&pc=LZO&device_sequence_no=1
3002806535-2023-00147 | - | 05/18/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that the patient underwent initial hip surgery. Post-operative hip dislocation was diagnosed approximately six (6) months after surgery. Revision surgery is planned. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ netherlands. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16952445&pc=LZO&device_sequence_no=1
3002806535-2023-00139 | 04/17/2023 | 05/16/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent uni to total knee conversion revision surgery due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: oxf twin-peg cmntd fem md pma; item#161469; lot#279160. Oxf anat brg lt md size 3 pma; item#159547; lot#744230. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00138, 3002806535 - 2023 - 00140. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16936966&pc=NRA&device_sequence_no=1
3002806535-2023-00140 | 04/17/2023 | 05/16/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent uni to total knee conversion revision surgery due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: oxf twin-peg cmntd fem md pma; item#161469; lot#279160. Oxf uni tib tray sz d lm pma; item#154724; lot#838200. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00138, 3002806535 - 2023 - 00139. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16936964&pc=NRA&device_sequence_no=1
3002806535-2023-00138 | 04/17/2023 | 05/16/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent uni to total knee conversion revision surgery due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: oxf uni tib tray sz d lm pma; item#154724; lot#838200. Oxf anat brg lt md size 3 pma; item#159547; lot#744230. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00139, 3002806535 - 2023 - 00140. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16936961&pc=NRA&device_sequence_no=1
3002806535-2023-00137 | 09/16/2008 | 05/15/2023 | 5354 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RECAP/MAGNUM CEMENTLESS ACETABULAR SHELL 56/50MM | HIP PROTHESIS | - | It was reported that a patient underwent an initial left mom hip surgery. Approximately nine years post-implantation follow up examinations resulted in elevated metal ion levels. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - m2a-magnum mod hd sz 50mm, item# 157450, lot# 1505536; tprloc 12/14 por lat 11x142, item# 650-0264, lot# 1476477; magnum tpr adpr ti 42- 50/+8mm, item# 130832, lot# 1338003. G2 ¿ foreign ¿ italy. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16929118&pc=MRA&device_sequence_no=1
3002806535-2023-00136 | 04/20/2023 | 05/12/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent total knee arthroplasty revision due to tibia collapse. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10: unknown oxford femoral component; item# unknown; lot# unknown. Oxf anat brg rt sm size 4 pma; item#159569; lot#1282040. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : item number and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16917695&pc=NRA&device_sequence_no=1
3002806535-2023-00133 | 04/20/2023 | 05/12/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN TAPERLOC | HIP PROSTHESIS | - | It was reported that the patient underwent a total hip arthroplasty and approximately 17 (seventeen) years later a revision surgery was performed due to a large amount of osteolysis and massive bone loss ¿like a pseudo-tumour¿. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: the united kingdom. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. H3 other text : device location is unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16917649&pc=LZO&device_sequence_no=1
3002806535-2023-00134 | 04/17/2023 | 05/11/2023 | 24 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+6MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient underwent an initial right hip arthroplasty. A week later the patient dislocated his hip. Eleven days post-implantation, revision hip surgery was performed. The ceramic head was split in half in the patient. It was replaced with a new ceramic head and new liner. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D-10 - 3. 2mmx30mm rnglc+ acet drl bit; item# 31323230, lot# 65819457. G7 osseoti 4 hole shell 54mm f; item# 110010245, lot# 65421942. Trilogy bone scr 6. 5x25; item# 00625006525, lot# 65521242. G7 vit e neutral lnr 36mm f; item# 30103606, lot# 65052048. Tprlc 133 t1 pps so 11x142mm; item# 51103110, lot# 7075445. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16909765&pc=LZO&device_sequence_no=1
3002806535-2023-00132 | 04/13/2023 | 05/09/2023 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI PH3 TIB IMPACTOR | KNEE INSTRUMENT | - | It was reported that during a left partial knee arthroplasty the impactor fractured and the fractured piece fell to the ground. The patient did not retain any foreign body. The instrument was still usable so the surgeon was able to finish the surgery. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16894020&pc=LXH&device_sequence_no=1
3002806535-2023-00131 | 02/22/2023 | 05/09/2023 | 76 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 40MM | HIP PROSTHESIS,SEMI-CONSTRAINED | - | It was reported that the patient underwent a left hip arthroplasty, and subsequently was revised due to recurring dislocations approximately 4 months after initial implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ ¿australia". D10 - medical product - g7 neutral e1 liner 36mm e; catalog #:010000857; lot#:7273174. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16893556&pc=LZO&device_sequence_no=1
3002806535-2023-00130 | 11/24/2022 | 05/04/2023 | 161 | Injury | Manufacturer | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that approximately three (3) weeks after initial right hip total arthroplasty patient underwent repositioning with anesthesia due to dislocation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2: foreign- netherlands. D10: associated products: item number: 30103608, item name:g7 vit e neutral lnr 36mm h, item lot: 65454505. Item number: 010000668, item name: g7 pps ltd acet shell 62h, item lot: r7313004a. Item number: 51-103150, item name: tprlc 133 t1 pps so 15x150mm, item lot:7198025. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16867265&pc=LZO&device_sequence_no=1
3002806535-2023-00126 | 11/22/2022 | 04/28/2023 | 157 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient experienced a dislocation approximately 15 days post right total hip arthroplasty and was corrected with a closed reduction. Subsequently, the patient experiences the second dislocation approximately 3 weeks later which was also corrected with a closed reduction. Subsequently, a revision surgery was performed approximately a month later due to recurring dislocations. Patient reports severe pain, difficulty with activities of daily living, difficulty ambulating with walker and has a limp and decreased range of motion. Attempts have been made and no further information has been provided. | (b)(4). D10 - medical devices: g7 vit e neutral lnr 32mm e; item# 30103205; lot# 65224183. G7 pps ltd acet shell 52e; item# 010000663; lot# 7289876. Tprlc 133 fp type1 pps so 7. 0; item# 51-100070; lot# 7240457. G2 - foreign: the netherlands. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. A review of the device manufacturing records confirmed no abnormalities or deviations. Device is used for treatment. Medical records were provided and reviewed by a health care professional. Contributing factors of event: pain, difficult with adl¿s, decreased rom, multiple dislocations, right hip revision due to multiple dislocation. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16830905&pc=LZO&device_sequence_no=1
3002806535-2023-00125 | 04/20/2023 | 04/27/2023 | 7 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TRL BRG W/SLOTS LGE 3MM | KNEE INSTRUMENT | - | It was reported that the one corner on the anterior end of the trial bearing broke off while the surgeon was trying to insert it into the joint. No negative impact on the patient was noted. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ united kingdom. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16821944&pc=NRA&device_sequence_no=1
3002806535-2023-00124 | 03/13/2023 | 04/27/2023 | 45 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | HIP PROSTHESIS | - | It was reported that patient underwent repositioning with anesthesia due to dislocation about four (4) months after initial surgery. No additional complications reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - item#: 30103608, lot#: 65454505, item name: g7 vit e neutral lnr 36mm h item#: 010000668, lot#: r7313004a, item name: g7 pps ltd acet shell 62h item#: 51-103150, lot#: 7198025, item name: tprlc 133 t1 pps so 15x150mm g2 ¿ foreign ¿ netherlands. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16821638&pc=LZO&device_sequence_no=1
3002806535-2023-00123 | 03/28/2023 | 04/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that approximately six (6) years after the initial surgery, a patient underwent a right sided unicompartmental knee arthroplasty to total knee arthroplasty conversion for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: oxf uni tib tray sz c rm pma; item# 154723; lot# 007610. Oxf anat brg rt md size 3 pma; item# 159575; lot# 942780. Multiple mdr reports were filed for this event: 3002806535-2023-00121; 3002806535-2023-00122. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16805292&pc=NRA&device_sequence_no=1
3002806535-2023-00122 | 03/28/2023 | 04/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that approximately six (6) years after the initial surgery, a patient underwent a right sided unicompartmental knee arthroplasty to total knee arthroplasty conversion for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: oxf uni tib tray sz c rm pma; item# 154723; lot# 007610. Oxf twin-peg cmntd fem md pma; item# 161469; lot# 120790. Multiple mdr reports were filed for this event: 3002806535-2023-00121; 3002806535-2023-00123. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16805159&pc=NRA&device_sequence_no=1
3002806535-2023-00121 | 03/28/2023 | 04/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that approximately six (6) years after the initial surgery, a patient underwent a right sided unicompartmental knee arthroplasty to total knee arthroplasty conversion for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D10 - medical devices: oxf anat brg rt md size 3 pma; item# 159575; lot# 942780. Oxf twin-peg cmntd fem md pma; item# 161469; lot# 120790. Multiple mdr reports were filed for this event: 3002806535-2023-00122; 3002806535-2023-00123. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16804964&pc=NRA&device_sequence_no=1
3002806535-2023-00120 | 03/27/2023 | 04/25/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+6MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that a patient fell while skating which led to the hip ceramic head implant breaking and fracturing. A revision surgery was then performed to remove and replace the damaged implant. No fragments were left in the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16804928&pc=LZO&device_sequence_no=1
3002806535-2023-00119 | 04/03/2023 | 04/19/2023 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | OXFORD CEMENTED FEMORALS | - | It was reported that the patient underwent an intial uni knee arthorplasty and approximately 9 (nine) years later a revision surgery was performed due to lateral wear and pain. Attempts have been made and no further information has been provided. | (b)(4). D10 ¿ medical devices: oxf anat brg rt md size 3 pma; item# 159575; lot# 785710. Oxf uni tib tray sz c rm pma; item# 154723; lot# 840570. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00117, 3002806535 - 2023 - 00118. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16767998&pc=NRA&device_sequence_no=1
3002806535-2023-00118 | 04/03/2023 | 04/19/2023 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported that the patient underwent an intial uni knee arthorplasty and approximately 9 (nine) years later a revision surgery was performed due to lateral wear and pain. Attempts have been made and no further information has been provided. | (b)(4). D10 ¿ medical devices: oxf anat brg rt md size 3 pma; item# 159575; lot# 785710. Oxf twin-peg cmntd fem md pma; item# 161469; lot# 413400. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00117, 3002806535 - 2023 - 00119. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16767994&pc=NRA&device_sequence_no=1
3002806535-2023-00117 | 04/03/2023 | 04/19/2023 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | OXFORD HXLPE BEARINGS | - | It was reported that the patient underwent an initial uni knee arthorplasty and approximately 9 (nine) years later a revision surgery was performed due to lateral wear and pain. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: ¿ medical devices: oxf uni tib tray sz c rm pma; item# 154723; lot# 840570 oxf twin-peg cmntd fem md pma; item# 161469; lot# 413400 multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00118, 3002806535 - 2023 - 00119. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16767985&pc=NRA&device_sequence_no=1
3002806535-2023-00116 | 03/20/2023 | 04/19/2023 | 30 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL RIGHT SIZE 3 | KNEE PROSTHESIS | - | It was reported that the intraoperative traveling nurse opened the product package and found that the inner sterile package was broken. Suspecting that it had been contaminated, did not use this product. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). G2 ¿ foreign ¿ china. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16766867&pc=NRA&device_sequence_no=1
3002806535-2023-00093 | 03/16/2023 | 04/14/2023 | 29 | Injury | Manufacturer | ZIMMER SWITZERLAND MANUFACTURING GMBH waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | STD BRL 3-HOLE PLATE 135 DEG | HIPLOC COMPRESSION PLATE | - | It was reported the patient underwent a trauma procedure (8) years ago. Additionally it has been indicated the patient needs to be revised due to an unknown reason. Plate removal required as confirmed by surgeon. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: hiploc compression screw 36mm, item# 236036, lot# 2015041444 std brl lag scrw 12. 5x110mm, item# 236110, lot# 2015092337 report source ¿ foreign ¿ australia multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00094 3002806535-2023-00092 the device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16742494&pc=JDO&device_sequence_no=1
3002806535-2023-00081 | 12/03/2018 | 04/13/2023 | 1592 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROSTHESIS | - | It was reported that the patient underwent a right total hip arthroplasty performed about 5 years ago. The patient was revised 9 months later due to an undersized stem implanted. | (b)(4). Concomitant medical product(s) - acetabular cup, item# 110017108, lot# 3972309; acetabular liner, item# 010000860, lot# 6109006; head adapter, item# 6501066, lot#2920994; femoral stem item# 51101170, lot# 3661564. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Not confirmed. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16731790&pc=LZO&device_sequence_no=1
3002806535-2023-00115 | 03/22/2023 | 04/13/2023 | 22 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN DRILL BIT | HIP INSTRUMENT | - | It was reported that the longest drill bit in screwdriver set is broken. No consequences or impact to the patient. Attempts have been made and no further information has been provided. | (b)(4). Foreign: poland. Product won't be returned by costumer due to hospital policy. Visual examination of the photograph provided from the hospital shows what appears to be six drill bits within an instrument tray. The longest drill bit visible in the photograph appears to show signs of the shank, which may indicate multiple use of the drill bit. Due to the quality of the photograph, a fracture cannot be confirmed. Item and lot numbers are not visible in the photograph. Review of manufacturing records cannot be performed without product identification. Devices are used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16731423&pc=HTW&device_sequence_no=1
3002806535-2023-00114 | 03/22/2023 | 04/13/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | OXFORD HXLPE BEARINGS | - | It was reported that patient underwent initial uni knee arthroplasty and approximately 8 (eight) years later a revision surgery was performed due to pain and lateral wear. Attempts have been made and no further information has been provided. | (b)(4). D10 ¿ medical devices: oxf uni tib tray sz c rm pma; item# 154723; lot# 197130. Oxf twin-peg cmntd fem md pma; item# 161469; lot# 252880. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00112, 3002806535 - 2023 - 00113, 3002806535 - 2023 - 00114. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16730384&pc=NRA&device_sequence_no=1
3002806535-2023-00113 | 03/22/2023 | 04/13/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | OXFORD CEMENTED FEMORALS | - | It was reported that patient underwent initial uni knee arthroplasty and approximately 8 (eight) years later a revision surgery was performed due to pain and lateral wear. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: oxf uni tib tray sz c rm pma; item# 154723; lot# 197130. Oxf anat brg rt md size 4 pma; item# 159576; lot# 717420. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00112, 3002806535 - 2023 - 00114. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16730382&pc=NRA&device_sequence_no=1
3002806535-2023-00112 | 03/22/2023 | 04/13/2023 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported that patient underwent initial uni knee arthroplasty and approximately 8 (eight) years later a revision surgery was performed due to pain and lateral wear. Attempts have been made and no further information has been provided. | (b)(4). D10 ¿ medical devices: oxf twin-peg cmntd fem md pma; item# 161469; lot# 252880. Oxf anat brg rt md size 4 pma; item# 159576; lot# 717420. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00112, 3002806535-2023-00113, and 3002806535-2023-00114. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16730377&pc=NRA&device_sequence_no=1
3002806535-2023-00111 | 03/13/2023 | 04/12/2023 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM STD NECK, TYPE 1 TAPER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that initial right total hip arthroplasty was performed. The patient underwent repositioning with anesthesia due to dislocation approximately three (3) weeks later. The patient dislocated for a second time due to unknown reasons requiring repositioning approximately two (2) weeks later. Subsequently the patient dislocated their hip a third time requiring repositioning approximately three (3) months later. There have been no additional complications reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item#: 30103608, lot#: 65454505, item name: g7 vit e neutral lnr 36mm h, item#: 010000668, lot#: r7313004a, item name: g7 pps ltd acet shell 62h, item#: 51-103150, lot#: 7198025, item name: tprlc 133 t1 pps so 15x150mm. Report source- netherlands. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16723879&pc=LZO&device_sequence_no=1
3002806535-2023-00110 | 12/12/2022 | 04/12/2023 | 121 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM STD NECK, TYPE 1 TAPER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that initial right total hip arthroplasty was performed. The patient underwent repositioning with anesthesia due to dislocation approximately three (3) weeks later. The patient dislocated for a second time due to unknown reasons requiring repositioning approximately two (2) weeks later. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item#: 30103608, lot#: 65454505, item name: g7 vit e neutral lnr 36mm h. Item#: 010000668, lot#: r7313004a, item name: g7 pps ltd acet shell 62h. Item#: 51-103150, lot#: 7198025, item name: tprlc 133 t1 pps so 15x150mm. Report source: netherlands. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16723837&pc=LZO&device_sequence_no=1
3002806535-2023-00109 | 11/24/2022 | 04/12/2023 | 139 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM STD NECK, TYPE 1 TAPER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that initial right total hip arthroplasty was performed. The patient underwent repositioning with anesthesia due to dislocation approximately three (3) weeks later. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item#: 30103608, lot#: 65454505, item name: g7 vit e neutral lnr 36mm h, item#: 010000668, lot#: r7313004a, item name: g7 pps ltd acet shell 62h, item#: 51-103150, lot#: 7198025, item name: tprlc 133 t1 pps so 15x150mm. Report source: netherlands. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16723836&pc=LZO&device_sequence_no=1
3002806535-2023-00091 | 12/29/2022 | 04/12/2023 | 104 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNI FEMORAL SM | KNEE PROTHESIS | - | It was reported that patient underwent right knee revision surgery due to luxation, approximately five (5) years and nine (9) months from initial implantation. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: oxf anat brg rt sm size 5 pma item#159570 lot#3032998. Oxf uni tib tray sz c rm pma item#154723 lot#3922418. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00017-1. The femoral component remains implanted and therefore not returned so visual and dimensional evaluations could not be performed. A review of the device manufacturing records for the femoral component confirmed no abnormalities or deviations. Device is used for treatment. Radiographs were provided and reviewed by a radiologist. Significant findings include subluxation of the femoral component as well as displaced polyethylene lining on the later time point. The original time points demonstrate possible improper sizing of the hardware for the bones. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16722472&pc=NRA&device_sequence_no=1
3002806535-2023-00107 | 03/13/2023 | 04/11/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that an implant for the right side was implanted during left knee implantation. As the surgeon approved the range of motion, he decided to keep it. Attempts have been made and no further information has been provided. | (b)(4). D10 ¿ associated products; oxf twin-peg cmntd fem md pma , item# 161469, lot# 085950. Oxf anat brg lt md size 3 pma, item# 159547, lot# 887160. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16711800&pc=NRA&device_sequence_no=1
3002806535-2023-00108 | 03/16/2023 | 04/11/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient underwent revision surgery due to infection less than one year post operatively. The head and liner were exchanged during revision surgery. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). G2-foreign-australia. D10. Item#:01. 0000. 860 ;lot#:6962500 ;item name: g7 neutral e1 liner 36mm h ; item#:01. 0000. 708 ;lot#:7025724 ;item name: g7 bonemaster ltd acet shl 62h; item#: 51-117200 ;lot#: 3577532;item name: tprlc 133 mp type1 bm ho 20. 0; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16711533&pc=LZO&device_sequence_no=1
3002806535-2023-00106 | 03/08/2023 | 04/07/2023 | 30 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF FEM TRIAL REMOVER | KNEE INSTRUMENT | - | It was reported that during knee arthroplasty spring in femoral trial remover broke. No patient harm. Attempts have been made and no further information has been provided. | (b)(4). Foreign- japan investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16701405&pc=LXH&device_sequence_no=1
3002806535-2023-00104 | 03/13/2023 | 04/06/2023 | 24 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF MP 04MM DRILL | KNEE INSTRUMENT | - | It was reported that the drill broke inside the femur during use. Surgeon was able to remove the broken drill from the bone without leaving any debris. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign canada. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without lot number being provided. Device is used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16687868&pc=HTW&device_sequence_no=1
3002806535-2023-00105 | 03/09/2023 | 04/06/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 28MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was reported that patient underwent revision due to retear of the gluteus medius, dislocation, and disassociation of the dual mobility components. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: associated products: item reference: 650-1068, item name: cer option type 1 tpr sleve +6, item lot: 3066235. Item reference: 110024464, item name: g7 dual mobility liner 44mm f, item lot: 725530. Item reference: 110031000, item name: longevity dm bearing 28x44mm, item lot: 65509335. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16687538&pc=LZO&device_sequence_no=1
3002806535-2023-00100 | 11/08/2022 | 04/05/2023 | 148 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due to an implant fracture, approximately 9 years post-implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford femoral; item#unknown; lot#unknown, unknown oxford bearing; item#unknown; lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00098, 3002806535 - 2023 - 00099. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681660&pc=NRA&device_sequence_no=1
3002806535-2023-00099 | 11/08/2022 | 04/05/2023 | 148 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due to an implant fracture, approximately 9 years post-implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford femoral; item#unknown; lot#unknown, unknown oxford tibial; item#unknown; lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00098, 3002806535 - 2023 - 00100. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681658&pc=NRA&device_sequence_no=1
3002806535-2023-00098 | 11/08/2022 | 04/05/2023 | 148 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due to an implant fracture, approximately 9 years post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Unknown oxford bearing; item#unknown; lot#unknown. Unknown oxford tibial; item#unknown; lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00099 3002806535 - 2023 - 00100. Other text : part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681653&pc=NRA&device_sequence_no=1
3002806535-2023-00096 | 09/18/2014 | 04/05/2023 | 3121 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN SIZE 4 OXFORD BEARING | KNEE PROSTHESIS | - | It was reported initial right unicompartmental knee replacement. Subsequently revised. During the revision noted wear to the bearing. The bearing was exchanged with the unknown component. The remaining unknown oxford components remained implanted. No contributing conditions, just worn and fractured implants attempts have been made and all additional information received has been included in this report. | (b)(4). G2- australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681293&pc=NRA&device_sequence_no=1
3002806535-2023-00103 | 03/08/2023 | 04/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a knee revision on their right side for unknown reasons, approximately four (4) years after the initial surgery. No known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf uni tib tray sz b rm pma; item# 154721; lot# 717080. Oxf twin-peg cmntd fem sm pma; item# 161468; lot# 790170. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00101, 3002806535-2023-00102. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681031&pc=NRA&device_sequence_no=1
3002806535-2023-00101 | 03/08/2023 | 04/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a knee revision on their right side for unknown reasons, approximately four (4) years after the initial surgery. No known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf twin-peg cmntd fem sm pma; item# 161468; lot# 790170. Oxf anat brg rt sm size 5 pma; item# 159570; lot# 837010. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00102, 3002806535-2023-00103. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16681001&pc=NRA&device_sequence_no=1
3002806535-2023-00102 | 03/08/2023 | 04/05/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a knee revision on their right side for unknown reasons, approximately four (4) years after the initial surgery. No known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: medical devices: oxf uni tib tray sz b rm pma; item#: 154721; lot#: 717080. Oxf anat brg rt sm size 5 pma; item#: 159570; lot#: 837010. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00101, 3002806535-2023-00103. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16680950&pc=NRA&device_sequence_no=1
3002806535-2023-00097 | 08/03/2017 | 04/05/2023 | 2071 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN SIZE 3 OXFORD BEARING | KNEE PROSTHESIS | - | It was reported initial left unicompartmental knee replacement. Subsequently revised due to implant fracture. During the revision the bearing was found to be in 2 pieces. The bearing was exchanged with unknown component. The remaining unknown oxford components remained implanted. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16679801&pc=NRA&device_sequence_no=1
3002806535-2023-00095 | 10/18/2022 | 04/04/2023 | 168 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN TAPERLOC HIP STEM | HIP PROSTHESIS | - | It was reported that patient underwent revision surgery of hip prosthesis on the right side due to unknown reasons. No further event information available at the time of this report. | (b)(4). Report source¿ foreign ¿ ireland. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of manufacturing records cannot be performed without product identification. Device is used for treatment. Medical records were not provided. With the available information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. Device not returned, insufficient information. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16672535&pc=JDI&device_sequence_no=1
3002806535-2023-00090 | 03/08/2023 | 04/03/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee replacement at an unknown date. Subsequently, after 18 years, the patient underwent revision surgery for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source-foreign-australia. Device evaluated by mfr: device evaluation could not be performed as part# and lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00088; 3002806535 -2023 -00089. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16665951&pc=NRA&device_sequence_no=1
3002806535-2023-00089 | 03/08/2023 | 04/03/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee replacement at an unknown date. Subsequently, after 18 years, the patient underwent revision surgery for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source-foreign-australia. Device evaluated by mfr: device evaluation could not be performed as part# and lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00088, 3002806535 -2023 -00090. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16665949&pc=NRA&device_sequence_no=1
3002806535-2023-00088 | 03/08/2023 | 04/03/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee replacement at an unknown date. Subsequently, after 18 years, the patient underwent revision surgery for unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | Cmp (b)(4). Report source foreign: australia. Other: device evaluation could not be performed as part# and lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00089. 3002806535 - 2023 - 00090. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16665948&pc=NRA&device_sequence_no=1
3002806535-2023-00083 | - | 03/28/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | M2A-MAGNUM RECAP CUP 50ODX44ID | HIP PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the hip, approximately four (4)years post-op. Attempts have been made and no further information has been provided. | This is a combination product (b)(4). Device was explanted in 2011, specific date is unknown. Concomitant medical products: m2a-magnum mod hd sz 44mm item# 157444 lot# 1322376, magnum tpr adpr ti 42-50/+8mm item# 130832 lot# 1118136. Report source ¿ spain. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2020-00065, 3002806535-2020-00066. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16625969&pc=KWA&device_sequence_no=1
3002806535-2023-00082 | - | 03/27/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN TAPERLOC SHORT STEM | HIP PROSTHESIS | - | It was reported that the patient is experiencing pain, approximately 12 years post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source-foreign- poland. Device evaluation could not be performed as part# and lot# unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16617964&pc=LZO&device_sequence_no=1
3002806535-2023-00080 | 03/03/2023 | 03/24/2023 | 21 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that approximately four years after initial right partial knee surgery, the patient underwent a conversion to a total knee arthroplasty due to poly dislocation, pain, limited range of motion, swelling, clunking noise when getting out of bed, instability of the polyethylene, radiographic evidence of mobile bearing posterior knee with collapse of the medial compartment with metal-on-metal. During the revision, extensive scaring was noted, the poly was found dislocated posteriorly, and all components were revised without complications. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem sm pma item# 161468; lot# 265800. Oxf uni tib tray sz d rm pma item# 154725; lot# 942050. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16608272&pc=NRA&device_sequence_no=1
3002806535-2023-00054 | 01/01/2015 | 03/24/2023 | 3004 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TAPERLOC BMPC 13.5X147MM 12/14 | HIP PROSTHESIS | - | It was reported in a retrospective clinical study that preceding a total hip arthroplasty, infection and an abscess occurred four years¿ post-surgery. Chirurgical drainage without opening the joint capsule was required to resolve the complication. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical device(s): acet. Comp. Rloc spdrcp 52mm/23, item# 1047-52, lot# unknown; batcup arc 10#lr 32i/d t23 rnl, item# 1076-23, lot# unknown; ceramic alu head 12/14 32mm s, item# 1010-35, lot# unknown. Report source ¿ foreign ¿ spain. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Device remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16608078&pc=MRA&device_sequence_no=1
3002806535-2023-00079 | 02/22/2023 | 03/22/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TINBN COATED VANGUARD INTERLOK PS ANATOMIC FEMORAL COMPONENT 65MM LEFT | KNEE PROSTHESIS | - | It was reported that patient has allergic reaction postoperatively, in unknown timeframe after knee surgery. Metallic allergy reaction is suspected. No known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: e1 vngd ps+ tib brg 71/75x10, item#: ep-183740, lot#: 127810; series a pat std 37 3 peg, item#: 184768, lot#: 689980; biomet bone cment r 1x40 jp, item#: 110035372, lot#: z07bac0902; tnbn pol cocr fin tib tray 71, item# :141253tnbn, lot#: 2020091428. Report source: foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16590605&pc=KRR&device_sequence_no=1
3002806535-2023-00078 | 02/23/2023 | 03/22/2023 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI SILICONE ANKLE STRAP | KNEE INSTRUMENT | - | It was reported that, the silicone ankle strap was damaged during surgery. The surgery was completed without delay. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ poland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16589867&pc=LXH&device_sequence_no=1
3002806535-2023-00074 | 02/21/2023 | 03/22/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL | KNEE PROSTHESIS | - | It was reported that the patient underwent a second revision due to for loosening of the implant. Attempts have been made and all additional information received has been included in this report. | (b)(4). Concomitant medical products: unknown oxford femoral; item# unknown; lot# unknown, unknown oxford bearing; item# unknown; lot# unknown. Report source: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. multiple mdr reports were filed for this event, please see the associated reports: 3002806535 - 2023 - 00070, 3002806535 - 2023 - 00072. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16588926&pc=NRA&device_sequence_no=1
3002806535-2023-00070 | 02/21/2023 | 03/22/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL | KNEE PROSTHESIS | - | It was reported that the patient underwent a second revision due to for loosening of the implant. Attempts have been made and all additional information received has been included in this report. | (b)(4). Unknown oxford bearing; item# unknown; lot# unknown. Unknown oxford tibial; item# unknown; lot# unknown. Report source: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. multiple mdr reports were filed for this event, please see the associated reports: 3002806535 - 2023 - 00072. 3002806535 - 2023 - 00074. Part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16588923&pc=NRA&device_sequence_no=1
3002806535-2023-00072 | 02/21/2023 | 03/22/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient underwent a second revision due to for loosening of the implant. Attempts have been made and all additional information received has been included in this report. | (b)(4). Concomitant medical products: unknown oxford femoral; item# unknown; lot# unknown. Unknown oxford tibial; item# unknown; lot# unknown. Report source: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00070 3002806535-2023-00074 part and lot number unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16588922&pc=NRA&device_sequence_no=1
3002806535-2023-00069 | 02/20/2023 | 03/21/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT LG SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery after nearly 10-year post-implantation of right knee implant due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products ¿ associated products; oxf uni tib tray sz e rm pma, item# 154727, lot# 2974499. Oxf twin-peg cmntd fem lg pma, item# 161470, lot# 567250. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00067, 3002806535 -2023 -00068. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583930&pc=NRA&device_sequence_no=1
3002806535-2023-00067 | 02/20/2023 | 03/21/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ E RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery after nearly 10-year post-implantation of right knee implant due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products; oxf twin-peg cmntd fem lg pma, item# 161470 lot# 567250. Oxf anat brg rt lg size 3 pma, item# 159582, lot# 2327803. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00068, 3002806535 -2023 -00069. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583926&pc=NRA&device_sequence_no=1
3002806535-2023-00068 | 02/20/2023 | 03/21/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery after nearly 10-year post-implantation of right knee implant due to unknown reason. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products ¿ associated products; oxf uni tib tray sz e rm pma, item# 154727, lot# 2974499. Oxf anat brg rt lg size 3 pma, item# 159582, lot# 2327803. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2023 -00067, 3002806535 -2023 -00069. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16583374&pc=NRA&device_sequence_no=1
3002806535-2023-00075 | 02/22/2023 | 03/21/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that patient underwent an initial uni knee arthroplasty and subsequently a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: unknown oxford tibial component; item# unknown; lot# unknown. Unknown oxford bearing; item# unknown; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00076, and 3002806535 - 2023 - 00077. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16581438&pc=NRA&device_sequence_no=1
3002806535-2023-00076 | 02/22/2023 | 03/21/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that patient underwent an initial uni knee arthroplasty and subsequently a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford femoral component; item# unknown; lot# unknown; unknown oxford tibial component; item# unknown; lot# unknown. Multiple mdr reports were filed for this event: 3002806535 - 2023 - 00075; 3002806535 - 2023 - 00077. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Device location unknown. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16581428&pc=NRA&device_sequence_no=1
3002806535-2023-00077 | 02/22/2023 | 03/21/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that patient underwent an initial uni knee arthroplasty and subsequently a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: unknown oxford femoral component; item# unknown; lot# unknown, unknown oxford bearing; item# unknown; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00075, 3002806535 - 2023 - 00076. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16581423&pc=NRA&device_sequence_no=1
3002806535-2023-00073 | 10/25/2021 | 03/20/2023 | 511 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a right oxford partial knee replacement and subsequently suffered reactions to it within three days post-operatively. Patient has experienced the following since the operation: bilateral hearing loss, vision loss, thyroid issues, muscle spasms, skin reactions, pain, fatigue, and swelling at the joint. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt md size 4 pma; item#: 159576 lot#: 486170. Unknown oxford tibial component; item#: unknown lot#: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00066. 3002806535 - 2023 - 00071. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16571910&pc=NRA&device_sequence_no=1
3002806535-2023-00071 | 10/25/2021 | 03/20/2023 | 511 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a right oxford partial knee replacement and subsequently suffered reactions to it within three days post-operatively. Patient has experienced the following since the operation: bilateral hearing loss, vision loss, thyroid issues, muscle spasms, skin reactions, pain, fatigue, and swelling at the joint. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt md size 4 pma; item#: 159576 lot#: 486170. Unknown oxford femoral component; item#: unknown lot#: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00066, 3002806535 - 2023 - 00073. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16571794&pc=NRA&device_sequence_no=1
3002806535-2023-00066 | 10/25/2021 | 03/20/2023 | 511 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a right oxford partial knee replacement and subsequently suffered reactions to it within three days post-operatively. Patient has experienced the following since the operation: bilateral hearing loss, vision loss, thyroid issues, muscle spasms, skin reactions, pain, fatigue, and swelling at the joint. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibial component; item#: unknown lot#: unknown. Unknown oxford femoral component; item#: unknown lot#: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00071 and 3002806535-2023-00073 remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16571554&pc=NRA&device_sequence_no=1
3002806535-2023-00065 | 02/22/2023 | 03/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the right knee due to unknown reason, approximately six (6) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz b rm pma item# 154721 lot# 626230. Oxf twin-peg cmntd fem md pma item# 161469 lot# 952430. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00063, 3002806535 - 2023 - 00064. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16563662&pc=NRA&device_sequence_no=1
3002806535-2023-00064 | 02/22/2023 | 03/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the right knee due to unknown reason, approximately six (6) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxf uni tib tray sz b rm pma, item# 154721, lot# 626230. Oxf anat brg rt md size 3 pma, item# 159575, lot# 139450. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00063, and 3002806535-2023-00065. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16563658&pc=NRA&device_sequence_no=1
3002806535-2023-00063 | 02/22/2023 | 03/17/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the right knee due to unknown reason, approximately six (6) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem md pma item# 161469 lot# 952430, oxf anat brg rt md size 3 pma item# 159575 lot# 139450. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00064, 3002806535 - 2023 - 00065. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16563657&pc=NRA&device_sequence_no=1
3002806535-2023-00062 | 02/27/2023 | 03/13/2023 | 14 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD UNI TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due to loosening of the tibial component. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products; oxford uni femoral component, item#: unknown lot#: unknown. Oxford uni bearing component, item#:unknown lot#:unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16529185&pc=NRA&device_sequence_no=1
3002806535-2023-00061 | 02/13/2023 | 03/13/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 9 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent revision of right partial knee implant to total knee implant due to dislocation of anatomical meniscal bearing, approximately 7 weeks post-op. Attempts have been made and no further information has been provided. | (b)(4). Associated products: item reference:(b)(4), item name: oxf uni tib tray sz b rm pma, item lot:528620. Item reference:(b)(4), item name: oxf twin-peg cmntd fem md pma, item lot:494860. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16528992&pc=NRA&device_sequence_no=1
3002806535-2023-00059 | 02/16/2023 | 03/13/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported the surgeon sent a photo of an oxford x-ray with what appeared to be a broken femur. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibial tray; catalog #:unknown lot #:unknown. Unknown oxford bearing; catalog #:unknown lot #:unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16528989&pc=NRA&device_sequence_no=1
3002806535-2023-00060 | 02/07/2023 | 03/10/2023 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that a patient underwent a revision 10 days post hip arthroplasty due to leg length discrepancy. The original stem was inserted further, and the head was exchanged. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: act artic e1 hip brg 28x46mm item# ep-200152; lot# 616890. Report source: foreign ¿ australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16521121&pc=LZO&device_sequence_no=1
3002806535-2023-00045 | 02/13/2023 | 03/10/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNK FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient was revised from a left oxford partial knee to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibia, catalog #:unknown, lot#:unknown; unknown oxford bearing, catalog #:unknown, lot#:unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00043, 3002806535-2023-00044. Investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16520355&pc=NRA&device_sequence_no=1
3002806535-2023-00044 | 02/13/2023 | 03/10/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNK OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient was revised from a left oxford partial knee to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford bearing, catalog #:unknown, lot#:unknown; unknown oxford femoral, catalog #:unknown, lot#:unknown. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535-2023-00043, 3002806535-2023-00045. Investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16520197&pc=NRA&device_sequence_no=1
3002806535-2023-00043 | 02/13/2023 | 03/10/2023 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNK OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient was revised from a left oxford partial knee to a total knee due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibia, catalog #:unknown, lot#:unknown; unknown oxford femoral, catalog #:unknown, lot#:unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00044, 3002806535-2023-00045. Investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16519796&pc=HRY&device_sequence_no=1
3002806535-2023-00056 | 02/06/2023 | 03/06/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | OXFORD CEMENTED FEMORALS | - | It was reported, that the patient underwent a left partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to pain. Contributing factor was disease progression. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical device: oxf uni tib tray sz d lm pma; item# 154724; lot# 478380, oxf anat brg lt lg size 3 pma; item# 159554; lot# 561070. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00057, 3002806535 - 2023 - 00058. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16484013&pc=NRA&device_sequence_no=1
3002806535-2023-00057 | 02/06/2023 | 03/06/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | OXFORD CEMENTED TIBIAL TRAY | - | It was reported, that the patient underwent a left partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to pain. Contributing factor was disease progression. Attempts have been made and no further information has been provided. | (b)(4). Medical device: oxf twin-peg cmntd fem lg pma; item# 161470; lot# 372730. Oxf anat brg lt lg size 3 pma; item# 159554; lot# 561070. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00056, 3002806535 - 2023 - 00058. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16484003&pc=NRA&device_sequence_no=1
3002806535-2023-00058 | 02/06/2023 | 03/06/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT LG SIZE 3 PMA | OXFORD HXLPE BEARINGS | - | It was reported, that the patient underwent a left partial knee arthroplasty and approximately 5 years later a revision surgery was performed due to pain. Contributing factor was disease progression. Attempts have been made and no further information has been provided. | (b)(4). Medical device: oxf twin-peg cmntd fem lg pma; item# 161470; lot# 372730. Oxf uni tib tray sz d lm pma; item# 154724; lot# 478380. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00056 , 3002806535 - 2023 - 00057. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16484002&pc=NRA&device_sequence_no=1
3002806535-2023-00055 | 02/05/2023 | 03/03/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ AA RM PMA | KNEE PROTHESIS | - | It was reported that the patient was revised from an oxford partial knee to a vanguard total knee do to tibial tray loosening approximately eight (8) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). D-10 oxf twin-peg cmntd fem sm pma item# 161468, lot# 514930, oxf anat brg rt sm size 3 pma item# 159568 ,lot# 459490. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16476617&pc=NRA&device_sequence_no=1
3002806535-2023-00053 | 02/13/2023 | 03/01/2023 | 16 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF FEM TRIAL REMOVER | KNEE INSTRUMENTS | - | It was reported that during the procedure, an element on the tip of the instrument was fractured. This event is related to malfunction that could potentially lead to serious injury. due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16458485&pc=LXH&device_sequence_no=1
3002806535-2023-00049 | 02/01/2023 | 02/28/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported, that: the patient underwent a revision due to dislocation, approximately 5 years 7 months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item#:154725 ;lot#:337110 ;item name: oxf uni tib tray sz d rm pma, item#:161469 ;lot#:737850 ;item name: oxf twin-peg cmntd fem md pma. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16450903&pc=NRA&device_sequence_no=1
3002806535-2023-00051 | 02/02/2023 | 02/28/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported that a patient underwent an initial left oxford pks, and approximately 5 years later a revision surgery was performed due to loosening of the implants. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem md pma, item# 161469, lot# 421950. Oxf anat brg lt md size 4 pma, item# 159548, lot# 847380. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00050 and 3002806535-2023-00052. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16450760&pc=NRA&device_sequence_no=1
3002806535-2023-00052 | 02/02/2023 | 02/28/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | OXFORD HXLPE BEARINGS | - | It was reported that a patient underwent an initial left oxford pks, and approximately 5 years later a revision surgery was performed due to loosening of the implants. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: oxf twin-peg cmntd fem md pma; item# 161469; lot# 421950. Oxf uni tib tray sz c lm pma; item# 154722; lot# 140470. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00050, and 3002806535 - 2023 - 00051. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16450759&pc=NRA&device_sequence_no=1
3002806535-2023-00050 | 02/02/2023 | 02/28/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | OXFORD CEMENTED FEMORALS | - | It was reported that a patient underwent an initial left oxford pks, and approximately 5 years later a revision surgery was performed due to loosening of the implants. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: oxf uni tib tray sz c lm pma; item# 154722; lot# 140470, oxf anat brg lt md size 4 pma; item# 159548; lot# 847380. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00051, 3002806535-2023-00052. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16450757&pc=NRA&device_sequence_no=1
3002806535-2023-00047 | 06/29/2021 | 02/27/2023 | 608 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OPTION TAPER SLEEVE, STD NECK, TYPE 1 TAPER | HIP PROSTHESIS | - | It was reported that a patient had a two-stage revision due to metal related pathology about two (2) weeks after initial surgery. Subsequently underwent an i&d due to a fungal and bacterial infection. The head, taper, and bearing were exchanged without complication. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - biolox delta option head, 28mm, item# 650-1055, lot# 3061992; vivacit-e dm bearing 28x50mm, item# 110031015, lot# 64760199. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00046. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16442314&pc=LZO&device_sequence_no=1
3002806535-2023-00046 | 06/29/2021 | 02/27/2023 | 608 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA OPTION HEAD, 28MM | HIP PROSTHESIS | - | It was reported that a patient had a two-stage revision due to metal related pathology about two (2) weeks after initial surgery. Subsequently underwent an i&d due to a fungal and bacterial infection. The head, taper, and bearing were exchanged without complication. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: option taper sleeve, std neck, type 1 taper, item# 650-1066, lot# 3048765; vivacit-e dm bearing 28x50mm, item# 110031015, lot# 64760199. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00047. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16442312&pc=LZO&device_sequence_no=1
3002806535-2023-00048 | - | 02/27/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported the inlay (bearing) fractured after cementless knee implantation in 2008, without adequate trauma. No remedial action has been reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16442278&pc=NRA&device_sequence_no=1
3002806535-2023-00037 | 01/18/2022 | 02/27/2023 | 405 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported, that the patient underwent a bearing revision due to dislocation, approximately 3 months post-op. The patient dislocated the bearing anteriorly during sleep. Revised for a thicker bearing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz e rm pma item # 154727, lot # 903100, oxf twin-peg cmntd fem md pma # 161469, lot # 716400. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16441979&pc=NRA&device_sequence_no=1
3002806535-2023-00042 | - | 02/24/2023 | - | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANTI-IMPINGEMENT GD MED | BMET GENERAL KNEE INSTRUMENTS | - | It was reported, that the mill guide instrument fractured during the surgery. This event is related to malfunction that could potentially lead to serious injury. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- foreign: poland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16434132&pc=LXH&device_sequence_no=1
3002806535-2023-00041 | 01/19/2023 | 02/14/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL RIGHT SIZE 6 | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision due to dislocation about one (1) year and two (2) months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Oxford pks twin peg femur - small, item# 161468, lot# unknown; oxford pks cemented tibia, alpha, standard - rm size c, item# 154723, lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16370677&pc=NRA&device_sequence_no=1
3002806535-2023-00038 | 01/31/2023 | 02/14/2023 | 14 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK | KNEE INSTRUMENT | - | It was reported that while using the t-handle to remove 0 spigot from patient during bilateral unicondular knee replacement, peg for insertion into the spigot sheared off at the connection point to the t-handle. No impact on patient was noted. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign - canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16370100&pc=HRY&device_sequence_no=1
3002806535-2023-00023 | 12/28/2022 | 02/13/2023 | 47 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that an oxford pks was revised to a total knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt lg size 4 pma item # 159583, lot # 886550, oxf twin-peg cmntd fem lg pma item # 161470, lot # 530620. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00021, 3002806535-2023-00022. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16361523&pc=NRA&device_sequence_no=1
3002806535-2023-00022 | 12/28/2022 | 02/13/2023 | 47 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT LG SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that an oxford pks was revised to a total knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products- oxf twin-peg cmntd fem lg pma item # 161470, lot # 530620 oxf uni tib tray sz c rm pma item # 154723, lot # 730530. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00021, 3002806535-2023-00023. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16361507&pc=NRA&device_sequence_no=1
3002806535-2023-00040 | 01/13/2023 | 02/13/2023 | 31 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | INTRAMEDULLARY ROD LINK | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that: the hook for removing the im rod broke off during the operation, nothing remained in the patient. No consequences or impact to the patient. At the same time, a tip of the connecting piece broke off during the operation. Nothing remained in the patient and could be removed from the im. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products : intramedullary rod removal hook with t-handle item# 32-401111, lot# unknown. Report source: foreign : germany. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00039. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16361020&pc=JWH&device_sequence_no=1
3002806535-2023-00039 | 01/13/2023 | 02/13/2023 | 31 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | INTRAMEDULLARY ROD REMOVAL HOOK WITH T-HANDLE | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that: the hook for removing the im rod broke off during the operation, nothing remained in the patient. No consequences or impact to the patient. At the same time, a tip of the connecting piece broke off during the operation. Nothing remained in the patient and could be removed from the im. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products : intramedullary rod link, item# 32-422822, lot# unknown. Report source : foreign : germany. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00040. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16361018&pc=JWH&device_sequence_no=1
3002806535-2023-00035 | 01/06/2023 | 02/03/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of partial knee prosthesis due to unknown reasons, approximately 9 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: associated products: item reference:161469, item name:oxf twin-peg cmntd fem md pma, lot number:441890. Item reference:159575, item name:oxf anat brg rt md size 3 pma, lot number:331530. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00029, 3002806535-2023-00036. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16289003&pc=NRA&device_sequence_no=1
3002806535-2023-00036 | 01/06/2023 | 02/03/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of partial knee prosthesis due to unknown reasons, approximately 9 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products:¿ item reference:154723, item name:oxf uni tib tray sz c rm pma, lot number:001020, item reference:161469, item name:oxf twin-peg cmntd fem md pma, lot number:441890. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00029, 3002806535-2023-00035. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16288998&pc=NRA&device_sequence_no=1
3002806535-2023-00029 | 01/06/2023 | 02/03/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of partial knee prosthesis due to unknown reasons, approximately 9 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: associated products: item reference:154723, item name:oxf uni tib tray sz c rm pma, lot number:001020. Item reference:159575, item name:oxf anat brg rt md size 3 pma, lot number:331530. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00035, 3002806535-2023-00036. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16288962&pc=NRA&device_sequence_no=1
3002806535-2023-00020 | 03/15/2005 | 01/31/2023 | 6531 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | MH SOLID SHELL HA/PC 56MM LN24 | HIP PROTHESIS | - | It was reported that the patient underwent revision of right hip due to loosening and dislocation, approximately one (1) month from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: australia. D-10 28mm mod hd std neck tp1 taper item#163662 lot#792660 modular head component 28mm item#11105914 lot#792660. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16271128&pc=LZO&device_sequence_no=1
3002806535-2023-00034 | 01/11/2023 | 01/31/2023 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the left knee due to aseptic loosening and dislocation of mobile bearing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products : oxf uni tib tray sz b lm pma item# 154720 lot# 960970; oxf anat brg lt sm size 7 pma item# 159544 lot# 299900. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00025 and 3002806535 - 2023 - 00033. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16270867&pc=NRA&device_sequence_no=1
3002806535-2023-00025 | 01/11/2023 | 01/31/2023 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the left knee due to aseptic loosening and dislocation of mobile bearing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf anat brg lt sm size 7 pma item# 159544 lot# 299900, oxf twin-peg cmntd fem sm pma item# 161468 lot# 591120. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00025, 3002806535 - 2023 - 00033, 3002806535 - 2023 - 00034. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16270866&pc=NRA&device_sequence_no=1
3002806535-2023-00033 | 01/11/2023 | 01/31/2023 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 7 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the left knee due to aseptic loosening and dislocation of mobile bearing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz b lm pma item# 154720 lot# 960970. Oxf twin-peg cmntd fem sm pma item# 161468 lot# 591120. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00025. 3002806535-2023-00034. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16270863&pc=NRA&device_sequence_no=1
3002806535-2023-00030 | 01/05/2023 | 01/31/2023 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 28MM STD NK | HIP PROSTHESIS | - | It was reported that: patient underwent hip revision surgery due to the dislocation about three (3) years after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - act artic e1 hip brg 28x44mm, item# ep-200150, lot# 530940; m2a-magnum pf cup 50odx44id, item# us157850, lot# 249530; echo b-mtrc mp fp so 9, item# 193009, lot# 681140. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16270061&pc=LZO&device_sequence_no=1
3002806535-2023-00026 | 01/09/2023 | 01/27/2023 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent knee revision surgery due to unknown reason, approximately four (4) years three (3) months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem md pma item#161469 lot#036750 oxf anat brg lt md size 3 pma item#159547 lot#826020. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00027, 3002806535-2023-00028. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16251649&pc=NRA&device_sequence_no=1
3002806535-2023-00027 | 01/09/2023 | 01/27/2023 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROTHESIS | - | It was reported that the patient underwent knee revision surgery due to unknown reason, approximately four (4) years three (3) months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz b lm pma item#154720 lot#690780. Oxf anat brg lt md size 3 pma item#159547 lot#826020. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00026, 3002806535-2023-00028. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16251648&pc=NRA&device_sequence_no=1
3002806535-2023-00028 | 01/09/2023 | 01/27/2023 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent knee revision surgery due to unknown reason, approximately four (4) years three (3) months from initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz b lm pma item#154720 lot#690780. Oxf twin-peg cmntd fem md pma item#161469 lot#036750. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00026, 3002806535-2023-00027. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16251646&pc=NRA&device_sequence_no=1
3002806535-2023-00021 | 12/28/2022 | 01/25/2023 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | It was reported that an oxford pks was revised to a total knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt lg size 4, pma item#:159583, lot#: 886550. Oxf uni tib tray sz c rm, pma item#: 154723, lot#: 730530. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00022 and 3002806535-2023-00023. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16241394&pc=NRA&device_sequence_no=1
3002806535-2023-00024 | 01/04/2023 | 01/25/2023 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK | KNEE INSTRUMENT | - | It was reported that the hook part of the handle broke during the case. The patient did not retain any parts and nothing fell into the patient. No harm to the patient and no delay. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16239768&pc=LXH&device_sequence_no=1
3002806535-2023-00019 | 01/10/2023 | 01/24/2023 | 14 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent an initial oxford pks, and approximately 10 months post op underwent a revision due to poly spin out (dislocation). Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16231616&pc=NRA&device_sequence_no=1
3002806535-2023-00017 | 12/29/2022 | 01/17/2023 | 19 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 5 PMA | KNEE PROTHESIS | - | It was reported that patient underwent right knee revision surgery due to luxation, approximately five (5) years and nine (9) months from initial implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source foreign: germany. Concomitant medical products: oxford uni femoral sm item#154600 lot#j3877349; oxf uni tib tray sz c rm pma item#154723 lot#3922418. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16181905&pc=NRA&device_sequence_no=1
3002806535-2023-00016 | 12/20/2022 | 01/12/2023 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNCEMENTED TIBIAL IMPACTOR | NON MODULAR TIBIAL TRAY IMPACTOR (CEMENTLESS APPLICATION) | - | It was reported that upon completion of the knee arthroplasty, post operative x-rays showed a fragment of the implant inserter was retained in the patient. No health impact to the patient has been reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16144640&pc=MDM&device_sequence_no=1
3002806535-2023-00015 | 09/29/2021 | 01/11/2023 | 469 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | VANGUARD PS POROUS COATED/HA FEMORAL COMPONENT 62.5MM RIGHT | KNEE PROSTHESIS | - | It was reported that there was a suspicion of an allergic reaction. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: biomet finned pri stem 40mm item#141314 lot#000580. Biomet por pri tib tray 71mm item#141263 lot#021030. Series a pat std 31 3 peg item#184764 lot#127140. Vngd ps tib brg 10x71/75mm item#183640 lot#772280. Report source- france. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. . | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16137030&pc=MBH&device_sequence_no=1
3002806535-2023-00011 | 12/13/2022 | 01/09/2023 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPT TYPE 1 TPR SLEVE 0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was found out that taper sleeve did not match with ceramic head during hip surgery. It was reported that this was due to the +6mm trial head, and the +6mm sleeves not recreating the same construct. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | This is a combination product (b)(4). Report source foreign: australia. Concomitant medical product(s): cer option type 1 tpr sleve +6 item#650-1068 lot#2911080. Cer option type 1 tpr sleve +6 item#650-1068 lot#3066236. +6mm mod hd endo ii trial item#31-401161 lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00009, and 3002806535-2023-00010. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120217&pc=LZO&device_sequence_no=1
3002806535-2023-00009 | 12/13/2022 | 01/09/2023 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPTION TYPE 1 TPR SLEVE +6 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was found that taper sleeve did not match with ceramic head during hip surgery. It was reported that this was due to the +6mm trial head, and the +6mm sleeves not recreating the same construct. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: australia. Concomitant medical products: cer option type 1 tpr sleve +6 item# 650-1068, lot#3066236. Cer opt type 1 tpr sleve 0mm item# 650-1066, lot#2020020470. +6mm mod hd endo ii trial item# 31-401161, lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00010 & 3002806535-2023-00011. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120215&pc=LZO&device_sequence_no=1
3002806535-2023-00010 | 12/13/2022 | 01/09/2023 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPTION TYPE 1 TPR SLEVE +6 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | - | It was found that taper sleeve did not match with ceramic head during hip surgery. It was reported that this was due to the +6mm trial head, and the +6mm sleeves not recreating the same construct. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: australia. Concomitant medical products: cer option type 1 tpr sleve +6 item#650-1068 lot#2911080. Cer opt type 1 tpr sleve 0mm item#650-1066 lot#2020020470. +6mm mod hd endo ii trial item#31-401161 lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00009, 3002806535-2023-00011. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120213&pc=LZO&device_sequence_no=1
3002806535-2023-00014 | 12/29/2022 | 01/09/2023 | 11 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery of a partial knee due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item number:159547, item name: oxf anat brg lt md size 3 pma, lot number: 618300; item number:161469, item name: oxf twin-peg cmntd fem md pma a, lot number: 214390. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00013; 3002806535-2023-00014. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120155&pc=NRA&device_sequence_no=1
3002806535-2023-00013 | 12/29/2022 | 01/09/2023 | 11 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery of a partial knee due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item number:159547, item name: oxf anat brg lt md size 3 pma, lot number: 618300; item number:154722, item name: oxf uni tib tray sz c lm pma, lot number: 394270. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00012; 3002806535-2023-00014. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120149&pc=NRA&device_sequence_no=1
3002806535-2023-00012 | 12/29/2022 | 01/09/2023 | 11 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery of a partial knee due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: associated products: item number:161469, item name: oxf twin-peg cmntd fem md pma, lot number:214390;item number:154722, item name: oxf uni tib tray sz c lm pma, lot number:394270. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2023-00013; 3002806535-2023-00014. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16120144&pc=NRA&device_sequence_no=1
3002806535-2022-00519 | 12/13/2022 | 01/09/2023 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNI TWIN-PEG FEMORAL LG | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the knee due to implant fracture. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign - austria. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16119803&pc=NRA&device_sequence_no=1
3002806535-2023-00008 | 11/24/2022 | 01/05/2023 | 42 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM LEFT SIZE 5 | KNEE PROSTHESIS | - | It was reported that patient underwent left knee revision approximately fifteen (15) years post initial surgery due to the pain which was caused by the metal marker embedded in the bearing having detached into the joint. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: france. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16101881&pc=NRA&device_sequence_no=1
3002806535-2023-00005 | - | 01/04/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent diagnostic intervention (biopsy for metal allergy) greater than six months from initial surgery. The patient also experiences pain and swelling. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf twin-peg cmntd fem md pma ; item#: 161469; lot#: unknown; item name: oxf anat brg lt md size 4 pma , item#: 159548 ; lot#: unknown; item name: refobacin bc r 1x40 us; item#: 110034355; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00003; 3002806535 - 2023 - 00004. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16093128&pc=NRA&device_sequence_no=1
3002806535-2023-00004 | - | 01/04/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent diagnostic intervention (biopsy for metal allergy) greater than six months from initial surgery. The patient also experiences pain and swelling. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf anat brg lt md size 4 pma; item#: 159548 ; lot#: unknown, item name: oxf uni tib tray sz d lm pma , item#: 154724; lot#: unknown, item name: refobacin bc r 1x40 us; item#: 110034355; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00003, 3002806535 - 2023 - 00005. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16093125&pc=NRA&device_sequence_no=1
3002806535-2023-00003 | - | 01/03/2023 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent diagnostic intervention (biopsy for metal allergy) greater than six months from initial surgery. The patient also experiences pain and swelling. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf twin-peg cmntd fem md pma ; item#: 161469; lot#: unknown; item name: oxf uni tib tray sz d lm pma, item#: 154724; lot#: unknown; item name: refobacin bc r 1x40 us; item#: 110034355; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated report: 0001822565 - 2023 - 00004. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16086012&pc=NRA&device_sequence_no=1
3002806535-2022-00508 | 12/05/2022 | 01/03/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ AA LM PMA | KNEE PROSTHESIS | - | It was reported that over 7 and a half years post implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf anat brg lt md size 5 pma; item#: 159549; lot#: unknown, item name: oxf twin-peg cmntd fem md pma, item#: 161469 lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00507, 3002806535 - 2023 - 00002. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16085468&pc=NRA&device_sequence_no=1
3002806535-2023-00002 | 12/05/2022 | 01/03/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that over 7 and a half years post implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf uni tib tray sz aa lm pma; item#: 159531; lot#: unknown, item name: oxf twin-peg cmntd fem md pma, item#: 161469 lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00507, 3002806535 - 2022 - 00508. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16085467&pc=NRA&device_sequence_no=1
3002806535-2022-00507 | 12/05/2022 | 01/03/2023 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that over 7 and a half years post implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: oxf uni tib tray sz aa lm pma; item#: 159531; lot#: unknown, item name: oxf anat brg lt md size 5 pma, item#: 159549; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2023 - 00002, 3002806535 - 2022 - 00508. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16085466&pc=NRA&device_sequence_no=1
3002806535-2023-00001 | 12/14/2022 | 01/03/2023 | 20 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENT | - | It was reported that during knee prosthesis implantation it was noticed that one of the lugs on tibial inserter is missing and therefore its no longer holding the tibial plate. The surgeon inserted it without the inserter. Attempts have been made and no further information has been provided. | (b)(4). Report source¿ foreign ¿ germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16085002&pc=NRA&device_sequence_no=1
3002806535-2022-00518 | 11/10/2022 | 12/29/2022 | 49 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNI TWIN-PEG FEMORAL MD | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery approximately seven (7) years post-implantation due to implant fracture of the femoral component. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: germany. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16067163&pc=NRA&device_sequence_no=1
3002806535-2022-00516 | 12/06/2022 | 12/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ AA LM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision surgery due to unknown reasons, approximately eight (8) months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg lt sm, size 4, pma, item#: 159541, lot#: 811030, oxf twin-peg cmntd fem sm, pma, item#: 161468, lot#: 947260. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00514, 3002806535-2022-00515. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16060017&pc=NRA&device_sequence_no=1
3002806535-2022-00514 | 12/06/2022 | 12/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 4 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision surgery due to unknown reasons, approximately eight (8) months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: product id: oxf twin-peg cmntd fem sm pma, item#: 161468, lot#: 947260; product id: oxf uni tib tray sz aa lm pma, item#: 159531, lot#: 297320. Multiple mdr reports were filed for this event: 3002806535-2022-00515 and 3002806535-2022-00516. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16060014&pc=NRA&device_sequence_no=1
3002806535-2022-00515 | 12/06/2022 | 12/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision surgery due to unknown reasons, approximately eight (8) months after initial surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg lt sm size 4 pma item#159541 lot#811030, oxf uni tib tray sz aa lm pma item#159531 lot#297320. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00514, 3002806535-2022-00516. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16060013&pc=NRA&device_sequence_no=1
3002806535-2022-00517 | 12/07/2022 | 12/28/2022 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE D | KNEE PROSTHESIS | - | It was reported that during knee surgery, the nurse opened the sterile outer package and found signs of wear on the base plate of the unicondylar tibial plateau. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ china. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16058945&pc=NRA&device_sequence_no=1
3002806535-2022-00506 | 11/28/2022 | 12/22/2022 | 24 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 4 PMA | UNICONDYLAR KNEE PROSTHESIS | - | It was reported that a patient experienced painful knee twisting trauma which resulted in bearing dislocation. A revision surgery took place where only the bearing was replaced, approximately three years after initial implantation. The bearing was reported to be worn. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Foreign source: germany. Oxf uni cmntls tib sz d rm item#166577 lot#6195770, oxford ph3 cementless fem sz m item#154926 lot#6459441. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16032524&pc=NRA&device_sequence_no=1
3002806535-2022-00513 | 09/02/2021 | 12/20/2022 | 474 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 5 PMA | OXFORD HXLPE BEARINGS | - | It was reported that expired product was implanted. No consequences or impact to the patient. Attempts have been made and no further information has been provided. | (b)(4). Report source: foreign: mexico. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16012568&pc=NRA&device_sequence_no=1
3002806535-2022-00512 | 12/06/2022 | 12/19/2022 | 13 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | PROSTHESIS, HIP | - | It was reported that patient underwent revision due to ceramic head fracture, 2 weeks post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item reference:30103605, item name:g7 vit e neutral lnr 36mm e, lot number: 65516876. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16003243&pc=LZO&device_sequence_no=1
3002806535-2022-00510 | 11/29/2022 | 12/19/2022 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OOXF UNI TIB TRAY SZ E RM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to the unknown reason, approximately 4 years 9 months post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - oxford pks twin peg femur - large; item# 161470, lot# 633310; oxford pks anatomic arcom meniscal brg - large right size 7; item# 159586, lot# 845260. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00509; 3002806535-2022-00511. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16003104&pc=NRA&device_sequence_no=1
3002806535-2022-00509 | 11/29/2022 | 12/19/2022 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to the unknown reason, approximately 4 years 9 months post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - oxford pks cemented tibia, alpha, standard - rm size e; item# 154727, lot# 725850; oxford pks anatomic arcom meniscal brg - large right size 7; item# 159586, lot# 845260. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00510; 3002806535-2022-00511. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16003099&pc=NRA&device_sequence_no=1
3002806535-2022-00511 | 11/29/2022 | 12/19/2022 | 20 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT LG SIZE 7 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent knee revision surgery due to the unknown reason, approximately 4 years 9 months post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxford pks twin peg femur - large; item#: 161470, lot#: 633310. Oxford pks cemented tibia, alpha, standard - rm size e; item#: 154727, lot#: 725850. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00509, 3002806535-2022-00510. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16003094&pc=NRA&device_sequence_no=1
3002806535-2022-00483 | 11/24/2022 | 12/19/2022 | 25 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, METAL/CERAMIC/POLYMER | - | It was reported that patient had an initial right total hip arthroplasty and approximately 3 weeks post implantation underwent manipulation under anesthesia due to subluxation. Due diligence is in progress for this complaint; to date no additional on the reported event is unavailable at this time. | (b)(4). Concomitant medical products: g7 pps ltd acet shell 62h, item# 010000668; lot# r7313004a. G7 vit e neutral lnr 36mm h, item# 30103608; lot#65454505. Tprlc 133 t1 pps so 15x150mm, item# 51-103150; lot#7198025. Report source ¿ foreign ¿ netherlands. Customer has indicated that the product will not be returned to zimmer biomet for investigation. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Device evaluated by manufacturer: remains implanted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=16002548&pc=LZO&device_sequence_no=1
3002806535-2022-00504 | 11/16/2022 | 12/16/2022 | 30 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENT | - | It was reported that during surgery, while impacting the tibial implant, a piece of the instrument broke off. Nothing left in the patient. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source ¿ foreign ¿ canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15994423&pc=HRY&device_sequence_no=1
3002806535-2022-00505 | 12/05/2022 | 12/15/2022 | 10 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM LEFT SIZE 3 | - | It was reported that the patient underwent revision surgery due to pain and dislocation, approximately 3 years post-implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxford uni twin-peg femoral md cocr med item#166942 lot#j6466696; oxf uni tib tray sz b lm pma item#154720 lot#6576513. Report source - japan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15985850&pc=NRA&device_sequence_no=1
3002806535-2022-00503 | 11/21/2022 | 12/15/2022 | 24 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision from partial to total knee prosthesis due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item number: 159540 , item name:oxf anat brg lt sm size 3 pma, lot number #:711080. Item number: 154722 , item name: oxf uni tib tray sz c lm pma, lot number #:732320. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00501, 3002806535-2022-00502. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15984916&pc=NRA&device_sequence_no=1
3002806535-2022-00502 | 11/21/2022 | 12/15/2022 | 24 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision from partial to total knee prosthesis due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item number: 159540 , item name:oxf anat brg lt sm size 3 pma, lot number #:711080. Item number: 161468 , item name:oxf twin-peg cmntd fem sm pmaa, lot number #:079700. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00501, 3002806535-2022-00503. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15984910&pc=NRA&device_sequence_no=1
3002806535-2022-00501 | 11/21/2022 | 12/14/2022 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision from partial to total knee prosthesis due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: associated products: item number: 154722 , item name: oxf uni tib tray sz c lm pma, lot number:732320, item number: 161468 , item name:oxf twin-peg cmntd fem sm pmaa, lot number:079700. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00502, 3002806535-2022-00503. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15975335&pc=NRA&device_sequence_no=1
3002806535-2022-00500 | 10/15/2022 | 12/13/2022 | 59 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT X-SM 3MM PMA | OXFORD HXLPE BEARINGS | - | It was reported that the patient underwent an initial right unicondylar knee replacement. Approximately six months later revision surgery was performed due to dislocation of the bearing. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: oxford uni twin-peg femoral xs; item# 166940; lot# 6971292; oxf uni tib tray sza rm; item# 154719; lot# 7006094. Report source - foreign: china. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Discarded. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15967365&pc=NRA&device_sequence_no=1
3002806535-2022-00499 | 11/13/2022 | 12/12/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 36MM G | HIP PROSTHESIS | - | It was reported that approximately one month post-implantation, the patient developed an infection for unknown reasons. The infected area was flushed with antibiotics. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: delta ceramic fem hd 36/+3mm, item#650-0662, lot#3100098. G7 pps ltd acet shell 58g, item#010000666, lot#7096119. Trilogy bone scr 6. 5x35, item#00-6250-065-35, lot# 65292812. Trilogy bone scr 6. 5x20, item#00-6250-065-20, lot#65135483. Trilogy bone scr 6. 5x20, item#00-6250-065-20, lot#64700874. Report source: china. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00492. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958392&pc=LZO&device_sequence_no=1
3002806535-2022-00495 | 11/15/2022 | 12/12/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL RIGHT SIZE 3 | - | It was reported that the patient was revised from an oxford uni to a full knee due to patella pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00493, 3002806535 - 2022 - 00494. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958130&pc=NRA&device_sequence_no=1
3002806535-2022-00493 | 11/15/2022 | 12/12/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | OXFORD PKS TWIN PEG FEMUR - SMALL | - | It was reported that the patient was revised from an oxford uni to a full knee due to patella pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00494; 3002806535 - 2022 - 00495. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958125&pc=NRA&device_sequence_no=1
3002806535-2022-00494 | 11/15/2022 | 12/12/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - RM SIZE B | - | It was reported that the patient was revised from an oxford uni to a full knee due to patella pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00493; 3002806535 - 2022 - 00495. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15958110&pc=NRA&device_sequence_no=1
3002806535-2022-00492 | 11/13/2022 | 12/12/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+3MM | PROSTHESIS, HIP | - | It was reported that approximately one month post-implantation, the patient developed an infection for unknown reasons. The infected area was flushed with antibiotics. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: g7 pps ltd acet shell 58g, item#010000666, lot#7096119; trilogy bone scr 6. 5x35, item#00-6250-065-35, lot# 65292812; trilogy bone scr 6. 5x20, item#00-6250-065-20, lot#65135483; trilogy bone scr 6. 5x20, item#00-6250-065-20, lot#64700874. Report source ¿ china. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15957905&pc=LZO&device_sequence_no=1
3002806535-2022-00491 | 11/21/2022 | 12/12/2022 | 21 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 5 PMA | OXFORD HXLPE BEARINGS | - | It was reported that the patient underwent an initial partial knee replacement. Approximately 2 months later a revision was performed due to dislocation of the poly bearing. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15957577&pc=NRA&device_sequence_no=1
3002806535-2022-00480 | 11/15/2022 | 12/08/2022 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that six and a half years post initial partial left knee replacement the patient underwent revision surgery due to bearing dislocation and implant fracture. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: australia. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15938779&pc=NRA&device_sequence_no=1
3002806535-2022-00485 | 11/09/2022 | 12/07/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT LG SIZE 4 PMA | OXFORD HXLPE BEARINGS | - | It was reported that patient underwent a partial knee arthroplasty and approximately 8 years later a revision surgery was performed due to metallosis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz c lm pma; item# 154722; lot# 813810. Oxf twin-peg cmntd fem lg pma; item# 161470; lot# 816300. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00486, 3002806535 - 2022 - 00487. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15930633&pc=NRA&device_sequence_no=1
3002806535-2022-00487 | 11/09/2022 | 12/07/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported that patient underwent a partial knee arthroplasty and approximately 8 years later a revision surgery was performed due to metallosis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg lt lg size 4 pma; item# 159555; lot# 948510, oxf twin-peg cmntd fem lg pma; item# 161470; lot# 816300. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00485, 3002806535 - 2022 - 00486. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15930627&pc=NRA&device_sequence_no=1
3002806535-2022-00486 | 11/09/2022 | 12/07/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | OXFORD CEMENTED FEMORALS | - | It was reported that patient underwent a partial knee arthroplasty and approximately 8 years later a revision surgery was performed due to metallosis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf anat brg lt lg size 4 pma; item# 159555; lot# 948510; oxf uni tib tray sz c lm pma; item# 154722; lot# 813810. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00485; 3002806535 - 2022 - 00487. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15930625&pc=NRA&device_sequence_no=1
3002806535-2022-00489 | 09/14/2022 | 12/05/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OPTION TAPER SLEEVE, -3MM NECK, TYPE 1 TAPER | PROSTHESIS, HIP | - | It was reported that the patient was revised 23-days post-implantation, due to a superficial infection. Attempts have been made and no further information has been provided. | (b)(4). Medical products; biolox delta option head, 40mm, item# 650-1058, lot# 3041313. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00488. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15914603&pc=LZO&device_sequence_no=1
3002806535-2022-00488 | 09/14/2022 | 12/05/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA OPTION HEAD, 40MM | PROSTHESIS, HIP | - | It was reported that the patient was revised 23-days post-implantation, due to a superficial infection. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: option taper sleeve, -3mm neck, type 1 taper, item# (b)(4), lot# 3087469. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00489. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15914596&pc=LZO&device_sequence_no=1
3002806535-2022-00484 | 11/07/2022 | 12/05/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENT | - | It was reported that the instrument fractured during surgery. The broken piece remained in the patient. Additional surgery to remove the instrument's broken piece was performed the next day. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15913816&pc=LXH&device_sequence_no=1
3002806535-2022-00481 | - | 11/30/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 40MM | PROSTHESIS, HIP | - | It was reported that patient underwent revision due to head and sleeve being lost in patient body during initial surgery. It was concluded that product may have traveled up into the abdomen. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: associated products: item number:51-107220,item name: tprlc 133 mp type1 pps ho 22. 0, lot number:3616830, item number:110017109,item name:g7 finned 4 hole shell 64h, lot number:7115768, item number:98-b001-009-41,item name:unknown product, lot number: unknown, item number:30104008, item name:g7 vit e neutral lnr 40mm h, lot number:65100260, item number:650-1058, item name:cer bioloxd option hd 40mm, lot number:3096108, item number:650-1067, item name:cer option type 1 tpr sleve +3, lot number:3035290, item number:650-1066, item name:cer opt type 1 tpr sleve 0mm, lot number:3048726. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00482. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15884337&pc=LZO&device_sequence_no=1
3002806535-2022-00482 | - | 11/30/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPT TYPE 1 TPR SLEVE 0MM | PROSTHESIS, HIP | - | It was reported that patient underwent revision due to head and sleeve being lost in patient body during initial surgery. It was concluded that product may have traveled up into the abdomen. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: associated products: item number:51-107220,item name: tprlc 133 mp type1 pps ho 22. 0, lot number:3616830, item number:110017109,item name:g7 finned 4 hole shell 64h, lot number:7115768, item number:98-b001-009-41,item name:unknown product, lot number: unknown, item number:30104008, item name:g7 vit e neutral lnr 40mm h, lot number:65100260, item number:650-1058, item name:cer bioloxd option hd 40mm, lot number:3096108, item number:650-1067, item name:cer option type 1 tpr sleve +3, lot number:3035290, item number:650-1058, item name: cer bioloxd optionhd 40mm, lot number:3095934. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00481. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15884334&pc=LZO&device_sequence_no=1
3002806535-2022-00479 | 11/07/2022 | 11/30/2022 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/+3MM T1 | HIP PROSTHESIS | - | It was reported that patient underwent revision of hip prosthesis due to ceramic head fracture, approximately 8 years post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source - foreign: united kingdom. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15884331&pc=LZO&device_sequence_no=1
3002806535-2022-00458 | 11/01/2022 | 11/28/2022 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENT | - | It was reported that during surgery a part of the instrument broke off. Nothing fell into the patient and the surgery was completed without any delay. No health consequences to patient were reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source - (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15868568&pc=LXH&device_sequence_no=1
3002806535-2022-00478 | 10/28/2022 | 11/25/2022 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD 4MM DRILL | KNEE INSTRUMENTS | - | It was reported that the drill fractured during surgery. This event is related to a malfunction that could potentially lead to a sterility issue. Due diligence is in progress for this complaint; to date, whatever additional information received has been included in this report. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15864787&pc=HTW&device_sequence_no=1
3002806535-2022-00477 | 10/27/2022 | 11/25/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 36MM F | HIP PROSTHESIS | - | It was reported that during initial surgery, the ceramic liner fractured during impaction. The liner could not be removed from the cup/shell, so the surgeon removed both components resulting in prolonged surgery. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: g7 bonemaster ltd acet shl 54f, item# 010000704, lot# 7228681. Foreign source: belgium. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15864069&pc=LZO&device_sequence_no=1
3002806535-2022-00471 | 10/25/2022 | 11/23/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ E LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a knee arthroplasty and approximately five years later a revision surgery was performed due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf twin-peg cmntd fem lg pma; item# 161470; lot# unknown, oxf anat brg lt lg size 4 pma; item# 159555; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00474, 3002806535-2022-00475. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15850914&pc=NRA&device_sequence_no=1
3002806535-2022-00475 | 10/25/2022 | 11/23/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT LG SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee arthroplasty and approximately five years later a revision surgery was performed due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf uni tib tray sz e lm pma; item# 154726; lot#: unknown, oxf twin-peg cmntd fem lg pma; item# 161470; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00471, 3002806535-2022-00474. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15850880&pc=NRA&device_sequence_no=1
3002806535-2022-00474 | 10/25/2022 | 11/23/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee arthroplasty and approximately five years later a revision surgery was performed due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz e lm pma; item# 154726; lot#: unknown; oxf anat brg lt lg size 4 pma; item# 159555; lot# unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00471; 3002806535-2022-00475. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15850879&pc=NRA&device_sequence_no=1
3002806535-2022-00476 | 11/14/2022 | 11/23/2022 | 9 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/+3MM T1 | HIP PROSTHESIS | - | It was reported that 2-week post-implantation, the patient underwent revision surgery due to implant fracture. It was proved by the x-ray taken during post op follow up visit. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15850550&pc=LZO&device_sequence_no=1
3002806535-2022-00470 | 10/31/2022 | 11/22/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision due to pain, approximately nine (9) years three (3) months after initial surgery. Total knee conversion was performed during the revision. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf uni tib tray sz c rm pma item#154723 lot#2800104; oxf anat brg rt md size 3 pma item#159575 lot#2944793. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00472; 3002806535-2022-00473. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15842907&pc=NRA&device_sequence_no=1
3002806535-2022-00473 | 10/31/2022 | 11/22/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision due to pain, approximately nine (9) years three (3) months after initial surgery. Total knee conversion was performed during the revision. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem md pma item#: 161469 lot#: 2920019; oxf uni tib tray sz c rm pma item#: 154723 lot#: 2800104. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00470; 3002806535-2022-00472. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15842885&pc=NRA&device_sequence_no=1
3002806535-2022-00472 | 10/31/2022 | 11/22/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent revision due to pain, approximately nine (9) years three (3) months after initial surgery. Total knee conversion was performed during the revision. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). D-10 oxf twin-peg cmntd fem md pma item#161469 lot#2920019. Oxf anat brg rt md size 3 pma item#159575 lot#2944793. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00470, 3002806535-2022-00473. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15842875&pc=NRA&device_sequence_no=1
3002806535-2022-00469 | 10/24/2022 | 11/21/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL | - | It was reported that the patient underwent revision surgery due to pain and implant fracture, approximately 6 months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15833210&pc=NRA&device_sequence_no=1
3002806535-2022-00460 | 10/16/2022 | 11/21/2022 | 36 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA LNR 36MM 54-56MM | HIP PROSTHESIS | - | It was reported that 8-month post-implantation, the patient underwent revision surgery due to implant fracture. Attempts have been made and no further information has been provided. | (b)(4). Foreign report source: china. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15832470&pc=LZO&device_sequence_no=1
3002806535-2022-00468 | 08/31/2022 | 11/17/2022 | 78 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL TRAY RM | KNEE PROTHESIS | - | It was reported that patient underwent revision due to lateral knee pain. Right knee total knee arthroplasty was performed during the operation. No further event information available at the time of this report. | (b)(4). Concomitant medical products: unknown oxford single peg femur #item unknown #lot unknown. Oxf anat brg rt lg size 3 pma item#159582 #lot 1332637. No product was returned, or pictures provided; visual and dimensional evaluations could not be performed. Unknown tibial tray: part and lot identification are necessary for review of device history records, neither were provided. Device is used for treatment. Medical records/radiographs were provided and reviewed by a health care professional. Review of the available records identified the following: here is advanced lateral compartment osteoarthritis. A joint effusion is noted. Bone quality is osteopenic. There is no evidence of joint overstuffing, but this would be best evaluated by earlier radiographs at the time of surgery. The current valgus alignment of the medial unicompartmental arthroplasty could lead to lateral compartment stress and osteoarthritis. The root cause of the reported issue is attributed to disease progression. No corrective actions, preventive actions, or field actions resulted after investigation of this event. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00466, 3002806535-2022-00467. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15813102&pc=NRA&device_sequence_no=1
3002806535-2022-00466 | 08/31/2022 | 11/17/2022 | 78 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD SINGLE PEG FEMUR | KNEE PROTHESIS | - | It was reported that patient underwent revision due to lateral knee pain. Right knee total knee arthroplasty was performed during the operation. No further event information available at the time of this report. | (b)(4). Concomitant medical products: oxf anat brg rt lg size 3 pma, item#159582, #lot 1332637; unknown oxford tibial tray rm, #item unknown, #lot unknown. No product was returned, or pictures provided; visual and dimensional evaluations could not be performed. Unknown oxford single peg: part and lot identification are necessary for review of device history records, neither were provided. Device is used for treatment. Medical records/radiographs were provided and reviewed by a health care professional. Review of the available records identified the following: here is advanced lateral compartment osteoarthritis. A joint effusion is noted. Bone quality is osteopenic. There is no evidence of joint overstuffing, but this would be best evaluated by earlier radiographs at the time of surgery. The current valgus alignment of the medial unicompartmental arthroplasty could lead to lateral compartment stress and osteoarthritis. The root cause of the reported issue is attributed to disease progression. No corrective actions, preventive actions, or field actions resulted after investigation of this event. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00467, 3002806535-2022-00468. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15813099&pc=NRA&device_sequence_no=1
3002806535-2022-00467 | 08/31/2022 | 11/17/2022 | 78 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT LG SIZE 3 PMA | KNEE PROTHESIS | - | T was reported that patient underwent revision due to lateral knee pain. Right knee total knee arthroplasty was performed during the operation. No further event information available at the time of this report. | (b)(4). Concomitant medial products: unknown oxford single peg femur #item unknown #lot unknown. Unknown oxford tibial tray rm #item unknown #lot unknown. No product was returned, or pictures provided; visual and dimensional evaluations could not be performed. Review of the device history record for item identified no deviations or anomalies during manufacturing related to the reported event. Device is used for treatment. Medical records/radiographs were provided and reviewed by a health care professional. Review of the available records identified the following: here is advanced lateral compartment osteoarthritis. A joint effusion is noted. Bone quality is osteopenic. There is no evidence of joint overstuffing, but this would be best evaluated by earlier radiographs at the time of surgery. The current valgus alignment of the media unicompartmental arthroplasty could lead to lateral compartment stress and osteoarthritis. The root cause of the reported issue is attributed to disease progression. No corrective actions, preventive actions, or field actions resulted after investigation of this event. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00466, 3002806535-2022-00468. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15813098&pc=NRA&device_sequence_no=1
3002806535-2022-00465 | 08/01/2022 | 11/16/2022 | 107 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CERAMIC HEAD | HIP PROSTHESIS | - | It was reported that the patient underwent a revision of right hip prosthesis due to dislocation and instability. During the procedure, extensive scar tissue was noted throughout the joint, tissue damaged to the posterior capsule and heterotopic bone to femur and acetabulum. The head and liner were exchanged without complications. | (b)(4). Concomitant medical products: g7 pps ltd acet shell 56f, item #: 010000665, lot #: 7001751, g7 vit e neutral lnr 36mm f, item #: 30103606, lot #: 65165414, stem, item #: 51104160, lot #: 7011414. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15805319&pc=LZO&device_sequence_no=1
3002806535-2022-00464 | 09/11/2019 | 11/15/2022 | 1161 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD MOD HD 36MM -3 NK | PROSTHESIS, HIP | - | It was reported that patient underwent revision surgery due to pain. Patient underwent a replacement of the head component the same day as the initial total hip arthroplasty. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: tloc 133 mp sp t1 ppsho 6x97. 5, item# 51-109060, lot# 6541761. 50mm o. D. Size hh porous uncemented with cluster holes shell use with hh liners, item# 00875705001, lot# 64311388. 32mm i. D. Size hh neutral liner use with 50mm o. D. Size hh shell, item# 00885100932, lot# 64271009. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15795596&pc=LZO&device_sequence_no=1
3002806535-2022-00463 | 10/07/2022 | 11/10/2022 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient experienced a dislocation approximately 15 days post right total hip arthroplasty and was corrected with a closed reduction. Subsequently, the patient experiences the second dislocation approximately 3 weeks later which was also corrected with a closed reduction. Patient reports severe pain, difficulty with activities of daily living, difficulty ambulating with walker and has a limp and decreased range of motion. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: g7 vit e neutral lnr 32mm e; item# 30103205; lot# 65224183, g7 pps ltd acet shell 52e; item# 010000663; lot# 7289876, and tprlc 133 fp type1 pps so 7. 0; item# 51-100070; lot# 7240457. Report source: the netherlands. Customer has indicated that the product will not be returned to zimmer biomet for the investigation as the device remains implanted. The investigation is in process. Once the investigation is completed, a follow-up mdr will be submitted. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15770126&pc=LZO&device_sequence_no=1
3002806535-2022-00462 | 09/13/2022 | 11/10/2022 | 58 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 32/-3MM T1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that the patient experienced a dislocation approximately 15 days post right total hip arthroplasty and was corrected with a closed reduction. Subsequently, the patient experienced the second dislocation approximately 3 weeks later which was also corrected with a closed reduction. Patient reports severe pain, difficulty with activities of daily living, difficulty ambulating with walker and has a limp with a decrease in range of motion. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: g7 vit e neutral lnr 32mm e; item# 30103205; lot# 65224183, g7 pps ltd acet shell 52e; item# 010000663; lot# 7289876, tprlc 133 fp type1 pps so 7. 0; item# 51-100070; lot# 7240457. Foreign source: the netherlands. Customer has indicated that the product will not be returned to zimmer biomet for the investigation as the device remains implanted. The investigation is in process. Once the investigation is completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15770120&pc=LZO&device_sequence_no=1
3002806535-2022-00461 | 10/18/2022 | 11/10/2022 | 23 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | QUICK CONNECT TWT DRL 3.2X70MM | KNEE INSTRUMENT | - | It was reported that during surgery drill bit broken. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: poland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15769575&pc=HRY&device_sequence_no=1
3002806535-2022-00459 | 10/17/2022 | 11/09/2022 | 23 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | QUICK CONNECT TWT DRL 3.2X50MM | HIP INSTRUMENT | - | It was reported that drill bit was damaged during the surgery. No health consequences to patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source foreign: poland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15757943&pc=HTW&device_sequence_no=1
3002806535-2022-00457 | 06/14/2022 | 11/08/2022 | 147 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT | KNEE INSTRUMENTS | - | It was reported that one of the metal parts of the oxford cementless tibial impactor broke off during tibial component insertion and was left inside the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: united kingdom. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15750382&pc=HWA&device_sequence_no=1
3002806535-2022-00456 | 10/11/2022 | 11/07/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNI TWIN-PEG FEMORAL MD | KNEE PROSTHESIS | - | It was reported that loosening of the femoral component occurred. A fractured anterior peg could have been a contributing factor. The patient was revised to a total knee. | (b)(4). Report source: foreign ¿ canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15742240&pc=NRA&device_sequence_no=1
0009613350-2022-00571 | 02/08/2021 | 11/07/2022 | 637 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | . waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN PEG CMNTLS FMRL MD | KNEE PROTHESIS | - | It was reported in a clinical study that seven years after initial knee arthroplasty, a patient experienced severe pain and increased swelling in the knee. An aspiration of the joint was performed (results not provided). Cortisone injections were given which provided about 2-3 weeks of pain relief. Additionally, an injection of platelet rich plasma was also given. No revision has been noted to date. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf anat brg lt md size 3, pma, item#159547, lot#369770; oxford cementless tibia d lm, item#us166576, lot#3050476. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00454, 3002806535-2022-00455. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15741919&pc=NRA&device_sequence_no=1
3002806535-2022-00455 | 02/08/2021 | 11/07/2022 | 637 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD CEMENTLESS TIBIA D LM | KNEE PROTHESIS | - | It was reported in a clinical study that seven years after initial knee arthroplasty, a patient experienced severe pain and increased swelling in the knee. An aspiration of the joint was performed (results not provided). Cortisone injections were given which provided about 2-3 weeks of pain relief. Additionally, an injection of platelet rich plasma was also given. No revision has been noted to date. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf anat brg lt md size 3 pma item#159547 lot#369770, oxford cementless tibia d lm item#us166576 lot#3050476. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 0009613350-2022-00571, 3002806535-2022-00454. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15741918&pc=NRA&device_sequence_no=1
3002806535-2022-00454 | 02/08/2021 | 11/07/2022 | 637 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROTHESIS | - | It was reported in a clinical study that seven years after initial knee arthroplasty, a patient experienced severe pain and increased swelling in the knee. An aspiration of the joint was performed (results not provided). Cortisone injections were given which provided about 2-3 weeks of pain relief. Additionally, an injection of platelet rich plasma was also given. No revision has been noted to date. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf twin peg cmntls fmrl md item#161474 lot#2503468; oxford cementless tibia d lm item#us166576 lot#3050476. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 0009613350-2022-00571; 3002806535-2022-00455. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15741917&pc=NRA&device_sequence_no=1
3002806535-2022-00447 | 10/04/2022 | 11/04/2022 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due an unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: unknown oxford femoral component, item name: unknown oxford bearing. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00446 and 3002806535-2022-00448. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15726996&pc=NRA&device_sequence_no=1
3002806535-2022-00448 | 10/04/2022 | 11/04/2022 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: unknown oxford femoral component; item#: unknown; lot#: unknown, item name: unknown oxford tibial component; item#: unknown; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00446, 3002806535 - 2022 - 00447. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15726684&pc=NRA&device_sequence_no=1
3002806535-2022-00446 | 10/04/2022 | 11/04/2022 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent revision surgery due an unknown reason. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: item name: unknown oxford tibial component; item#: unknown; lot#: unknown, item name: unknown oxford bearing; item#: unknown; lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 - 2022 - 00447, 3002806535 - 2022 - 00448. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15725481&pc=NRA&device_sequence_no=1
3002806535-2022-00453 | 10/11/2022 | 11/03/2022 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | MH SOLID SHELL HA/PC 50MM LN23 | HIP PROSTHESIS | - | It was reported that a patient underwent a hip revision approximately 11 years post implantation due to the head wearing into the liner causing metallosis. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Report source: (b)(6). Concomitant medical products: unknown stem, e-poly 36mm +3 hiwall lnr sz23, item ep-108323, lot 026270, 36mm cocr mod hd -3mm, item 11-363661, lot 506070. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15722548&pc=KWA&device_sequence_no=1
3002806535-2022-00450 | 10/07/2022 | 11/02/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent an initial right knee arthroplasty. Subsequently, a revision procedure due to unknown reason was performed approximately two (2) years and five (5) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt sm size 3 pma item# 159568 lot# 863520; oxf uni tib tray sz c rm pma item# 154723 lot# 434000. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00451; 3002806535 - 2022 - 00452. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713559&pc=NRA&device_sequence_no=1
3002806535-2022-00451 | 10/07/2022 | 11/02/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent an initial right knee arthroplasty. Subsequently, a revision procedure due to unknown reason was performed approximately two (2) years and five (5) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem sm pma item# 161468 lot# 674360; oxf uni tib tray sz c rm pma item# 154723 lot# 434000. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00450; 3002806535 - 2022 - 00452. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713556&pc=NRA&device_sequence_no=1
3002806535-2022-00452 | 10/07/2022 | 11/02/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent an initial right knee arthroplasty. Subsequently, a revision procedure due to unknown reason was performed approximately two (2) years and five (5) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem sm pma item# 161468 lot# 674360; oxf anat brg rt sm size 3 pma item# 159568 lot# 863520. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00450; 3002806535 - 2022 - 00451; 3002806535 - 2022 - 00452. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713554&pc=NRA&device_sequence_no=1
3002806535-2022-00449 | 10/06/2022 | 11/02/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/-3MM | HIP PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the hip due to infection, approximately two (2) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: g7 bonemaster ltd acet shl 60g item# 010000707 lot# 6966249; g7 hi-wall e1 liner 36mm g item# 010000937 lot# 7210978. Report source: foreign - australia. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713500&pc=LZO&device_sequence_no=1
3002806535-2022-00441 | 10/07/2022 | 11/02/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNCEMENTED TIBIAL IMPACTOR | ORTHOPAEDIC IMPLANT IMPACTOR | - | It was reported that during the removal of the cementless impactor, the hook that secures the definitive tibial component to the instrument broke off. This was only noticed on post operative x-rays. No known impact or consequence to the patient. Attempts have been made and no further information has been provided. | (b)(4). Report source ¿ foreign ¿ united kingdom. This follow-up report is being submitted to relay additional and/or corrected information. No product was returned, or pictures provided; visual and dimensional evaluations could not be performed. Lot identification is necessary for review of device history records, lot identification was not provided. Radiographs were provided and reviewed by a health care professional. Review of the available records identified the following: x-ray attached to complaint does not appear to show the fragment of fractured instrument and could not be used to conclusively confirm reported event. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15713489&pc=HWA&device_sequence_no=1
3002806535-2022-00445 | 08/26/2022 | 10/31/2022 | 66 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD INTRAMEDULLARY ROD LINK | KNEE INSTRUMENT | - | It was reported that during the revision surgery of the knee, the two-prong connection guide broke. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source¿ foreign ¿ south africa. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15698453&pc=HRY&device_sequence_no=1
3002806535-2022-00440 | 10/03/2022 | 10/31/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 3 PMA | KNEE PROTHESIS | - | It was reported that the patient underwent conversion of partial to total knee implant due to unknown reason, approximately 4 years after initial implantation. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf uni tib tray sz b lm pma item#154720 lot#000450. Oxf twin-peg cmntd fem sm pma item#161468 lot#113140. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00439; 3002806535-2022-00438. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15698134&pc=NRA&device_sequence_no=1
3002806535-2022-00438 | 10/03/2022 | 10/31/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent conversion of partial to total knee implant due to unknown reason, approximately 4 years after initial implantation. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf uni tib tray sz b lm pma, item#154720, lot#000450; oxf anat brg lt sm size 3 pma, item#159540, lot#299880. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00439; 3002806535-2022-00440. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15698119&pc=NRA&device_sequence_no=1
3002806535-2022-00439 | 10/03/2022 | 10/31/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | KNEE PROTHESIS | - | It was reported that the patient underwent conversion of partial to total knee implant due to unknown reason, approximately 4 years after initial implantation. Due diligence is in progress for this event; to date no further information has been provided. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem sm pma item#161468, lot#113140 oxf anat brg lt sm size 3 pma item#159540 lot#299880. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00438, 3002806535-2022-00440. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15698111&pc=NRA&device_sequence_no=1
3002806535-2022-00444 | 10/05/2022 | 10/27/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of left knee in an unknown date due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibial component item# unknown lot# unknown. Unknown oxford bearing item# unknown lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00442, 3002806535 - 2022 - 00443. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15679524&pc=NRA&device_sequence_no=1
3002806535-2022-00442 | 10/05/2022 | 10/27/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of left knee in an unknown date due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford bearing item# unknown lot# unknown. Unknown oxford femoral component item# unknown lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00443, 3002806535 - 2022 - 00444. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15679518&pc=NRA&device_sequence_no=1
3002806535-2022-00443 | 10/05/2022 | 10/27/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of left knee in an unknown date due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibial component item# unknown lot# unknown. Unknown oxford femoral component item# unknown lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00442 and 3002806535-2022-00444. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15679499&pc=NRA&device_sequence_no=1
3002806535-2022-00433 | 10/10/2022 | 10/25/2022 | 15 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 36MM F | HIP PROSTHESIS | - | It was reported that the patient heard squeaking in the hip joint, and experienced pain. The ceramic liner was found to be fractured and disassociated from the shell (cup). A new ploy liner and head were implanted. The patient does not report falling or suffering any trauma. | (b)(4). Concomitant medical products: g7 bonemaster ltd acet shl 54f lot# unknown item# unknown. Modular head lot# unknown item# unknown. Report source¿ foreign ¿ ¿belgium" an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15665429&pc=LZO&device_sequence_no=1
3002806535-2022-00437 | 09/26/2022 | 10/24/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | PROSTHESIS, KNEE, FEMOROTIBIAL, UNICOMPARTMENTAL | - | It was reported that the patient underwent revision surgery due to dislocation of the bearing. A 4mm bearing revised to a 5mm bearing. Due diligence is in progress for this complaint; to date, whatever additional information received has been included in this report. | (b)(4). Item#161469; lot#414420; item name: oxf twin-peg cmntd fem md pma; item#154724; lot#080510; item name: oxford pks cocr size d lm std. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15657593&pc=NRA&device_sequence_no=1
3002806535-2022-00436 | 10/03/2022 | 10/21/2022 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that 3 years after knee implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4) multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00434. 3002806535-2022-00435. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15649368&pc=NRA&device_sequence_no=1
3002806535-2022-00435 | 10/03/2022 | 10/21/2022 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that 3 years after knee implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00434, 3002806535 -2022 -00436. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15649364&pc=NRA&device_sequence_no=1
3002806535-2022-00434 | 10/03/2022 | 10/21/2022 | 18 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that 3 years after knee implantation, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00435, 3002806535 -2022 -00436. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15649357&pc=NRA&device_sequence_no=1
3002806535-2022-00425 | 09/19/2022 | 10/18/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPTION TYPE 1 TPR SLEVE +6 | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the left hip due to a pseudotumor. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: bearing 28 mm i. D. 46 mm o. D. Size g , item# 110031001 , lot# 64515040; cer bioloxd option hd 28mm , item# 650-1055 , lot# 3087211. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00432. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15624630&pc=LZO&device_sequence_no=1
3002806535-2022-00432 | 09/19/2022 | 10/18/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 28MM | HIP PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the left hip due to a pseudotumor. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: bearing 28 mm i. D. 46 mm o. D. Size g , item# 110031001 , lot# 64515040. Cer option type 1 tpr sleve +6 , item# 650-1068 , lot# 3063681. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00425. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15624401&pc=LZO&device_sequence_no=1
3002806535-2022-00426 | 09/20/2022 | 10/18/2022 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TRL BRG W/SLOTS MED 3MM | KNEE INSTRUMENT | - | It was reported that during the surgery when the surgeon was trialing, the bearing trial chipped. No pieces felt in patient. No harm reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign - (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15622581&pc=NRA&device_sequence_no=1
3002806535-2022-00430 | 09/19/2022 | 10/14/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent initial oxford uni knee surgery. Approximately 5 years later a revision surgery was performed due pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: medical device: oxf anat brg lt sm size 3 pma; item# 159540; lot# 989610. Oxf uni tib tray sz c lm pma; item# 154722; lot# 336710. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00428, 3002806535-2022-00429. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Sent for cultures. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15602786&pc=NRA&device_sequence_no=1
3002806535-2022-00428 | 09/19/2022 | 10/14/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent initial oxford uni knee surgery. Approximately 5 years later a revision surgery was performed due pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf uni tib tray sz c lm pma, item# 154722; lot# 336710. Oxf twin-peg cmntd fem sm pma, item# 161468; lot# 101130, multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00429. 3002806535 - 2022 - 00430. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Sent for cultures. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15602768&pc=NRA&device_sequence_no=1
3002806535-2022-00429 | 09/19/2022 | 10/14/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that a patient underwent initial oxford uni knee surgery. Approximately 5 years later a revision surgery was performed due pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: oxf anat brg lt sm size 3 pma; item# 159540; lot# 989610. Oxf twin-peg cmntd fem sm pma; item# 161468; lot# 101130. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00428, 3002806535 - 2022 - 00430. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Sent for cultures. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15602767&pc=NRA&device_sequence_no=1
3002806535-2022-00427 | 09/30/2022 | 10/13/2022 | 13 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROSTHESIS | - | It was reported that the patient underwent revision due to implant fracture. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: medical devices. Cer option type 1 tpr sleve -6 , item# 650-1064 , lot# 356070. Epoly 36mm rlc lnr mrom sz24 , item# ep-105994 , lot# 838400. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15595224&pc=LZO&device_sequence_no=1
3002806535-2022-00424 | 09/15/2022 | 10/10/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial oxford uni partial knee replacement surgery. Approximately 8 months later, a revision was performed due to implant dislocation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4) the device will not be returned for analysis. However; an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Discarded. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15574226&pc=NRA&device_sequence_no=1
3002806535-2022-00421 | - | 10/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | KNEE PROSTHESIS | - | The patient underwent an initial left knee hemiarthroplasty. Subsequently, the patient experienced loosening, swelling, and difficulty ambulating. No revision procedure has occurred at this time. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: item#: 161470; lot#: unknown, item#: 159554; lot#: unknown. Therapy date: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2022 -00420, 3002806535 -2022 -00422. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15560392&pc=NRA&device_sequence_no=1
3002806535-2022-00420 | - | 10/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM LG PMA | KNEE PROSTHESIS | - | The patient underwent an initial left knee hemiarthroplasty. Subsequently, the patient experienced loosening, swelling, and difficulty ambulating. No revision procedure has occurred at this time. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: item#: 154724; lot#: unknown, item#: 159554; lot#: unknown. Therapy date: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2022 -00421, 3002806535 -2022 -00422. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15560386&pc=NRA&device_sequence_no=1
3002806535-2022-00422 | - | 10/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT LG SIZE 3 PMA | KNEE PROSTHESIS | - | The patient underwent an initial left knee hemiarthroplasty. Subsequently, the patient experienced loosening, swelling, and difficulty ambulating. No revision procedure has occurred at this time. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | Cmp-(b)(4). Concomitant medical products: item#: 161470; lot#: unknown; item#: 154724; lot#: unknown. Therapy date: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see the associated reports: 3002806535 -2022 -00420; 3002806535 -2022 -00421. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15560367&pc=NRA&device_sequence_no=1
3002806535-2022-00412 | 02/29/2012 | 10/04/2022 | 3870 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | EXCEED ABT E1 MULLER CUP 28X56 | HIP PROSTHESIS | - | It was reported in a clinical study that the patient underwent a revision procedure approximately 1 month post implantation due to dislocation where the head was exchanged and impingement of the front left acetabulum due to osteophytes was removed. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: unknown modular head component 28mm + 6mm, unknown mallory head stem 12x165mm. Foreign source: netherlands. Study - clinical study ortho. Cr. Gh005. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00413. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of the device history record identified no deviations or anomalies during manufacturing. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15534975&pc=KWA&device_sequence_no=1
3002806535-2022-00419 | 09/07/2022 | 10/03/2022 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD TRIAL BRNG LRG SZ 4/5 | KNEE INSTRUMENT | - | It was reported, that the trial bearing cracked where it screws into the handle. No consequences or impact to the patient. Attempts have been made and no further information has been provided. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15526863&pc=HRY&device_sequence_no=1
3002806535-2022-00411 | 09/13/2022 | 09/29/2022 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | MENISCAL UNICOMPARTMENTAL KNEE SYSTEM | - | It was reported that the patient underwent a right knee arthroplasty approximately six (6) years ago. Subsequently, the patient was revised from an oxford uni to a total knee due to unknown reasons. | (b)(4). Concomitant medical devices: oxf uni tib tray sz c rm pma item # 154723 lot# 518140, oxf anat brg rt md size 3 pma item # 159575 lot # 824600. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00408, 3002806535-2022-00410. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted h3 other text : discarded. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15508600&pc=NRA&device_sequence_no=1
3002806535-2022-00410 | 09/13/2022 | 09/29/2022 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 3 PMA | MENISCAL UNICOMPARTMENTAL KNEE SYSTEM | - | It was reported that the patient underwent a right knee arthroplasty approximately six (6) years ago. Subsequently, the patient was revised from an oxford uni to a total knee due to unknown reasons. | (b)(4). Concomitant medical products: oxf uni tib tray sz c rm pma item # 154723 lot# 518140, oxf twin-peg cmntd fem md pma item # 161469 lot#687500. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00408, 3002806535-2022-00411. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15508348&pc=NRA&device_sequence_no=1
3002806535-2022-00408 | 09/13/2022 | 09/29/2022 | 16 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C RM PMA | MENISCAL UNICOMPARTMENTAL KNEE SYSTEM | - | It was reported that the patient underwent a right knee arthroplasty approximately six years ago. Subsequently, the patient was revised from an oxford uni to a total knee due to unknown reasons. | (b)(4). Oxf anat brg rt md size 3 pma item #159575 lot#824600. Oxf twin-peg cmntd fem md pma item # 161469 lot#687500. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00410 and 3002806535-2022-00411. The device will not be returned for analysis. However; an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507951&pc=NRA&device_sequence_no=1
3002806535-2022-00416 | 09/06/2022 | 09/29/2022 | 23 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D LM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery of the left knee due to loosening, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf twin-peg cmntd fem lg item# 161470 lot# 933250, oxf anat brg lt lg size 3 pma item# 159554 lot# 151740, biomet bc r 1x40 us item# 110035368 lot# 915eak1907, biomet bc r 1x40 us item# 110035368 lot# az48ck2305. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507735&pc=NRA&device_sequence_no=1
3002806535-2022-00402 | 09/13/2022 | 09/29/2022 | 16 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK | KNEE INSTRUMENT | - | It was reported that the pin of handle broke while attempting to remove the intramedullary rod. No delay was reported. No impact or consequences to the patient or the user. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507648&pc=LXH&device_sequence_no=1
3002806535-2022-00403 | 09/12/2022 | 09/29/2022 | 17 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZA LM | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision surgery due to subsidence of the tibial baseplate 3 years post-implantation. Patient was revised to primary tka. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products; lps option femoral e-l , item# 00-5996-015-51 , lot# 63742529, st prc tib plt size 5 , item# 00-5980-047-01 , lot# j6828721, lps flex art surf ef5 6 14 , item# 00-5962-040-14 , lot# 64971772, stem implant 15mmdx75mm. Item# 00-5988-012-15 , lot# 64946296, all poly pat comp 32dia , item# 00-5972-065-32 , lot# 65341199, oxf twin-peg cmntd fem md pma , item# 161469 , lot# 266810, oxf anat brg lt md size 4 pma , item# 159548 , lot# 403220. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15507310&pc=NRA&device_sequence_no=1
3002806535-2022-00414 | 03/09/2022 | 09/29/2022 | 204 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TAPERLOC PRIMARY STEM, POROUS TI TYPE 12/14 TAPER C'LESS - 15.0 X 150MM | HIP PROSTHESIS | - | It was reported that the patient experienced persistent coxalgia and underwent revision of right hip prosthesis due to loosening of stem and acetabular cup. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: unknown ringloc cup item# unknown lot# unknown, unknown ceramic head item# unknown lot# unknown. Report sources: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated report: 3002806535 - 2022 - 00415. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15501670&pc=KWA&device_sequence_no=1
3002806535-2022-00415 | 03/09/2022 | 09/29/2022 | 204 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN RINGLOC CUP | HIP PROTHESIS | - | It was reported that the patient experienced persistent coxalgia and underwent revision of right hip prosthesis due to loosening of stem and acetabular cup. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: tprloc 12/14 por 15x150 item# 650-0322 lot# unknown. Unknown ceramic head item# unknown lot# unknown. Report source- italy. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00414. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15501440&pc=LZO&device_sequence_no=1
3002806535-2022-00413 | 02/29/2012 | 09/28/2022 | 3864 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN MODULAR HEAD COMPONENT 28MM+6MM | HIP PROSTHESIS | - | It was reported in a clinical study that the patient underwent a revision procedure approximately 1 month post implantation due to dislocation where the head was exchanged and impingement of the front left acetabulum due to osteophytes was removed. Attempts have been made and no further information has been provided. | (b)(4). Medical products exceed abt e1 muller cup 28x56; item: ep-122856: lot: 2241124. Unknown mallory head stem 12x165mm. Foreign netherlands study clinical study ortho. Cr. Gh005. Multiple mdr reports were filed for this event, please see associated report: 3002806535 -2022 -00412. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Medical records were not provided. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500924&pc=LZO&device_sequence_no=1
3002806535-2022-00407 | 08/31/2022 | 09/28/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | PROSTHESIS, KNEE, FEMOROTIBIAL | - | It was reported that the patient underwent revision surgery due to unknown reasons. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Comconitant medical devices: item#:154720; lot#: unknown; itemname: oxf uni tib tray sz b lm pma; item#:159548; lot#: unknown; itemname: oxf anat brg lt md size 4 pma. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00406 and 3002806535 -2022 -00409. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500525&pc=NRA&device_sequence_no=1
3002806535-2022-00391 | 09/06/2022 | 09/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM SM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery due to pain and wear approximately one and a half (1. 5) year post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical device: oxf uni tib tray sz c lm pma item#: 154722 lot#: 013320, oxf anat brg lt sm size 4 pma item#: 159541 lot#: 464920, pulsavc plus fan spray kit int item#: 00-5150-475-01 lot#: 64826482, quick-vac mixing bowl single item#: 00-5049-001-00 lot#: 56518519, oxf sawblade stryker cmntd 3pk item#: 506298 lot#: 362651. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00393 and 3002806535-2022-00394. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500503&pc=NRA&device_sequence_no=1
3002806535-2022-00394 | 09/06/2022 | 09/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT SM SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery due wear and pain approximately one and a half (1. 5) year post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf twin-peg cmntd fem sm pma item#: 161468 lot#: 649630, oxf uni tib tray sz c lm pma item#: 154722 lot#: 013320, pulsavc plus fan spray kit int item#: 00-5150-475-01 lot#: 64826482, quick-vac mixing bowl single item#: 00-5049-001-00 lot#: 56518519, oxf sawblade stryker cmntd 3pk item#: 506298 lot#: 362651. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00391 and 3002806535-2022-00393. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500406&pc=NRA&device_sequence_no=1
3002806535-2022-00393 | 09/06/2022 | 09/28/2022 | 22 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision surgery due wear and pain approximately one and a half (1. 5) year post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf twin-peg cmntd fem sm pma item#: 161468 lot#: 649630, oxf anat brg lt sm size 4 pma item#: 159541 lot#: 464920, pulsavc plus fan spray kit int item#: 00-5150-475-01 lot#: 64826482, quick-vac mixing bowl single item#: 00-5049-001-00 lot#: 56518519, oxf sawblade stryker cmntd 3pk item#: 506298 lot#: 362651. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00391 and 3002806535-2022-00394. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500385&pc=NRA&device_sequence_no=1
3002806535-2022-00409 | 08/31/2022 | 09/28/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 4 PMA | PROSTHESIS, KNEE, FEMOROTIBIAL | - | It was reported that the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant: item#:154720; lot#: unknown; itemname: oxf uni tib tray sz b lm pma; item#:161469; lot#: unknown; itemname: oxf twin-peg cmntd fem md pma; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00406, 3002806535 -2022 -00407. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500364&pc=NRA&device_sequence_no=1
3002806535-2022-00406 | 08/31/2022 | 09/28/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B LM PMA | PROSTHESIS, KNEE, FEMOROTIBIAL | - | It was reported that the patient underwent revision surgery due to an unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item#:161469; lot#: unknown; itemname: oxf twin-peg cmntd fem md pma; item#:159548; lot#: unknown; itemname: oxf anat brg lt md size 4 pma. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00407, 3002806535 -2022 -00409. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15500358&pc=NRA&device_sequence_no=1
3002806535-2022-00382 | 06/25/2019 | 09/28/2022 | 1191 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN EXCEED CUP ARCOMULLER 28X58 | HIP PROSTHESIS | - | It was reported that patient had a cup revision due to cup loosening approximately seven (7) years post op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: netherlands. Concomitant medical products: unknown modular head component 28mm + 6mm, item# unknown, lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15496884&pc=JDI&device_sequence_no=1
3002806535-2022-00404 | 04/05/2022 | 09/27/2022 | 175 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK | KNEE INSTRUMENT | - | It was reported that the pin of the handle hook broke. No impact or consequences to the patient or the user were reported. Attempts have been made and no further information has been provided. | (b)(4). This report is being submitted to relay additional and/or corrected information. Visual examination of the returned product identified that one of the pins has fractured. The fractured pin was not returned with the instrument. The area around the fracture has numerous indentations, however, it is not possible to determine how these occurred. Review of the device history record identified no deviations or anomalies during manufacturing. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15493202&pc=LXH&device_sequence_no=1
3002806535-2022-00405 | 09/02/2022 | 09/27/2022 | 25 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TRL BRG W/SLOTS MED 3MM MM | GENERAL KNEE INSTRUMENTS | - | It was reported that while the surgeon was trying to take out trial bearing from the knee, the bearing broke. No consequences or impact to the patient. Attempts have been made and no further information has been provided. | (b)(4). Report source foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15492983&pc=HRY&device_sequence_no=1
3002806535-2022-00397 | 06/29/2022 | 09/19/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN CERAMIC HEAD 28 MM | HIP PROSTHESIS | - | It was reported that the patient was revised due to polyethylene wear approximately 15 years post implantation. The head and liner were removed and replaced. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: reference number: 11-105904, item name: arcom 28mm rngloc lnr hwall 24, lot number: 567700. Visual examination of the provided pictures identified the ceramic head has marks and lines on the internal taper, bottom flat and sphere but otherwise appears to be in good condition. Part and lot identification are necessary for review of device history records, neither were provided. Device is used for treatment. Radiographs were provided and reviewed by a health care professional. Review of the available records identified the following: overall size of both implants is appropriate. Polyethylene wear of the acetabular cups is noted. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15444586&pc=LZO&device_sequence_no=1
3002806535-2022-00390 | 06/18/2022 | 09/19/2022 | 93 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED | - | It was reported that a patient was diagnosed with post-operative left popliteal, deep vein thrombosis despite fragmin prescription. Symptoms of calf tension and minimal redness began three days after left total hip arthroplasty. The thrombosis was resolved without complications; all implants remained in place. Attempts have been made and additional information on the reported event is unavailable at this time. | (b)(4). Foreign source: netherlands. Concomitant medical products: item name: cer option type 1 tpr sleve -6; item#:650-1064; lot#3100139, item name: g7 vit e neutral lnr 36mm f; item#:30103606; lot#65256767; item name: g7 pps solid acet shell 66i; item#:010000650; lot#7194784; item name: tprlc 133 t1 pps ho 16x152mm; item#:51-104160;lot# 7142517. The device will not be returned for analysis, as the device remains implanted; however, an investigation of the reported event is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00392. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15444509&pc=LZO&device_sequence_no=1
3002806535-2022-00392 | 06/18/2022 | 09/19/2022 | 93 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER OPTION TYPE 1 TPR SLEVE -6 | PROSTHESIS, HIP, SEMI-CONSTRAINED | - | It was reported that a patient was diagnosed with post-operative left popliteal, deep vein thrombosis despite fragmin prescription. Symptoms of calf tension and minimal redness began three days after left total hip arthroplasty. The thrombosis was resolved without complications; all implants remained in place. Attempts have been made and additional information on the reported event is unavailable at this time. | (b)(4). Foreign source: netherlands. Concomitant medical products: item name: delta ceramic fem hd 36/0mm; item#:650-0661; lot#3111491; item name: g7 vit e neutral lnr 36mm f; item#:30103606; lot#65256767; item name: g7 pps solid acet shell 66i; item#:010000650; lot#7194784; item name: tprlc 133 t1 pps ho 16x152mm; item#:51-104160;lot# 7142517. The device will not be returned for analysis, as the device remains implanted; however, an investigation of the reported event is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00390. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15444171&pc=LZO&device_sequence_no=1
3002806535-2022-00396 | - | 09/14/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN BIOLOX DELTA HEAD | HIP PROSTHESIS | - | It was reported in a journal article that one patient, within the fixed bearing group, experienced a deep joint infection. No further discussion or information regarding intervention or outcome was provided. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: unknown liner. Unknown cup. Unknown stem. Foreign: country: south korea. Literature: moon, n. H. , do, m. U. , kim, j. S. Et al. The usefulness of dual mobility cups in primary total hip arthroplasty patients at a risk of dislocation. Sci rep 12, 774 (2022). Https://doi. Org/10. 1038/s41598-022-04774-2. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15418008&pc=LPH&device_sequence_no=1
3002806535-2022-00395 | - | 09/14/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN BIOLOX DELTA HEAD | HIP PROSTHESIS | - | It was reported in a journal article that four patients, within the fixed bearing group, experienced a dislocation. No further discussion or information regarding intervention or outcome was provided. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: unknown liner, unknown cup, unknown stem. Foreign source: south korea. Literature: moon, n. H. , do, m. U. , kim, j. S. Et al. The usefulness of dual mobility cups in primary total hip arthroplasty patients at a risk of dislocation. Sci rep 12, 774 (2022). Https://doi. Org/10. 1038/s41598-022-04774-2. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15417647&pc=LPH&device_sequence_no=1
3002806535-2022-00388 | 08/24/2022 | 09/14/2022 | 21 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TIB GAP GAUGE MED 1MM | GENERAL KNEE INSTRUMENT | - | It was reported that when sizing the femur, the oxford tibial gap gauge tip broke off. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign ¿ south africa. Investigation of this incident is currently ongoing. A followup/ final report will be submitted when additional information becomes available. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15415799&pc=LXH&device_sequence_no=1
3002806535-2022-00387 | 05/03/2019 | 09/13/2022 | 1229 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 40MM | HIP PROSTHESIS | - | It was reported that patient underwent a left total hip revision due to unknown reason, approximately five (5) years post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Investigation of this incident is currently ongoing. A followup/final report will be submitted when additional information becomes available. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15409544&pc=LZO&device_sequence_no=1
3002806535-2022-00385 | 06/22/2022 | 09/12/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT X-SM SZ 3 PMA | KNEE PROTHESIS | - | It was reported the patient underwent a left knee revision 27 months post-implantation due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem xs pma , item# 161467 , 492840, oxf uni tib tray sz aa lm pma , item# 159531 , lot# 105880. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00384, 3002806535 -2022 -00386. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15402117&pc=NRA&device_sequence_no=1
3002806535-2022-00386 | 06/22/2022 | 09/12/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ AA LM PMA | KNEE PROTHESIS | - | It was reported the patient underwent a left knee revision 27 months post-implantation due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products; oxf twin-peg cmntd fem xs pma , item# 161467 , 492840, oxf anat brg lt x-sm sz 3 pma, item# 159790 , lot# 600760. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00384, 3002806535 -2022 -00385. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15402109&pc=NRA&device_sequence_no=1
3002806535-2022-00384 | 06/22/2022 | 09/12/2022 | 82 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM XS PMA | KNEE PROTHESIS | - | It was reported the patient underwent a left knee revision 27 months post-implantation due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical products; oxf anat brg lt x-sm sz 3 pma , item# 159790 , lot# 600760. Oxf uni tib tray sz aa lm pma , item# 159531 , lot# 105880. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00385. 3002806535 -2022 -00386. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15402106&pc=NRA&device_sequence_no=1
3002806535-2022-00383 | 08/16/2022 | 09/12/2022 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF MP 04MM DRILL | KNEE INSTRUMENT | - | It was reported that drill broke. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: poland. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15400923&pc=HTW&device_sequence_no=1
3002806535-2022-00381 | 08/29/2022 | 09/12/2022 | 14 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROTHESIS | - | It was reported that approximately 9 years post-operative, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford bearing, unknown oxford femoral component. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00379, 3002806535 -2022-00380. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15400237&pc=NRA&device_sequence_no=1
3002806535-2022-00379 | 08/29/2022 | 09/12/2022 | 14 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROTHESIS | - | It was reported that approximately 9 years post-operative, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products. Unknown oxford femoral component, unknown oxford tibial component. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00380, 3002806535 -2022 -00381. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15399232&pc=NRA&device_sequence_no=1
3002806535-2022-00380 | 08/29/2022 | 09/12/2022 | 14 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROTHESIS | - | It was reported that approximately 9 years post-operative, the patient underwent revision surgery due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical products: unknown oxford bearing, unknown oxford tibial component. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 -2022 -00379, 3002806535 -2022 -00381. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15399221&pc=NRA&device_sequence_no=1
3002806535-2022-00378 | 08/08/2022 | 09/08/2022 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 4 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of knee prosthesis due to bearing dislocation in unknown time frame post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: china. Concomitant medical products: associated devices: reference number:166941, device name: oxford uni twin-peg femoral sm, lot:6737518, reference number: 154719, device name: oxf uni tib tray sza rm, lot:6681212. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15376344&pc=NRA&device_sequence_no=1
3002806535-2022-00377 | - | 09/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient experienced pain and swelling approximately 2. 5 months after implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item name: unknown oxford femoral component; item#:unknown; lot#: unknown; item name: unknown oxford tibial component; item#:unknown; lot#: unknown; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00375, 3002806535 - 2022 - 00376. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15368174&pc=NRA&device_sequence_no=1
3002806535-2022-00376 | - | 09/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient experienced pain and swelling approximately 2. 5 months after implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant products: item name: unknown oxford femoral component; item#: unknown; lot#: unknown; item name: unknown oxford bearing; item#: unknown; lot#: unknown; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00375, 3002806535 - 2022 - 00377. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15368172&pc=NRA&device_sequence_no=1
3002806535-2022-00375 | - | 09/07/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient experienced pain and swelling approximately 2. 5 months after implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item name: unknown oxford tibial component; item#:unknown; lot#: unknown; item name: unknown oxford bearing; item#:unknown; lot#: unknown; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00376, 3002806535 - 2022 - 00377. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15368168&pc=NRA&device_sequence_no=1
3002806535-2022-00374 | 08/24/2022 | 09/06/2022 | 13 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RHK 16X80 CEMENTED STEM | HIP PROSTHESIS | - | It was reported, that an expired product has been implanted approximately 1 month after the sterile expiration date. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign - spain. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15359986&pc=KWA&device_sequence_no=1
3002806535-2022-00371 | 08/03/2022 | 09/06/2022 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that a patient was revised from an oxford partial knee system to a total knee due to fracture and wear of the oxford components. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: united kingdom. Concomitant medical products: associated devices: reference number: unknown, product name: unknown oxford femoral, lot number: unknown. Reference number: unknown, product name: unknown oxford tibial component, lot number: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00370, 3002806535-2022-00372. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15359440&pc=NRA&device_sequence_no=1
3002806535-2022-00370 | 08/03/2022 | 09/06/2022 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that a patient was revised from an oxford partial knee system to a total knee due to fracture and wear of the oxford components. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign- united kingdom. Associated devices: reference number: unknown, product name: unknown oxford bearing, lot number: unknown. Reference number: unknown, product name: unknown oxford tibial component, lot number: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00371, 3002806535-2022-00372. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15359420&pc=NRA&device_sequence_no=1
3002806535-2022-00372 | 08/03/2022 | 09/06/2022 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that a patient was revised from an oxford partial knee system to a total knee due to fracture and wear of the oxford components. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: united kingdom. Concomitant medical products: associated devices: reference number: unknown, product name: unknown oxford bearing, lot number: unknown. Reference number: unknown, product name: unknown oxford femoral component, lot number: unknown. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00370, 3002806535-2022-00371. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15359400&pc=NRA&device_sequence_no=1
3002806535-2022-00373 | 08/04/2022 | 09/05/2022 | 32 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF MP SLOT POST CUT GD MED | KNEE INSTRUMENTS | - | It was reported: that the postoperative x-ray showed that there was a foreign body in the joint cavity as result of the pin from the instrument detaching during the procedure. The patient underwent surgery to remove the foreign body. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Foreign source: taiwan. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15355254&pc=LXH&device_sequence_no=1
3002806535-2022-00369 | 08/09/2022 | 09/05/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL RIGHT SIZE 4 | KNEE PROSTHESIS | - | It was reported that the patient had an initial right knee arthroplasty and approximately 1. 5 years later a revision surgery was performed due to the luxation of the inlay. It was also reported that the surface of the explanted product was not as smooth as expected. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source- foreign: austria. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15354614&pc=NRA&device_sequence_no=1
3002806535-2022-00368 | 07/29/2022 | 09/01/2022 | 34 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 28MM B | G7 ACETABULAR CUP SYSTEM - CERAMIC | - | It was reported, that the patient underwent a hip replacement surgery and approximately 2 years later a revision surgery was performed due to ceramic liner fracture. Attempts have been made and no further information has been provided. | (b)(4). Foreign: japan. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15333817&pc=LZO&device_sequence_no=1
3002806535-2022-00366 | 08/02/2022 | 08/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 6 PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of a knee prosthesis due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products list: item reference:161469, item name: oxf twin-peg cmntd fem md pma, lot number:216590. Item reference:154722, item name: oxf uni tib tray sz c lm pma, lot number:169180. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00364, 3002806535-2022-00365. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15317904&pc=NRA&device_sequence_no=1
3002806535-2022-00363 | 05/10/2022 | 08/30/2022 | 112 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RHK 12X80 CEMENTED STEM | PROSTHESIS, KNEE | - | It was reported that a patient underwent a total knee arthroplasty. Approximately eight years and two months post implantation, the patient underwent revision surgery due to loosening of the femoral components with a periprosthetic fracture of the femoral condyle. During the revision procedure, the surgeon noted synovitis and metallosis and a dislocation of the tibial plateau to the offset adapter. All components were revised. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign-germany. Item name: rhk 12x80 cemented stem item#: 159405 lot#: 2084361. Item name: biomet tib block 6mm 63 item#: 141741 lot#: 639830. Item name: biomet tib block 6mm 63 item#: 141741 lot#: 203910. Item name: vngd sskpsc tib brg s 12x63/67 item#: 183822 lot#: 486500. Item name: agc mod tib locking screw item#: 153103 lot#: 089550. Item name: offset tib tray 5. 0mm adaptor item#:141491 lot#: 551170. Item name: biomet offset tibial tray 63mm item#: 141481 lot#: 640750. Item name: vngd dist fem aug 60x5 rl/lm item#: 184102 lot#: 558410. Item name: vngd dist fem aug 60x5 ll/rm item#:184122 lot#: 346530. Item name: vngd ssk psc intlk fmrl 60 rt item#: 183302 lot#: 667480. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15317903&pc=HRY&device_sequence_no=1
3002806535-2022-00365 | 08/02/2022 | 08/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ C LM PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of a knee prosthesis due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Associated products list: item reference:161469, item name: oxf twin-peg cmntd fem md pma, lot number:216590. Item reference:159550, item name: oxf anat brg lt md size 6 pma, lot number:658360. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00364, 3002806535-2022-00366. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15317843&pc=NRA&device_sequence_no=1
3002806535-2022-00364 | 08/02/2022 | 08/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that patient underwent revision of knee prosthesis due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: associated products list: item reference:154722, item name: oxf uni tib tray sz c lm pma, lot number:169180. Item reference:159550, item name: oxf anat brg lt md size 6 pma lot number:658360. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00365 and 3002806535-2022-00366. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15317838&pc=NRA&device_sequence_no=1
3002806535-2022-00362 | 04/19/2022 | 08/23/2022 | 126 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN EXCEED CUP 54 GENERIC | HIP PROSTHESIS | - | It was reported, that: the patient has undergone revision surgery due to recurring dislocations of freedom constrained liner in an exceed size 54 cup. This has been treated with closed reduction, but it is decided to revise the liner to a different constrained liner solution, to solve the luxation. When opening the patient, the head is in the liner, but when trying to luxate the head out of the cup the liner comes out of the socket of the cup and seems like it was not locked into the cup. Furthermore, the head has many visual scratches when inspected and it seems that there is a deformation of the inner rim of the liner. The old locking ring is taken out and replaced with a new one and the r/c ii constrained liner is inserted in the socket and locked in, checking that the liner is locked. A new 32mm cocr head t1 is attached to the stem and together with the locking ring on the constrained liner attached in the constrained liner. The surgeon wants to have the liner inspected as he doesn¿t find a connection between the dislocations and the findings when the patient is operated on. | (b)(4). Concomitant medical products: freedom constr hd 36mm t1 std, item number: 11-107018, lot number: 224690, freedom ringloc -x +5 54/36mm, item number: 11-253654, lot number: 6491327. Foreign source: denmark. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00227, 3002806535-2022-00362. The head, liner, and ringloc shell locking ring were returned for evaluation. The head and liner were still assembled when received. The liner was observed to exhibit score marks and scratches on the scallops and on the outer circumference. The rim of the liner was also observed to have been scratched and deformed in one area. The head has multiple surface scratches throughout one hemisphere with what appears to be metal transfer being evident also. The locking ring was observed to be over-compressed at the split point and exhibited a score mark at the split point. Device is used for treatment. The investigation confirmed that the products used during surgery are compatible and a legally marketed combination of products. This combination would not be expected to contribute to the reported issue. The revision surgery report states that when attempting to luxate the femoral component during surgery, the whole construction came out (stem/head/insert). There is a small wear mark on the liner and the head is scratched in several places. The locking ring is stuck at one side while it¿s mobile (as it should be) at the other side. The stem is fixed well. Radiographs were provided and reviewed by a health care professional. Dislocation was confirmed. No contributing factors were detected. Based on the investigation and provided information, a definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15275051&pc=JDH&device_sequence_no=1
3002806535-2022-00361 | - | 08/19/2022 | - | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD IMPLANTS | UNKNOWN KNEE PROSTHESIS | - | It was reported that the wrong implants have been opened because the identification of implants is not clear enough. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign - canada. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Part and lot identification are necessary for review of device history records, neither were provided device is used for treatment. A definitive root cause cannot be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15258256&pc=HRY&device_sequence_no=1
3002806535-2022-00360 | - | 08/19/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BEARING LARGE RIGHT SIZE 4 | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial knee arthroplasty and approximately seven weeks post-implantation the bearing dislocated. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxford ph3 cementless fem sz l item# 154927 lot# 7134937. Oxf uni cmntls tib sz d rm item# 166577 lot# 6949214. Report source: foreign - united kingdom. Investigation of this incident is currently ongoing. A followup/ final report will be submitted when additional information becomes available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15256605&pc=NRA&device_sequence_no=1
3002806535-2022-00359 | 07/20/2022 | 08/18/2022 | 29 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI PH3 TIB IMPACTOR | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that during surgery the black tip of the tibial impactor broke. No information about patient harm has been reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | Cmp-(b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15247150&pc=HWA&device_sequence_no=1
3002806535-2022-00356 | 06/24/2022 | 08/17/2022 | 54 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD UNCEMENTED TIBIAL IMPACTOR | OXFORD UNCEMENTED KNEE SYSTEM INSTRUMENTS | - | It was reported that during the implantation, a small part of the instrument broke off inside the patient. This was only discovered in a post-op x-ray. The patient was returned to theatre and the surgeon removed the remnants of the instrument. Due diligence is in progress for this complaint; to date no additional information or product has been received. | Cmp-(b)(4). Foreign - united kingdom. Investigation of this incident is currently ongoing. A followup/ final report will be submitted when additional information becomes available. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15238961&pc=HRY&device_sequence_no=1
3002806535-2022-00357 | 07/09/2022 | 08/17/2022 | 39 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN CERAMIC HEAD | HIP PROSTHESIS | - | It was reported that patient underwent revision of hip implant due to dislocation, which occurred due to a fall. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign-switzerland. Associated products: item number: unknown, item name: unknown liner, lot number: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15238805&pc=KWA&device_sequence_no=1
3002806535-2022-00355 | 07/18/2022 | 08/16/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN-PEG CMNTD FEM MD PMA | KNEE PROSTHESIS | - | It was reported that approximately six years after initial implantation, the patient underwent revision from partial to total knee prosthesis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item:154725, item name: oxf uni tib tray sz d rm pma, lot number: 306450; item:159577, item name: oxf anat brg rt md size 5 pma, lot number: 685630. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00353; 3002806535-2022-00354. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15232055&pc=NRA&device_sequence_no=1
3002806535-2022-00354 | 07/18/2022 | 08/16/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ D RM PMA | KNEE PROSTHESIS | - | It was reported that approximately six years after initial implantation, the patient underwent revision from partial to total knee prosthesis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item:159577, item name: oxf anat brg rt md size 5 pma, lot number: 685630, item:161469, item name: oxf twin-peg cmntd fem md pma, lot number: 221770. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00353, 3002806535-2022-00355. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15232043&pc=NRA&device_sequence_no=1
3002806535-2022-00353 | 07/18/2022 | 08/16/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT MD SIZE 5 PMA | KNEE PROSTHESIS | - | It was reported that approximately six years after initial implantation, the patient underwent revision from partial to total knee prosthesis. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: item:154725, item name: oxf uni tib tray sz d rm pma, lot number:306450. Item:161469, item name: oxf twin-peg cmntd fem md pma, lot number: 221770. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00354, 3002806535-2022-00355. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15232038&pc=NRA&device_sequence_no=1
3002806535-2022-00350 | 07/08/2022 | 08/10/2022 | 33 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | PROSTHESIS, KNEE | - | It was reported that the patient was revised due to implant fracture. | (b)(4). Foreign: germany. Medical product unknown. Oxford twin peg fem medium. Catalog#: unknown, lot#:unknown. Medical product unknown. Oxford tib catalog#: unknown. Lot#:unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15197804&pc=NRA&device_sequence_no=1
3002806535-2022-00348 | 07/11/2022 | 08/09/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 28MM | PROSTHESIS, HIP, SEMI-CONSTRAINED | - | It was reported patient underwent a hip revision approximately 5 years post implantation due to a dislocation. The head and liner were replaced. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Cat# 110024465 lot# 985900 g7 dual mobility liner 46mm g. Cat# 650-1067 lot# 2900844 cer option type 1 tpr sleve. Cat# ep-200152 lot# 038300 act artic e1 hip brg 28x46mm s52 dia28. Customer has indicated that the product is in process of being returned to zimmer biomet for investigation. Once the investigation has been completed, a follow-up mdr will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15190756&pc=LZO&device_sequence_no=1
3002806535-2022-00347 | 07/12/2022 | 08/08/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE AA | KNEE PROSTHESIS | - | It was reported, that: the patient underwent a revision of left knee due to unknown reasons, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10: oxf anat brg lt x-sm sz 5 pma item# 159792 lot# 298510. Oxf twin-peg cmntd fem xs pma item# 161467 lot# 778210. Oxford resection 3pk hall hub item# 506078 lot# 778210. Oxf uni fm/cm tb/uni item# 98-b002-006-01 lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00345, 3002806535 - 2022 - 00346. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15182471&pc=NRA&device_sequence_no=1
3002806535-2022-00346 | 07/12/2022 | 08/08/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS TWIN PEG FEMUR - XSMALL | KNEE PROSTHESIS | - | It was reported, that: the patient underwent a revision of left knee due to unknown reasons, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10: oxf anat brg lt x-sm sz 5 pma item# 159792 lot# 298510. Oxf uni tib tray sz aa lm pma item# 159531 lot# 487710. Oxford resection 3pk hall hub item# 506078 lot# 778210. Oxf uni fm/cm tb/uni item# 98-b002-006-01 lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00345, 3002806535 - 2022 - 00347. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15182423&pc=NRA&device_sequence_no=1
3002806535-2022-00345 | 07/12/2022 | 08/08/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - XSMALL LEFT SIZE 5 | KNEE PROSTHESIS | - | It was reported, that: the patient underwent a revision of left knee due to unknown reasons, approximately two (2) years post-op. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). D10: oxf twin-peg cmntd fem xs pma item# 161467 lot# 778210. Oxf uni tib tray sz aa lm pma item# 159531 lot# 487710. Oxford resection 3pk hall hub item# 506078 lot# 778210. Oxf uni fm/cm tb/uni item# 98-b002-006-01 lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00346, 3002806535 - 2022 - 00347. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15182392&pc=NRA&device_sequence_no=1
3002806535-2022-00344 | 06/27/2022 | 08/08/2022 | 42 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | INTRAMEDULLARY ROD REMOVAL HOOK WITH T-HANDLE | KNEE ARTHROPLASTY | - | It was reported that the instrument was broken. No impact or consequence to the patient have been reported. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: foreign: germany. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15182367&pc=HWA&device_sequence_no=1
3002806535-2022-00342 | 05/17/2022 | 08/03/2022 | 78 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP | - | It was reported that the patient was revised approximately 9 months post implantation due to femoral head implant fracture. Attempts have been made and no further information has been provided. | (b)(4). Foreign- argentina. Item name: g7 screw 6. 5mm x 25mm; item# 010000998; lot# 6861266; item name: g7 pps ltd acet shell 58g; item# 010000666; lot# 6883595; g7 neutral arcomxl lnr 36mm g; item# 010000742; lot# 6628991; taperloc por lat fmrl 12. 5x145; item# 11-103206; lot# 939270; investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15158811&pc=LZO&device_sequence_no=1
3002806535-2022-00343 | 07/06/2022 | 08/02/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BI-POLAR 28 CUP 50MM | HIP PROSTHESIS | - | It was reported, that the sap system identified that one (1) expired product has been implanted approximately three months after the sterile expiration date. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign spain. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of the device history record(s) identified no deviations or anomalies during manufacturing. Device is used for treatment. Medical records were not provided. The root cause of the reported issue is attributed to user issuing expired product. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15144398&pc=JDI&device_sequence_no=1
3002806535-2022-00341 | 05/17/2019 | 07/29/2022 | 1169 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA CER LNR 32MM G | HIP PROSTHESIS | - | It was reported that the patient underwent initial left hip arthroplasty. The patient experienced dislocation and subluxation which resulted in a closed reduction approximately 7 months post-op. Subsequently, 5 months later the patient underwent a revision due to dislocation. | (b)(4). Concomitant products: delta cer fem hd 32/- 3mm t1 item# 650-1163 lot# 2015050213; echo por fmrl lat nc 12x140mm. Item# 192112 lot #478700; g7 pps ltd acet shell 58g. Item# 010000666 lot #: 39061720. Report source: foreign ¿ sweden. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. The reason for dislocation could not be established based on the provided radiographs. The manufacturing history records for the biolox delta femoral head and acetabular liner have been checked and verified that the components were manufactured and sterilized in accordance with the applicable specifications. A definitive root cause cannot be determined. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2019-00500. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15129330&pc=LZO&device_sequence_no=1
3002806535-2022-00337 | 07/07/2022 | 07/27/2022 | 20 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIBIAL IMPACTOR | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that the black plastic impactor tip split in a half upon impaction. No known harm or consequences to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign source: canada. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15110822&pc=LXH&device_sequence_no=1
3002806535-2022-00339 | 06/27/2022 | 07/26/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL TRAY | KNEE PROSTHESIS | - | It was reported that patient underwent the knee revision surgery due to bone fracture caused by trauma. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Source - foreign - united kingdom. Medical products: unknown oxford bearing item#: unknown, lot#: unknown; unknown oxford femoral component item#: unknown, lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00338 and 3002806535-2022-00340. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15102899&pc=HSA&device_sequence_no=1
3002806535-2022-00340 | 06/27/2022 | 07/26/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that patient underwent the knee revision surgery due to bone fracture caused by trauma. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source: (b)(4). Unknown oxford tibial tray item#: unknown, lot#: unknown. Unknown oxford bearing item#: unknown, lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00338, 3002806535 - 2022 - 00339, 3002806535 - 2022 - 00340. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15102888&pc=HSA&device_sequence_no=1
3002806535-2022-00338 | 06/27/2022 | 07/26/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that patient underwent the knee revision surgery due to bone fracture caused by trauma. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Source - foreign - united kingdom. Medical products: unknown oxford tibial tray item#: unknown, lot#: unknown; unknown oxford femoral component item#: unknown, lot#: unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00339 and 3002806535-2022-00340. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15102873&pc=HSA&device_sequence_no=1
3002806535-2022-00335 | - | 07/25/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL TRAY | KNEE PROSTHESIS | - | It was reported that the patient underwent a left partial oxford knee surgery approximately eight years prior. Subsequently, the patient stated he recently started experiencing pain and weakness in the leg. No revision surgery was planned. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00334 and 3002806535-2022-00336. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15092239&pc=JWH&device_sequence_no=1
3002806535-2022-00334 | - | 07/25/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported that the patient underwent a left partial oxford knee surgery approximately eight years prior. Subsequently, the patient stated he recently started experiencing pain and weakness in the leg. No revision surgery was planned. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022¿00334, 3002806535-2022¿00335, 3002806535-2022¿00336. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15092238&pc=JWH&device_sequence_no=1
3002806535-2022-00336 | - | 07/25/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a left partial oxford knee surgery approximately eight years prior. Subsequently, the patient stated he recently started experiencing pain and weakness in the leg. No revision surgery was planned. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00334 and 3002806535 - 2022 ¿ 00335. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15092225&pc=JWH&device_sequence_no=1
3002806535-2022-00332 | 06/29/2022 | 07/25/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM LEFT SIZE 4 | KNEE ARTHROPLASTY | - | It was reported, the patient had a partial knee arthroplasty implantation surgery and approximately 8 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: oxf twin-peg cmntd fem md pma; item# 161469; lot# 703020; oxf uni tib tray sz c lm pma; item# 154722; lot# 810720; oxford resection 3pk strk 6 h; item# 506124; lot# 948229. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00331 and 3002806535-2022-00333. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15091774&pc=NRA&device_sequence_no=1
3002806535-2022-00333 | 06/29/2022 | 07/25/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE C | KNEE ARTHROPLASTY | - | It was reported, the patient had a partial knee arthroplasty implantation surgery and approximately 8 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: oxf twin-peg cmntd fem md pma; item# 161469; lot# 703020. Oxf anat brg lt md size 4 pma; item# 159548; lot# 376980. Oxford resection 3pk strk 6 h; item# 506124; lot# 948229. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022¿00331 and 3002806535-2022¿00332. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15091773&pc=NRA&device_sequence_no=1
3002806535-2022-00331 | 06/29/2022 | 07/25/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS TWIN PEG FEMUR - MEDIUM | KNEE ARTHROPLASTY | - | It was reported, the patient had a partial knee arthroplasty implantation surgery and approximately 8 years later a revision surgery was performed due to unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: oxf anat brg lt md size 4 pma; item# 159548; lot# 376980. Oxf uni tib tray sz c lm pma; item# 154722; lot# 810720. Oxford resection 3pk strk 6 h; item# 506124; lot# 948229. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00332, 3002806535 - 2022 ¿ 00333. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15091772&pc=NRA&device_sequence_no=1
3002806535-2022-00330 | 06/29/2022 | 07/25/2022 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | QUICK RELEASE FEMORAL DRILL BIT, 4MM DIAMETER | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that the drill bit broke during use in the drill guide. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - medical devices: oxf sawblade stryker cmntd 3pk; item# 506298; lot# 395082. Oxf anat brg lt sm size 3 pma; item# 159540; lot# 935850. Oxf twin-peg cmntd fem sm pma; item# 161468; lot# 363390. Oxf uni tib tray sz aa lm pma; item# 159531; lot# 783780. Oxf uni fm/cm tb/uni; item# 98-b002-006-01; lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15091770&pc=HTW&device_sequence_no=1
3002806535-2022-00326 | 06/21/2022 | 07/22/2022 | 31 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM +6MM NECK, TYPE 1 TAPER | HIP PROSTHESIS | - | It was reported that patient underwent revision due to the infection approximately three (3) months post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15082234&pc=LZO&device_sequence_no=1
3002806535-2022-00328 | 06/24/2022 | 07/22/2022 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA OPTION HEAD, 28MM | HIP PROSTHESIS | - | It was reported that the ceramic head chipped after being assembled with the bearing and taper sleeve on the bearing compressor. No consequences or impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products; vivacit-e dm bearing 28x42mm item#: 110031011 lot#: 65254540 cer option type 1 tpr sleve +3 item#: 650-1067 lot#: 3087988. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15082129&pc=LZO&device_sequence_no=1
3002806535-2022-00329 | 01/29/2019 | 07/22/2022 | 1270 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TINBN COATED VANGUARD INTERLOK PS ANATOMIC FEMORAL COMPONENT 70MM LEFT | KNEE PROSTHESIS | - | It was reported via a clinical study that a patient underwent an initial left total knee arthroplasty. Approximately one year and two months post implantation, the patient underwent revision surgery due to pain. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Concomitant medical products: vngd ps tib brg 12x71/75mm; item#: 183642; lot#: 430510; tnbn pol cocr fin tib tray 71; item#141253tnbn; lot#: 2017051899; tibial stem; item#: unknown; lot#: unknown; refobacin bone cement r; item#: unknown; lot#: unknown; patella component; item#: unknown; lot#: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15082019&pc=JWH&device_sequence_no=1
3002806535-2022-00322 | - | 07/22/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/+6MM | PROSTHESIS, HIP | - | It was reported that the patient was revised due to unknown reasons approximately 3 months post implantation. Attempts have been made and additional information on the reported event is unavailable. | (b)(4). Initial surgery: (b)(6) 2021; revision surgery: (b)(6) 2021. Report source: source: australia. Concomitant medical products: part : 010000704, lot: 6721086, brand: acetabular shell; part: 010000936, lot: 6848116, brand: acetabular liner; part: 51-114180, lot: 6820522, brand: femoral stem. Customer has indicated that the product will not be returned to zimmer biomet for investigation. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15081659&pc=LZO&device_sequence_no=1
3002806535-2022-00325 | 10/12/2018 | 07/22/2022 | 1379 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI TIB TRAY SZ B RM PMA | KNEE PROSTHESIS | - | It was reported that the patient underwent a partial knee replacement. Subsequently, the patient was revised due to loosening of the tibial component, approximately 10 months post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: catalog number: 159575, item name: oxf anat brg rt md size 3 pma, lot number: 365570; catalog number: 161469, item name: oxf twin-peg cmntd fem md pma, lot number: 502190; catalog number: unknown cement, item name: unknown, lot number: unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15081652&pc=NRA&device_sequence_no=1
3002806535-2022-00318 | 04/16/2019 | 07/20/2022 | 1191 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | TINBN COATED VANGUARD INTERLOK PS ANATOMIC FEMORAL COMPONENT 60MM LEFT | VANGUARD COMPLETE KNEE SYSTEM | - | It was reported the patient underwent a left knee revision approximately a year-and-a-half post-implantation due to pain and limited range of motion. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical device: vngd ps tib brg 10x63/67mm; item# 183620; lot# 835790; unknown femoral stem; item# unknown; lot# unknown; unknown bone cement; item# unknown; lot# unknown; tnbn pol cocr fin tib tray 63; item# 141251tnbn; lot# 2017070093. Source foreign: italy. Due diligence is in progress for this complaint; to date no additional information or product has been received. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15062223&pc=JWH&device_sequence_no=1
3002806535-2022-00321 | 04/30/2022 | 07/19/2022 | 80 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - RM SIZE D | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the right knee approximately three (3) months post implantation due to pain, swelling and limitation of movement after extreme flexion of the knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf anat brg rt x-sm 4mm pma item# 160791 lot# 6741836; oxford uni twin-peg femoral xs item# 166940 lot# 6839631. Report source: foreign - china. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00319; 3002806535 - 2022 ¿ 00320. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15052343&pc=NRA&device_sequence_no=1
3002806535-2022-00320 | 04/30/2022 | 07/19/2022 | 80 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS HIGH FLEX TWIN PEG CEMENTED FEMUR - XSMALL | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the right knee approximately three (3) months post implantation due to pain, swelling and limitation of movement after extreme flexion of the knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf anat brg rt x-sm 4mm pma item# 160791 lot# 6741836. Oxf uni tib tray sz d rm pma item# 154725 lot# 6937758. Report source: foreign - china. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00319, 3002806535 - 2022 ¿ 00321. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15052330&pc=NRA&device_sequence_no=1
3002806535-2022-00319 | 04/30/2022 | 07/19/2022 | 80 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - XSMALL RIGHT SIZE 4 | KNEE PROSTHESIS | - | It was reported that the patient underwent a revision procedure of the right knee approximately three (3) months post implantation due to pain, swelling and limitation of movement after extreme flexion of the knee. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxford uni twin-peg femoral xs item# 166940 lot# 6839631, oxf uni tib tray sz d rm pma item# 154725 lot# 6937758. Foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00319, 3002806535-2022-00320, 3002806535-2022-00321. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15052326&pc=NRA&device_sequence_no=1
3002806535-2022-00313 | 06/22/2022 | 07/19/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a total knee arthroplasty surgery and approximately twenty five years later a revision surgery was performed due to bearing wear and loosening of the stem. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: unknown oxford femoral component; item# unknown; lot# unknown; unknown oxford bearing component; item# unknown; lot# unknown. (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00311, 3002806535-2022-00312, 3002806535-2022-00313. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15050006&pc=HRY&device_sequence_no=1
3002806535-2022-00312 | 06/22/2022 | 07/19/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a total knee arthroplasty surgery and approximately twenty five years later a revision surgery was performed due to bearing wear and loosening of the stem. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: unknown oxford femoral component; item# unknown; lot# unknown; unknown oxford tibial component; item# unknown; lot# unknown. (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00311, 3002806535-2022-00312, 3002806535-2022-00313. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15050005&pc=HRY&device_sequence_no=1
3002806535-2022-00311 | 06/22/2022 | 07/19/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient underwent a total knee arthroplasty surgery and approximately twenty five years later a revision surgery was performed due to bearing wear and loosening of the stem. Attempts have been made and no further information has been provided. | (b)(4). Report source- foreign: (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00312 and 3002806535 - 2022 - 00313. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15049989&pc=HRY&device_sequence_no=1
3002806535-2022-00304 | - | 07/18/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE PROSTHESIS | - | It was reported, that the patient suffered a bearing dislocation. A lidocaine injection was administered. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: oxf twin peg cmntls fmrlmd; item# 161474; lot# r2503466a; oxford cementless tibia dlm;item# us166576; lot# r3050476a. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15037550&pc=HRY&device_sequence_no=1
3002806535-2022-00310 | 05/20/2022 | 07/18/2022 | 59 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM +6MM NECK, TYPE 1 TAPER | PROSTHESIS, HIP | - | It was reported that the patient underwent a total hip arthroplasty and approximately a month later a revision surgery was performed due to infection. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical device: tprlc 133 t1 pps ho16x152mm; item# 51-104160; lot# 6299205; g7 vit e +5mm lnr 36mm g; item# 30153607; lot# 65207718; g7 pps ltd acet shell 58g; item# 010000666;lot# 7170807. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15037456&pc=LZO&device_sequence_no=1
3002806535-2022-00309 | 05/10/2022 | 07/18/2022 | 69 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE PROSTHESIS | - | It was reported that the patient had a knee replacement surgery and a revision surgery was needed for unknown reason. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: rhk 12x80 cemented stem; item# 159405; lot# 2084361. Foreign: (b)(6). An investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15037455&pc=HRY&device_sequence_no=1
3002806535-2022-00305 | 06/17/2022 | 07/15/2022 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM +6MM NECK, TYPE 1 TAPER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER | - | It was reported that during surgery the ceramic head broke while the surgeon was removing it from the stem taper. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products - medical devices: g7 vit e +5mm lnr 36mm g;item# 30153607; lot# 65056459. 3. 2mmx40mm rnglc+ acetdrl bit; item# 31-323240; lot# 937430. Trilogy bone scr 6. 5x50;item# 00-6250-065-50; lot# 63055150. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15026697&pc=LZO&device_sequence_no=1
3002806535-2022-00308 | 06/16/2022 | 07/15/2022 | 29 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | NON MODULAR TIBIAL TRAY IMPACTOR (CEMENTLESS APPLICATION) | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that the tibial impactor broke. The broken piece of the instrument remained in the patient. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15025962&pc=NRA&device_sequence_no=1
3002806535-2022-00307 | 06/20/2022 | 07/15/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 28MM +3.5MM NECK, 12/14 TAPER | HIP PROSTHESIS | - | It was reported that expired product has been implanted. No known impact to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Source foreign: italy. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15025898&pc=LZO&device_sequence_no=1
3002806535-2022-00306 | 05/10/2022 | 07/15/2022 | 66 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RHK 12 X 80 CEMENTED STEM | KNEE PROSTHESIS | - | It was reported that during removal of the tibial component, the tibial plateau broke off from the stem. Two holes had to be drilled into the prosthesis to remove the stem from the bone. Damage to the bone occurred while removing the stem. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15025840&pc=JWH&device_sequence_no=1
3002806535-2022-00303 | 06/03/2022 | 07/15/2022 | 42 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | PROSTHESIS, HIP | - | It was reported that the patient experienced dislocation due to improper motion approximately 1 week post implantation. The patient's hip was reduced at ed without incident. Due diligence is in progress for this complaint; to date no additional on the reported event is unavailable at this time. | (b)(4). Concomitant medical products: g7 pps ltd acet shell 62h, 010000668; 6906549, echo por fmrl lat nc 13x145mm; 192113; 043840, g7 longevity neutral 36mm h; 20103608; 64473176. Customer has indicated that the product will not be returned to zimmer biomet for investigation. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=15025472&pc=LZO&device_sequence_no=1
3002806535-2022-00302 | 03/19/2021 | 07/13/2022 | 481 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA OPTION HEAD, 28MM | PROSTHESIS, HIP | - | It was reported patient underwent initial left hip arthroplasty. Subsequently, the patient underwent a closed reduction under anesthesia approximately 2 days post implantation due to a left hip dislocation after suffering from a fall. The fall occurred while walking to the bathroom in a private residence. Additional information on the reported event is unavailable. | (b)(4). Concomitant medical products: 110024461, 186690, g7 dual mobility liner 38mm c; 110010242, 6882934, g7 osseoti 3 hole shell 48mm c; 650-1066, 3050922, cer opt type 1 tpr sleve 0mm; 51101080, 6341287, femoral stem. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14998424&pc=LZO&device_sequence_no=1
3002806535-2022-00294 | 05/14/2022 | 07/07/2022 | 54 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. s bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BEARING MEDIUM LEFT SIZE 4 | KNEE PROSTHESIS | - | It was reported the patient underwent a revision surgery of the left knee due to the bearing dislocation, approximately ten (10) months post-op. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942614&pc=NRA&device_sequence_no=1
3002806535-2022-00293 | 06/07/2022 | 07/07/2022 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RINGLOC-X 10 DEGREE ARCOM XL LINER 46/28MM (22) | HIP PROSTHESIS | - | It was reported that expired product has been implanted. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Foreign- poland. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942605&pc=JDI&device_sequence_no=1
3002806535-2022-00296 | 06/10/2022 | 07/07/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTLESS TIBIA LEFT ALPHA POROUS/HA + EXTRACTOR SLOT RIGHT MEDIAL S | KNEE PROSTHESIS | - | It was reported that the patient was revised from a partial knee to a total knee due to tibial collapse, approximately two (2) months post-op. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg rt md size 4 pma item#159576 lot#7123716, oxford ph3 cementless fem sz m item#154926 lot#7126473. Foreign: netherlands. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942470&pc=NRA&device_sequence_no=1
3002806535-2022-00297 | 01/04/2020 | 07/07/2022 | 915 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM RIGHT SIZE 3 | KNEE PROTHESIS | - | It was reported that the patient underwent a knee arthroplasty and developed infection within the same year. Revision surgery was performed and the bearing was changed. Investigation results are now available. Attempts have been made and no further information has been provided. | (b)(4). Medical devices: oxford ph3 cementlessfem sz m; item# 154926; lot# 6299751. Oxf uni cmntls tib sz d rm; item# 166577, lot# 6181847. Foreign: singapore. No product was returned or pictures provided; visual and dimensional evaluations could not be performed. Review of the device history records identified no deviations or anomalies during manufacturing that could be related to the reported event. The device is used for treatment. Medical records were not provided. A definitive root cause cannot be determined. It has been confirmed that the implant is not within the scope or subject of any field actions or recalls which could be attributed to reported event. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942310&pc=NRA&device_sequence_no=1
3002806535-2022-00285 | 06/10/2022 | 07/07/2022 | 27 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BEARING TRIAL WITH SLOTS, SMALL, 3MM | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported, that during the surgery the surgeon was trialing and when he removed the trial bearing a small part of it chipped. No delay or reported health consequences to patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14942193&pc=NRA&device_sequence_no=1
3002806535-2022-00298 | 03/25/2022 | 07/06/2022 | 103 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | RINGLOC-X 10 DEGREE ARCOM XL LINER 50/32MM (23) | HIP PROSTHESIS | - | It was reported that expired product has been implanted. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Source foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14927237&pc=JDI&device_sequence_no=1
3002806535-2022-00292 | 01/27/2022 | 07/06/2022 | 160 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS C'LESS TIBIA ALPHA POROUS/HA + EXTRACTOR SLOT - RM SIZE D | KNEE PROTHESIS | - | It was reported that a partial knee arthroplasty surgery was performed and approximately 6 years later the patient is still experiencing intermittent pain and swelling. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin peg cmntls fmrl md; item# 161474; lot# r3050584a. Oxf anat brg rt mdsize 3 pma; item# 159575; lot# 816870. Multiple mdr reports were filed for this event, please see associated reports 3002806535 - 2022 - 00290; 3002806535 - 2022 - 00291; 3002806535 - 2022 - 00292. The device history records for were reviewed for deviations and / or anomalies and no discrepancies related to the reported event were found. Review of x-ray images (6 week post-op and 3 year post-op) show overall fit, alignment, and bone quality appear normal. There are no signs of loosening, wear, radiolucency, or other contributing factors. Soft tissues are normal, no hardware or bone fractures present. There are no other concerns noted that would lead to the patient's complaint of pain/catching sensation. The root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14926783&pc=NRA&device_sequence_no=1
3002806535-2022-00290 | 01/27/2022 | 07/06/2022 | 160 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN PEG CMNTLS FMRL MD | KNEE PROTHESIS | - | It was reported that a partial knee arthroplasty surgery was performed and approximately 6 years later the patient is still experiencing intermittent pain and swelling. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf anat brg rt mdsize 3 pma; item# 159575; lot# 816870; oxf uni cmntls tib szd rm; item# (b)(4); lot# r3050482a. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00290, 3002806535-2022-00291, 3002806535-2022-00292. The device history records for were reviewed for deviations and / or anomalies and no discrepancies related to the reported event were found. Review of x-ray images (6week post-op and 3 year post-op) show overall fit, alignment, and bone quality appear normal. There are no signs of loosening, wear, radiolucency, or other contributing factors. Soft tissues are normal, no hardware or bone fractures present. There are no other concerns noted that would lead to the patient's complaint of pain/catching sensation. The root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14926775&pc=NRA&device_sequence_no=1
3002806535-2022-00291 | 01/27/2022 | 07/06/2022 | 160 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM RIGHT SIZE 3 | KNEE PROTHESIS | - | It was reported that a partial knee arthroplasty surgery was performed and approximately 6 years later the patient is still experiencing intermittent pain and swelling. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical devices: oxf twin peg cmntls fmrl md; item# (b)(4); lot# r3050584a; oxf uni cmntls tib szd rm; item# (b)(4); lot# r3050482a. Multiple mdr reports were filed for this event, please see associated reports: multiple mdr reports were filed for this event, please see associated reports 3002806535-2022-00290; 3002806535-2022-00291; 3002806535-2022-00292. The device history records for item # 159575, lot # 816870 were reviewed for deviations and/or anomalies with no anomalies/deviations identified. Review of x-ray images (6week post-op and 3 year post-op) show overall fit, alignment, and bone quality appear normal. There are no signs of loosening, wear, radiolucency, or other contributing factors. Soft tissues are normal, no hardware or bone fractures present. There are no other concerns noted that would lead to the patient's complaint of pain/catching sensation. The root cause of the event was unable to be determined. If any further information is found which would change or alter any conclusions or information, a supplemental report will be filed accordingly. Zimmer biomet will continue to monitor for trends. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14926773&pc=NRA&device_sequence_no=1
3002806535-2022-00295 | 06/04/2022 | 07/01/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 32/ 0MM T1 | PROSTHESIS, HIP | - | It was reported the patient underwent a hip revision due to dislocation approximately two days post implantation. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: g7 hi-wall e1 liner 32mm f; item# 010000928; item# 7014294; tprlc xr t1 pps 10x140mm; item 51-105100; item# 7057079. Source foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14884486&pc=LZO&device_sequence_no=1
3002806535-2022-00289 | - | 06/30/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN RECAP HIP SYSTEM | HIP PROTHESIS | - | It was reported that a claim for damages has been submitted, allegedly caused by a mom implant. Metallosis is mentioned but not yet confirmed. The patient walks with brachial braces and has partially recovered. The item remains implanted. Due diligence is in progress for this complaint. No additional information is available at this time. | (b)(4). Source: foreign: (b)(6). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14868636&pc=KXA&device_sequence_no=1
3002806535-2022-00278 | 05/30/2022 | 06/29/2022 | 30 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | NON MODULAR TIBIAL TRAY IMPACTOR (CEMENTLESS APPLICATION) | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported that during the unicompartmental knee surgery the instrument's back tab wasn't fixed anymore. The surgery was prolonged by 5 minutes and completed with another device. No other known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Report source - foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14851619&pc=NRA&device_sequence_no=1
3002806535-2022-00287 | - | 06/29/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTLESS TIBIA LEFT ALPHA POROUS/HA + EXTRACTOR SLOT LEFT MEDIAL SI | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial left knee arthroplasty and approximately two (2) years post-implantation is still experiencing pain and instability. Due diligence is in progress for this complaint; to date no additional information has been received. | (b)(4). Concomitant medical products: oxf anat brg lt md size 3 pma, item#159547, lot#6483393; oxford uni twin-peg femoral md, item#166942, lot#j6466714; palacos r+g 2x40g, item #66017569, lot #92804884. Report source foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022¿00286, 3002806535-2022¿00287, 3002806535-2022¿00288. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14849602&pc=NRA&device_sequence_no=1
3002806535-2022-00288 | - | 06/29/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS HIGH FLEX TWIN PEG CEMENTED FEMUR MEDIUM | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial left knee arthroplasty and approximately two (2) years post-implantation is still experiencing pain and instability. Due diligence is in progress for this complaint; to date no additional information has been received. | (b)(4). Concomitant medical products: oxf anat brg lt md size 3 pma item#159547 lot#6483393, oxf uni cmntls tib sz c lm item#166574 lot#3922603, palacos r+g 2x40g item #66017569 lot #92804884. Report source: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00286, 3002806535 - 2022 ¿ 00287. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14849596&pc=NRA&device_sequence_no=1
3002806535-2022-00286 | - | 06/29/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BEARING MEDIUM LEFT SIZE 3 | KNEE PROSTHESIS | - | It was reported that the patient underwent an initial left knee arthroplasty and approximately two (2) years post-implantation is still experiencing pain and instability. Due diligence is in progress for this complaint; to date no additional information has been received. | (b)(4). Concomitant medical products: oxf uni cmntls tib sz c lm item#166574 lot#3922603, oxford uni twin-peg femoral md item#166942 lot#j6466714, palacos r+g 2x40g item #66017569 lot #92804884. Report source: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 ¿ 00287, 3002806535 - 2022 ¿ 00288. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14849593&pc=NRA&device_sequence_no=1
0009613350-2022-00338 | 05/10/2022 | 06/28/2022 | 49 | Injury | Manufacturer | BIOMET UK LTD. sulzer allee 8 sulzer industrie park winterthur 8404 SZ 8404 | BIOMET UK LTD. sulzer allee 8 sulzer industrie park winterthur 8404 SZ 8404 | UNKNOWN CERAMIC INSERT | HIP PROSTHESIS | - | It was reported the patient underwent a right hip revision approximately 6 years post implantation due to implant breakage while walking accompanied with pain, noise, and sensation of sand. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: unknown ceramic head; item# unknown; lot# unknown, 56mm o. D. Size kk porous uncemented with uni-hole shell use with kk liners; item# 00875705600; lot# 62945617. Report source: (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 0009613350-2022-00337, 0009613350-2022-00338. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14841587&pc=LZO&device_sequence_no=1
3002806535-2022-00284 | 05/30/2022 | 06/28/2022 | 29 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | INTRAMEDULLARY ROD REMOVAL HOOK WITH T-HANDLE | KNEE ARTHROPLASTY | - | It was reported that while trying to remove intramedullary rod from femur with the t-handle, the tip of the t-handle broke. No further patient outcome provided. Attempts have been made and no further information has been made available at this time. Due diligence is in progress for this complaint. | (b)(4). Foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14840866&pc=LXH&device_sequence_no=1
0009613350-2022-00337 | 05/10/2022 | 06/28/2022 | 49 | Injury | Manufacturer | BIOMET UK LTD. sulzer allee 8 sulzer industrie park winterthur 8404 SZ 8404 | BIOMET UK LTD. sulzer allee 8 sulzer industrie park winterthur 8404 SZ 8404 | UNKNOWN CERAMIC HEAD | HIP PROSTHESIS | - | It was reported the patient underwent a right hip revision approximately 6 years post implantation due to implant breakage while walking accompanied with pain, noise, and sensation of sand. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical devices: unknown ceramic insert; item# unknown; lot# unknown. 56mm o. D. Size kk porous uncemented with uni-hole shell use with kk liners; item# 00875705600; lot# 62945617. Report source - foreign: france. Multiple mdr reports were filed for this event, please see associated reports: 0009613350-2022-00338. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14839719&pc=LZO&device_sequence_no=1
3002806535-2022-00283 | 05/23/2022 | 06/24/2022 | 32 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL RIGHT SIZE 4 | KNEE ARTHROPLASTY | - | It was reported that the patient underwent a revision surgery due to instability approximately 4 years after implantation. The patient has been revised from a 4mm bearing to a 6mm bearing. No additional information is available at this time. | (b)(4). Report source foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14796632&pc=NRA&device_sequence_no=1
3002806535-2022-00282 | 05/20/2022 | 06/24/2022 | 35 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI PH3 TIB IMPACTOR | KNEE ARTHROPLASTY | - | It was reported that the pin that holds the black pad broke during impaction. The event occurred during surgery and no harm to the patient is reported. The surgeon was able to complete the procedure with the same device. No further patient outcome provided. | (b)(4). Report source: foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14796581&pc=HWR&device_sequence_no=1
3002806535-2022-00276 | 05/23/2022 | 06/17/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE C | KNEE ARTHROPLASTY | - | It was reported that the patient has been revised due to reoccurring hemarthrosis approximately 3. 5 years after the initial surgery. No additional information is available at this time. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00275, 3002806535-2022-00277. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14734133&pc=NRA&device_sequence_no=1
3002806535-2022-00275 | 05/23/2022 | 06/17/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS TWIN PEG FEMUR - SMALL | KNEE ARTHROPLASTY | - | It was reported that the patient has been revised due to reoccurring hemarthrosis approximately 3. 5 years after the initial surgery. No additional information is available at this time. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00276, 3002806535-2022-00277. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14734102&pc=NRA&device_sequence_no=1
3002806535-2022-00277 | 05/23/2022 | 06/17/2022 | 25 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL LEFT SIZE 3 | KNEE ARTHROPLASTY | - | It was reported that the patient has been revised due to reoccurring hemarthrosis approximately 3. 5 years after the initial surgery. Attempts have been made and no further information has been provided at this time. | (b)(4). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00275, 3002806535-2022-00276. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14734099&pc=NRA&device_sequence_no=1
3002806535-2022-00281 | 05/10/2022 | 06/17/2022 | 38 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD CEMENTLESS TIBIA D LM | KNEE ARTHROPLASTY | - | It was reported in a clinical study that the patient developed painful pes anserinus bursa on the left knee approximately seven years post-implantation. Patient is being treated with non-steroidal anti-inflammatory drugs. No revision or further outcome reported. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: oxf twin peg cmntls fmrl md item# 161474, lot# 3050585, oxf anat brg lt md size 3 pma med sz 3 item# 159547, lot# 825760. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00279, 3002806535 - 2022 ¿ 00280. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14730871&pc=NRA&device_sequence_no=1
3002806535-2022-00280 | 05/10/2022 | 06/17/2022 | 38 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM LEFT SIZE 3 | KNEE ARTHROPLASTY | - | It was reported in a clinical study that the patient developed painful pes anserinus bursa on the left knee approximately seven years post-implantation. Patient is being treated with non-steroidal anti-inflammatory drugs. No revision or further outcome reported. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: oxf twin peg cmntls fmrl md item# 161474 ,lot# 3050585, oxford cementless tibia d lm item# us166576, lot# 3050478. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00279, 3002806535 - 2022 - 00281. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14730861&pc=NRA&device_sequence_no=1
3002806535-2022-00279 | 05/10/2022 | 06/17/2022 | 38 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TWIN PEG CMNTLS FMRL MD | KNEE ARTHROPLASTY | - | It was reported in a clinical study that the patient developed painful pes anserinus bursa on the left knee approximately seven years post-implantation. Patient is being treated with non-steroidal anti-inflammatory drugs. No revision or further outcome reported. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical products: oxf anat brg lt md size 3 pma med sz 3 item# 159547, lot# 825760, oxford cementless tibia d lm item# us166576, lot# 3050478. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535 - 2022 - 00280, 3002806535 - 2022 - 00281. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14730785&pc=NRA&device_sequence_no=1
3002806535-2022-00274 | 05/16/2022 | 06/13/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA STANDARD, RIGHT MEDIAL SIZE D | KNEE PROSTHESIS | - | It was reported that a patient underwent an initial right knee arthroplasty. Subsequently, a revision procedure due to subsidence was performed, approximately six (6) years post-op. Attempts have been made and no further information has been provided. | (b)(4). Concomitant products: oxf twin-peg cmntd fem lg pma item#161470 lot#483330. Oxf anat brg rt lg size 4 pma item#159583 lot#418060. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14673777&pc=NRA&device_sequence_no=1
3002806535-2022-00271 | 05/10/2022 | 06/09/2022 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | HIP PROTHESIS | - | It was reported that the patient underwent a revision procedure of the right hip approximately one month post implantation due to subsidence of the femoral stem after limb length discrepancy was noted. During revision, a crack in the upper femoral shaft and the greater trochanter developed. These were repaired and the final components were replaced without further complication. Attempts have been made and additional information on the reported event is unavailable at this time. | (b)(4). Concomitant medical products: tloc 133 mp sp t1 pps ho 7x99 item#51-109070 lot#3652988. G7 pps ltd acet shell 52e item#010000663 lot#7061616. 36mm i. D. Size e neutral liner item#30103605 lot#65256574. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14641533&pc=LZO&device_sequence_no=1
3002806535-2022-00270 | 04/28/2022 | 06/09/2022 | 42 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/-3MM | HIP PROTHESIS | - | It was reported that the patient underwent a revision procedure of the left hip approximately three (3) days post implantation due to periprosthetic fracture after a fall two (2) days post implantation. The fracture was secured with a plate and cables, and the femoral head and stem were replaced without complication. Attempts have been made and additional information on the reported event is unavailable at this time. | (b)(4). Concomitant medical device: tprlc 133 t1 pps ho 17x154mm 4mm t1 item#51-104170 lot#6992243. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14641410&pc=LZO&device_sequence_no=1
3002806535-2022-00269 | 05/16/2022 | 06/08/2022 | 23 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF CMNTLS IMPLANT INSERT IMPACTOR | KNEE ARTHROPLASTY | - | It was reported that the knee instrument fractured upon impaction of the implant. The surgery was completed without any delay; there was no patient or user harm reported. Attempts have been made and no further information has been provided. | (b)(4). Foreign: canada. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14631874&pc=NRA&device_sequence_no=1
3002806535-2022-00268 | 04/27/2022 | 06/08/2022 | 42 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA HEAD, 36MM +3MM NECK, TYPE 1 TAPER | PROSTHESIS, HIP, SEMI-CONSTRAINED | - | It was reported that the patient was bending and twisting and felt a pop and increased pain on left hip approximately 1,5 month after initial left hip procedure. Anterior dislocation was confirmed; patient underwent a reduction under anesthesia. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical device: acetabular shell g7 3 hole, pps; catalog#010000664; lot#7167132, trilogy self-tapping bone screw; catalog#63759079; lot#63759079, acetabular g7 liner, longevity; catalog#20103606 ; lot#65255998, femoral stem echo bi-metric full proximal profile, lateralized; catalog#192111; lot#501010. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14631121&pc=LZO&device_sequence_no=1
3002806535-2022-00266 | 05/13/2022 | 06/08/2022 | 26 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF UNI FEMORAL DRILL 6.35 MM | OXFORD MICROPLASTY INSTRUMENTS | - | It was reported, that the drill broke off into drill guide during surgery. No known impact or consequence to the patient. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629808&pc=HTW&device_sequence_no=1
3002806535-2022-00251 | 05/06/2022 | 06/08/2022 | 33 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - RM SIZE B | OXFORD PARTIAL KNEE SYSTEM | - | It was reported that during an initial knee arthroplasty the bearing spun out and fell off to the posterior joint capsule side. The surgeon then decided to use a different system, and the femur and tibia components were removed. During removal, the patient¿s femur bone cracked. Surgery was delayed for about 40 minutes. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: oxf ph3 cementless fem sz xsm; item# 154912; lot# 7055221, oxf anat brg rt x-sm 4mm pma; item# 160791; lot# 7029590. Foreign: (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00249, 3002806535-2022-00252. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629805&pc=NRA&device_sequence_no=1
3002806535-2022-00252 | 05/06/2022 | 06/08/2022 | 33 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - XSMALL RIGHT SIZE 4 | OXFORD PARTIAL KNEE SYSTEM | - | It was reported that during an initial knee arthroplasty the bearing spun out and fell off to the posterior joint capsule side. The surgeon then decided to use a different system, and the femur and tibia components were removed. During removal, the patient¿s femur bone cracked. Surgery was delayed for about 40 minutes. Attempts have been made and no further information has been provided. | (b)(4) concomitant medical devices: oxf ph3 cementless fem sz xsm; item# 154912; lot# 7055221, oxf uni tib tray sz b rm pma; item# 154721; lot# 7083494. Reportv source foreign: (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00249, 3002806535-2022-00251. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629800&pc=NRA&device_sequence_no=1
3002806535-2022-00249 | 05/06/2022 | 06/08/2022 | 33 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF PH3 CEMENTLESS FEM SZ XSM | OXFORD PARTIAL KNEE SYSTEM | - | It was reported that during an initial knee arthroplasty the bearing spun out and fell off to the posterior joint capsule side. The surgeon then decided to use a different system, and the femur and tibia components were removed. During removal, the patient¿s femur bone cracked. Surgery was delayed for about 40 minutes. Attempts have been made and no further information has been provided. | (b)(4). Concomitant medical devices: oxf uni tib tray sz b rm pma; item# 154721; lot# 7083494, oxf anat brg rt x-sm 4mm pma; item# 160791; lot# 7029590. Foreign: (b)(6). Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00251, 3002806535-2022-00252. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14629797&pc=NRA&device_sequence_no=1
3002806535-2022-00257 | 05/12/2022 | 06/07/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - MEDIUM RIGHT SIZE 6 | OXFORD PARTIAL KNEE SYSTEM | - | It was reported, that the patient underwent a partial knee arthroplasty surgery and approximately 4. 5 years later the right mobile bearing was revised due to pain and instability. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Medical devices: oxf twin-peg cmntd fem md pma; item# 161469; lot# 372520; oxf uni tib tray sz c rm pma; item# 154723; lot# 568190. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14617584&pc=NRA&device_sequence_no=1
3002806535-2022-00267 | 05/09/2022 | 06/06/2022 | 28 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF TRL BRG W/SLOTS MED 3MM | BMET GENERAL KNEE INSTRUMENTS | - | It was reported that during initial right knee arthroplasty, when the surgeon took out the spacer from the knee, it broke. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). (b)(6). Investigation of this incident is currently ongoing. A followup/ final report will be submitted when additional information becomes available. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14601516&pc=NRA&device_sequence_no=1
3002806535-2022-00260 | 05/04/2022 | 06/01/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE B | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported, that: the patient underwent a revision due to the loosening and pain of the left knee, approximately one year post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant products: oxf anat brg lt sm size 3 pma item#159540 lot#075460; oxf twin-peg cmntd fem sm pma item#161468 lot#953870; biomet bone cement r 1x40us item#110035368 lot#005fae2408. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14561378&pc=NRA&device_sequence_no=1
3002806535-2022-00259 | 05/02/2022 | 06/01/2022 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE B | OXFORD CEMENTED TIBIAL TRAYS | - | It was reported, that the patient underwent a revision surgery due to unknown reason, approximately 3 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg lt sm size 4 pma item#159541 lot#017210, oxf twin-peg cmntd fem sm pma item#161468 lot#533960. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022- 00254 and 3002806535-2022-00258. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14560647&pc=NRA&device_sequence_no=1
3002806535-2022-00258 | 05/02/2022 | 06/01/2022 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS TWIN PEG FEMUR - SMALL | OXFORD CEMENTED FEMORALS | - | It was reported, that the patient underwent a revision surgery due to unknown reason, approximately 3 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf anat brg lt sm size 4 pma item#159541 lot#017210, oxf uni tib tray sz b lm pma item#154720 lot#775070. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022- 00254 and 3002806535-2022-00259. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14560646&pc=NRA&device_sequence_no=1
3002806535-2022-00254 | 05/02/2022 | 06/01/2022 | 30 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL LEFT SIZE 4 | OXFORD HXLPE BEARINGS | - | It was reported that the patient underwent a revision surgery due to unknown reason, approximately 3 years post-op. Due diligence is in progress for this complaint; to date no additional information or product has been received. | (b)(4). Concomitant medical products: oxf twin-peg cmntd fem sm pma item#161468 lot#533960, oxf uni tib tray sz b lm pma item#154720 lot#775070. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00258, 3002806535-2022-00259. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14560645&pc=NRA&device_sequence_no=1
3002806535-2022-00261 | 05/06/2022 | 06/01/2022 | 26 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 32/+3MM T1 | HIP PROTHESIS | - | It was reported that: the ceramic femoral head fractured 3 weeks postop, the complication required revision surgery. No additional information. | (b)(4). Report source ¿ foreign: (b)(6). Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14560016&pc=LZO&device_sequence_no=1
3002806535-2022-00263 | 05/05/2022 | 06/01/2022 | 27 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 28/ 0MM T1 | HIP PROTHESIS | - | It was reported that the patient was revised due to dislocation after a fall approximately 2 months post implantation. No additional information is available. | (b)(4). Concomitant medical products: arcos con sz a hi 50mm; catalog# 11-301310; lot#818290, vivacit-e dm bearing 28x42mm; catalog# 110031011; lot#65254557, trilogy bone scr 6. 5x40; catalog#00-6250-065-40; lot#65094425, g7 dual mobility liner 42mm e; catalog#110024463; lot#883140. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14559977&pc=LZO&device_sequence_no=1
3002806535-2022-00264 | - | 05/31/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD TIBIAL COMPONENT | KNEE ARTHROPLASTY | - | It was reported that the patient underwent a left knee arthroplasty. Subsequently, approximately fourteen years post-op, the patient started having some pain and difficulty walking. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford femoral component item# unknown; lot# unknown. Unknown oxford bearing item# unknown; lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. If multiple reports submitted for the event, use following statement: multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00262, 3002806535-2022-00265. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14546268&pc=HRY&device_sequence_no=1
3002806535-2022-00262 | - | 05/31/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD FEMORAL COMPONENT | KNEE ARTHROPLASTY | - | It was reported that the patient underwent a left knee arthroplasty. Subsequently, approximately fourteen years post-op, the patient started having some pain and difficulty walking. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford tibial component item# unknown; lot# unknown. Unknown oxford bearing item# unknown; lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. If multiple reports submitted for the event, use following statement: multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00264, 3002806535-2022-00265. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14546254&pc=HRY&device_sequence_no=1
3002806535-2022-00265 | - | 05/31/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | UNKNOWN OXFORD BEARING | KNEE ARTHROPLASTY | - | It was reported that the patient underwent a left knee arthroplasty. Subsequently, approximately fourteen years post-op, the patient started having some pain and difficulty walking. Due diligence is in progress for this complaint; to date, no additional information or product has been received. | (b)(4). Concomitant medical products: unknown oxford femoral component item# unknown; lot# unknown. Unknown oxford tibial component item# unknown; lot# unknown. The device will not be returned for analysis; however, an investigation of the reported event is in progress. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. If multiple reports submitted for the event, use following statement: multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00262, 3002806535-2022-00264. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14546249&pc=HRY&device_sequence_no=1
3002806535-2022-00247 | 05/02/2022 | 05/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS ANATOMIC ARCOM MENISCAL BRG - SMALL LEFT SIZE 3 | PROSTHESIS, KNEE | - | It was reported that the patient underwent a revision surgery approximately 8 months after the implantation due to pain and metal allergy. Additional information requested. | (b)(4). Concomitant medical products: unknown oxford femoral component; catalog# unknown; lot# unknown. Unknown oxford tibial component; catalog# unknown; lot# unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00245, 3002806535-2022-00246. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14539613&pc=NRA&device_sequence_no=1
3002806535-2022-00246 | 05/02/2022 | 05/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS CEMENTED TIBIA, ALPHA, STANDARD - LM SIZE A | PROSTHESIS, KNEE | - | It was reported that the patient underwent a revision surgery approximately 8 months after the implantation due to pain and metal allergy. Additional information requested. | (b)(4). Concomitant products: unknown oxford femoral component; catalog#unknown; lot#unknown; unknown oxford bearing; catalog#unknown; lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00245, 3002806535-2022-00247. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14539606&pc=NRA&device_sequence_no=1
3002806535-2022-00245 | 05/02/2022 | 05/30/2022 | 28 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXFORD PKS TWIN PEG FEMUR - SMALL | PROSTHESIS, KNEE | - | It was reported that the patient underwent a revision surgery approximately 8 months after the implantation due to pain and metal allergy. Additional information requested. | (b)(4). Concomitant products: unknown oxford tibial component; catalog#unknown; lot#unknown; unknown oxford bearing; catalog#unknown; lot#unknown. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated reports: 3002806535-2022-00246, 3002806535-2022-00247. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14539603&pc=NRA&device_sequence_no=1
3002806535-2022-00244 | - | 05/27/2022 | - | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOLOX DELTA OPTION HEAD, 28MM | PROSTHESIS, HIP, SEMI-CONSTRAINED | - | It was reported that the patient suffered pain and severe hematoma which complicated her recovery and led to a revision surgery. Attempts have been made and no further information has been provided. | (b)(4). Customer has indicated that the product will not be returned to zimmer biomet for investigation, as the device location is unknown. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No Answer Provided | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14521760&pc=LZO&device_sequence_no=1
3002806535-2022-00250 | 12/20/2017 | 05/27/2022 | 1619 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 28MM | HIP PROSTHESIS | - | It was reported that the patient became infected post left hip revision that required a secondary revision. Attempts have been made and no further information has been made available at this time. | (b)(4). Concomitant medical products: option taper sleeve, -6mm neck, type 1 taper; catalog# 50-1064; lot# 2897754. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated report: 3002806535-2022-00248. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14520698&pc=LZO&device_sequence_no=1
3002806535-2022-00248 | 12/20/2017 | 05/27/2022 | 1619 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OPTION TAPER SLEEVE -6MM NECK TYPE 1 TAPER | HIP PROSTHESIS | - | It was reported that the patient became infected post left hip revision that required a secondary revision. Attempts have been made and no further information has been made available at this time. | (b)(4). Concomitant medical products: biolox delta option head, 28mm; catalog# 650-1055; lot# 2886422. Investigation of this incident is currently ongoing. Once the investigation is completed, a supplemental medwatch 3500a will be submitted. Multiple mdr reports were filed for this event, please see associated report: 3002806535-2022-00250. | Device Returned to Manufacturer | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14520695&pc=LZO&device_sequence_no=1
3002806535-2022-00243 | 09/11/2018 | 05/16/2022 | 1343 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | BIOLOX DELTA CER OPT HD LG DIA | - | It was reported, that: a patient¿s legal counsel reported that the patient underwent left total hip arthroplasty on (b)(6) 2017. Legal counsel further reports patient underwent a revision procedure on (b)(6) 2018, due to dislocation and failed hip. This report is based on allegations set forth in plaintiff¿s complaint, and the allegations contained therein are unverified. | (b)(4). Combined initial/final report. This complaint was originally based on a 28mm ceramic head produced in usa so all previous reports were filed by the us zb team. Now the product has been identified as a 36mm, which is produced in (b)(6), it is now under (b)(6) control. The product has already been returned to zimmer biomet. Associated products: medical product: cer option type 1 tpr sleve +3, catalogue number: 650-1067, lot number: unknown. Complaint summary: investigation reopened as the product has now been returned. A visual inspection of the returned product shows this ceramic head to be a 36mm and not a 28mm as reported in the complaint, the ceramtec # is r36. The head shows signs of metal transfer on its surface indicating that articulation has taken place, most likely against the surface of the shell. There is also damage to the edge of the tapered adapter most likely caused whilst removing during the revision. Ceramtec coc review-: the component properties and the microstructure as obtained from the quality documents fulfil the requirements as specified at the time of production. There is no indication of any pre-existing material defect. A review of complaint history identified an additional 1 similar complaint about the reported item 650-1057 (lot number unconfirmed). Unable to perform complaint history for unknown shell as item/lot number is unknown. It has been confirmed that implant 650-1057 is not within the scope or subject of any field actions or recalls which could be attributed to the reported events. The likely condition of the devices when they left zimmer biomet is conforming to the specification. The reported event has not been confirmed. The root cause of the reported event cannot be determined with the information provided. No corrective actions, preventive actions, or field actions resulted after the investigation of this event. Multiple mdr reports were filed for this event previously, please see associated report numbers: 0001825034-2019-01678, 0001825034-2019-01678-1, 0001825034-2019-01678-2. If any additional information is discovered or received that may adjust any conclusions or data, a supplemental report will be rendered accordingly and the complaint reopened. Zimmer biomet will continue to monitor for trends. | Yes | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14405778&pc=JDI&device_sequence_no=1
3002806535-2022-00240 | 02/04/2017 | 05/12/2022 | 1923 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | CER BIOLOXD OPTION HD 36MM | HIP PROTHESIS | - | It was reported in medical notes received that patient experienced dislocation approximately 5 months post-implantation. It is unknown if medical intervention took place. | Mdr initial report. Product remains implanted. Concomitant medical products: item #010000938, lot #3123575, g7 hi-wall e1 liner 36mm h; item #010000669, lot #3350575, g7 pps ltd acet shell 64h; item #010000999, lot #3365486, g7 screw 6. 5mm x 30mm; item #010000999, lot #3369626, g7 screw 6. 5mm x 30mm; item #103207, lot #220990, taperloc por fmrl 13. 5x147; item #650-1068, lot # 3123575, cer option type 1 tpr sleve +6. The investigation is in progress. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14369242&pc=LZO&device_sequence_no=1
3002806535-2022-00241 | 04/28/2021 | 05/12/2022 | 379 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | BIOLOX DELTA CER FEM HD LG DIA | - | It was reported that: initial right total hip arthroplasty performed (b)(6) 2021. Subsequently, the patient developed a pulmonary embolism on (b)(6) 2021, requiring 1 day hospitalization. Pe resolved without further complication. Upon follow-up, the patient reported no problems and was very satisfied. Patient involved. | (b)(4). Initial report. Customer has indicated that the product will not be returned to zimmer biomet for investigation, as it remains implanted. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14368992&pc=LZO&device_sequence_no=1
3002806535-2022-00239 | 10/03/2021 | 05/11/2022 | 220 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG RT SM SIZE 3 PMA | OXFORD HXLPE BEARINGS | - | It was reported, that: the oxford bearing has been revised due to an unknown reason. Poly swap, patient involvement-revision. | (b)(4). Initial report. The client has indicated that the product is in process of being returned to zimmer biomet for an investigation. Medical product: oxf uni tib tray sza rm, catalogue number: 154719, lot number: 751980. Medical product: oxf twin-peg cmntd fem sm pma, catalogue number: 161468, lot number: 835430. The investigation is currently underway. Once the investigation has concluded, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14351097&pc=NRA&device_sequence_no=1
3002806535-2022-00238 | 05/02/2022 | 05/09/2022 | 7 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CERAMIC FEM HD 36/0MM | BIOLOX DELTA CER FEM HD LG DIA | - | It was reported, that: the patient underwent a primary tha on the right side on (b)(6) 2022. The patient had hip pain and was admitted to (b)(6) hospital. It was seen on the x-ray of the right hip, that the ceramic head was in multiple pieces and revision surgery. | (b)(4). Initial report. The client has indicated that the product will not be returned to zimmer biomet for an investigation. Associated products: medical product: bone screw self-tapping 6. 5 mm dia. 30 mm length, catalogue number: 00625006530, lot number: j7137592; medical product: tprlc 133 mp type1 pps ho 17. 0, catalogue number: 51-107170, lot number: 6674090; medical product: g7 neutral arcomxl lnr 36mm f, catalogue number: 010000741, lot number: 6897739; medical product: g7 pps ltd acet shell 54f, catalogue number: 010000664, lot number: 6897739; medical product: pr prm st/g7 cp/xl ln/cer, catalogue number: 98-b001-007-23, lot number: unknown. The investigation is currently underway. Once the investigation has concluded, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14333997&pc=LZO&device_sequence_no=1
3002806535-2022-00237 | 03/02/2022 | 05/07/2022 | 66 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | DELTA CER FEM HD 28/-3MM T1 | BIOLOX DELTA CERAMIC FEM HD | - | It was reported, that: on 02-mar-2022, a journal article was retrieved from the journal of orthopaedic surgery (2021) that reported a study from japan that looked at the impact of the surgical approach on postoperative dislocations. The purpose of the study was to evaluate the efficacy and safety of hemiarthroplasty performed using the cpp (conjoined tendon preserving posterior) approach for femoral neck fractures. The study reviewed 322 patients with femoral neck fractures. A bipolar hemiarthroplasty was performed in all patients using a press-fit taperloc stem and ringloc bipolar cup. Two patients were excluded from the study due to receiving a total hip at the time of surgery, leaving 320 patients for final review. No patients were excluded based on age, frailty, or cognitive status. The study population had a mean age of 83. 3 years at the time of surgery (range 57-104) and a mean bmi of 20. 7 (range 11. 4-28. 9). The final follow-up was conducted at 9. 1 +/- 1. 5 months after surgery (range 6-17 months). Complaint 7: the study reported an initial right hip hemiarthroplasty performed on (b)(6) 2018 on a (b)(6) female. After closing the wound during the initial procedure, the patient developed low cardiac output and oxygen levels. The patient was subsequently diagnosed with a pulmonary embolism. The pe was noted to have resolved. No further details have been provided. | (b)(4). Initial report. Report source, foreign - event occurred in (b)(6). The client has indicated that the product will not be returned to zimmer biomet for an investigation. Associated products: medical product: tprlc 133 fp type1 pps so 8. 0. Catalogue number: 51-100080. Lot number: unknown. Medical product: e1 ringloc bipolar 28x46mm. Catalogue number: 110010465. Lot number: unknown. The investigation is currently underway. Once the investigation has concluded, a follow-up mdr will be submitted. Remains implanted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14315851&pc=LZO&device_sequence_no=1
3002806535-2022-00236 | 04/14/2022 | 04/29/2022 | 15 | Injury | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF ANAT BRG LT MD SIZE 3 PMA | OXFORD HXLPE BEARINGS | - | It was reported that: a revision took place to remove the implant because the reference indicated on the package label is the wrong reference; the reference was for size 3m, whereas on the implant it is indicated size 3l (the lot number on the label and engraved on the implant is the same). | (b)(4). Initial report. Report source, foreign - event occurred in (b)(6). Customer has indicated that the product will not be returned to zimmer biomet for investigation. The investigation is in process. Once the investigation has been completed, a follow-up mdr will be submitted. | No | Yes | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14248686&pc=NRA&device_sequence_no=1
3002806535-2022-00235 | 04/25/2022 | 04/28/2022 | 3 | malfunction | Manufacturer | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | BIOMET UK LTD. waterton industrial estates bridgend CF31 3XA UK CF31 3XA | OXF KNEE PH3 I/M ROD RMVL HK L HK | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | - | It was reported, that: the instrument broke when being used. Another instrument was used from another set. There was a 10-minute delay. Patient involvement. | (b)(4). Initial report. Report source, foreign - event occurred in (b)(6). The client has indicated that the product is in process of being returned to zimmer biomet for an investigation. The investigation is currently underway. Once the investigation has concluded, a follow-up mdr will be submitted. | No | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=14236273&pc=LXH&device_sequence_no=1

## Inspections
City | State | Zip Code | Inspection End Date | Project Area | Classification
Bridgend | - | - | 2015-03-18 | Compliance: Devices | Voluntary Action Indicated (VAI)
Bridgend | - | - | 2009-07-02 | Compliance: Devices | Voluntary Action Indicated (VAI)
Bridgend | - | - | 2009-07-02 | Postmarket Assurance: Devices | No Action Indicated (NAI)

## Devices
Device Summary | Device Detail | TPLC Information | # Reported Events
Device Name | ENDOBON XENOGRAFT GRANULES | TPLC Device Problem | MDR
Device Class | 2 | Failure to Osseointegrate33 | 16
Product Code | LYC | Adverse Event Without Identified Device or Use Problem34 | 14
Regulation No | 872.39 | Patient-Device Incompatibility35 | 8
Premarket Submission No | K110449 | Insufficient Information36 | 7
Establishment Operations | Manufacturer | Product Quality Problem37 | 4
Recalls | 1
Device Name | TEMPLATE | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions124 | 5563
Product Code | HWT | No Consequences Or Impact To Patient125 | 1771
Regulation No | 888.48 | Insufficient Information126 | 235
Premarket Submission No | - | No Code Available127 | 139
Establishment Operations | Manufacturer | No Patient Involvement128 | 115
Recalls | -
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HSB | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTY | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | JDW | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | KTT | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | TRAUMA INTERNAL FIXATION SYSTEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | LXT | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K111663 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 17
Device Name | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | Break29 | 10678
Product Code | LXH | Device-Device Incompatibility30 | 3479
Regulation No | 888.45 | Fracture31 | 2599
Premarket Submission No | - | Material Deformation32 | 1333
Establishment Operations | Manufacturer | Material Twisted/Bent33 | 1264
Recalls | -
Device Name | STAPLE, FIXATION BONE; SCREW FIXATION BONE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K931155 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | BIT, DRILL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 3549
Product Code | HTW | No Consequences Or Impact To Patient129 | 2172
Regulation No | 888.45 | Foreign Body In Patient130 | 881
Premarket Submission No | - | No Code Available131 | 342
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient132 | 308
Recalls | -
Device Name | REAMER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions123 | 10673
Product Code | HTO | No Consequences Or Impact To Patient124 | 4137
Regulation No | 888.45 | No Code Available125 | 459
Premarket Submission No | - | Insufficient Information126 | 396
Establishment Operations | Manufacturer | Foreign Body In Patient127 | 321
Recalls | -
Device Name | TAP, BONE | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions65 | 214
Product Code | HWX | No Consequences Or Impact To Patient66 | 127
Regulation No | 888.45 | No Patient Involvement67 | 78
Premarket Submission No | - | Foreign Body In Patient68 | 26
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient69 | 22
Recalls | -
Device Name | STARTER, BONE SCREW | TPLC Device Problem | MDR
Device Class | 1 | Unspecified Infection30 | 1
Product Code | HWD | No Patient Involvement31 | 1
Regulation No | 888.45 | Foreign Body In Patient32 | 1
Premarket Submission No | - | No Known Impact Or Consequence To Patient33 | 1
Establishment Operations | Manufacturer | No Information34 | 1
Recalls | -
Device Name | GUIDE, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 1714
Product Code | FZX | No Consequences Or Impact To Patient129 | 784
Regulation No | 878.48 | No Patient Involvement130 | 390
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 198
Establishment Operations | Manufacturer | No Code Available132 | 148
Recalls | -
Device Name | SCREWDRIVER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 4583
Product Code | HXX | No Consequences Or Impact To Patient129 | 2515
Regulation No | 888.45 | No Patient Involvement130 | 881
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 867
Establishment Operations | Manufacturer | No Code Available132 | 314
Recalls | -
Device Name | FORCEPS, WIRE HOLDING | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions27 | 2
Product Code | HYA | No Patient Involvement28 | 1
Regulation No | 878.48 | Foreign Body In Patient29 | 1
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | INSTRUMENT, BENDING OR CONTOURING | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions37 | 38
Product Code | HXP | No Consequences Or Impact To Patient38 | 8
Regulation No | 888.45 | No Patient Involvement39 | 4
Premarket Submission No | - | No Known Impact Or Consequence To Patient40 | 1
Establishment Operations | Manufacturer | Not Applicable41 | 1
Recalls | -
Device Name | PLIERS, SURGICAL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions48 | 49
Product Code | HTC | No Consequences Or Impact To Patient49 | 33
Regulation No | 878.48 | No Patient Involvement50 | 23
Premarket Submission No | - | No Known Impact Or Consequence To Patient51 | 5
Establishment Operations | Manufacturer | Foreign Body In Patient52 | 2
Recalls | -
Device Name | RETRACTOR | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions95 | 190
Product Code | GAD | No Consequences Or Impact To Patient96 | 76
Regulation No | 878.48 | No Known Impact Or Consequence To Patient97 | 41
Premarket Submission No | - | No Patient Involvement98 | 33
Establishment Operations | Manufacturer | Foreign Body In Patient99 | 14
Recalls | -
Device Name | GAUGE, DEPTH | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions75 | 933
Product Code | HTJ | No Consequences Or Impact To Patient76 | 308
Regulation No | 888.43 | No Patient Involvement77 | 248
Premarket Submission No | - | No Known Impact Or Consequence To Patient78 | 24
Establishment Operations | Manufacturer | Insufficient Information79 | 18
Recalls | -
Device Name | TRAY, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions119 | 396
Product Code | FSM | No Consequences Or Impact To Patient120 | 297
Regulation No | 878.48 | No Known Impact Or Consequence To Patient121 | 38
Premarket Submission No | - | No Patient Involvement122 | 24
Establishment Operations | Manufacturer | No Code Available123 | 22
Recalls | -
Device Name | ACE HUMERAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K934643 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | AIM SUPRACONDYLAR NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K934962 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K882831 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY AE TK2 HIP SCREW SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem56 | 1765
Product Code | KTT | Break57 | 764
Regulation No | 888.3 | Device-Device Incompatibility58 | 380
Premarket Submission No | K972629 | Migration59 | 354
Establishment Operations | Manufacturer | Fracture60 | 129
Recalls | 2
Device Name | ACE RECONSTRUCTIVE PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K905595 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ACE CORTICAL BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K912598 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE A.I.M. TITANIUM TIBIAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K972183 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 2
Device Name | ACE AIM TTC FUSION NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K003797 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE BROOKER BONE PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K833865 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ZINDRICK IM ROD | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K871539 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 1
Device Name | REVERSE CUTTING CANNULATED CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K892214 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE CANNULATED SELF TAPPING CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K903810 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | 6.5MM SOLID CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K954485 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ZINDRICK IM ROD INSTRUMENTS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K882443 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | SMALL BONE LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K081546 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 1
Device Name | ACE 3.5MM CORTICAL BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K905048 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE HUMERUS & RADIUS PLATES | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K955472 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | THACKRAY TITANIUM MINI-PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K905774 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | DEPUY ACE TIMAX META PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K983853 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ANTEROLATERAL AND MEDIAL LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072832 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | POLYAX LOCKED PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K060969 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 10
Device Name | PE.R.I. II KNEE FRACTURE SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K003235 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | 4.5MM LOCKING BROAD & NARROW COMPRESSION PLATES, 4.5MM & 5.5MM CORTICAL & CANCELLOUS LOCKING SCREWS, RESPECTIVELY. | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072423 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | LOCKING ANATOMIC & COMPOSITE PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K083843 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | CLOSED FUSION PLATE, MTP PLATE, TALUS PLATE, MODEL 824077021, 22; 824677021, 22, 824071001, 824071101, 824671001, 82467 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101240 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | SCARF THREAD-HEAD SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K962236 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE UNIVERSAL AND TROCH ENTRY FEMORAL NAIL SYSTEMS | TPLC Device Problem | MDR
Device Class | 2+B452 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K033329 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | HFN 125 DEGREE 9-13 MM, 180MM, HFN 130 DEGREE 9-13 MM, HFN RH 125 DEGREE 9-13MM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K100238 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 100
Device Name | UNIVERSAL AND TROCH ENTRY FEMORAL NAILING SYSTEMS LINE EXTENSION | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K111444 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE BIO WIN TIBIAL NAILING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K013430 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | DISPENSER, CEMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions63 | 53
Product Code | KIH | No Consequences Or Impact To Patient64 | 41
Regulation No | 888.42 | Foreign Body In Patient65 | 9
Premarket Submission No | - | No Known Impact Or Consequence To Patient66 | 7
Establishment Operations | Manufacturer | Insufficient Information67 | 7
Recalls | -
Device Name | MIXER, CEMENT, FOR CLINICAL USE | TPLC Device Problem | MDR
Device Class | 1 | No Patient Involvement70 | 151
Product Code | JDZ | No Clinical Signs, Symptoms or Conditions71 | 135
Regulation No | 888.42 | No Consequences Or Impact To Patient72 | 35
Premarket Submission No | - | No Known Impact Or Consequence To Patient73 | 30
Establishment Operations | Manufacturer | No Information74 | 6
Recalls | -
Device Name | TROCHANTERIC NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K010780 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 84
Device Name | ACE VERSANAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K023115 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | DEPUY ACE VERSANAIL TIBIAL NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K032097 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | ACE NANCY NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem47 | 354
Product Code | HTY | Break48 | 196
Regulation No | 888.3 | Fracture49 | 71
Premarket Submission No | K032687 | Mechanical Jam50 | 44
Establishment Operations | Manufacturer | Material Twisted/Bent51 | 19
Recalls | -
Device Name | ACE ANTEGRADE RETROGRADE HUMERAL NAIL SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K033878 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | TROCHANTERIC NAIL SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K042325 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | 11
Device Name | MEMORY STAPLE | TPLC Device Problem | MDR
Device Class | 2 | Break52 | 394
Product Code | JDR | Adverse Event Without Identified Device or Use Problem53 | 229
Regulation No | 888.3 | Detachment of Device or Device Component54 | 122
Premarket Submission No | K060746 | Fracture55 | 74
Establishment Operations | Manufacturer | Material Frayed56 | 37
Recalls | -
Device Name | FRS SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K062352 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | SMALL FRAGMENT LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K072083 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | ANATOMIC LOCKING PLATING SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K082300 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | MULTIPLICITY OF PRODUCTS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K782123 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | ACE CANCELLOUS BONE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K872859 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | ACE OPEN-BACK HALO RING | TPLC Device Problem | MDR
Device Class | 2 | Device-Device Incompatibility33 | 25
Product Code | JEC | Detachment of Device or Device Component34 | 3
Regulation No | 888.3 | Break35 | 3
Premarket Submission No | K914858 | Mechanical Jam36 | 2
Establishment Operations | Manufacturer | Adverse Event Without Identified Device or Use Problem37 | 2
Recalls | -
Device Name | NANCY NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem47 | 354
Product Code | HTY | Break48 | 196
Regulation No | 888.3 | Fracture49 | 71
Premarket Submission No | K960642 | Mechanical Jam50 | 44
Establishment Operations | Manufacturer | Material Twisted/Bent51 | 19
Recalls | -
Device Name | TWIST-OFF SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K962233 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | MEMORY STAPLE (20 SIZES AVAILABLE) | TPLC Device Problem | MDR
Device Class | 2 | Break52 | 394
Product Code | JDR | Adverse Event Without Identified Device or Use Problem53 | 229
Regulation No | 888.3 | Detachment of Device or Device Component54 | 122
Premarket Submission No | K964226 | Fracture55 | 74
Establishment Operations | Manufacturer | Material Frayed56 | 37
Recalls | -
Device Name | TWIST-OFF (TM) SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K971069 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | DEPUY ACE AIM(R) TITANIUM SUPRACONDYLAR NAIL | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K974781 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | MODIFICATION OF DEPUY ACE TIMAX MEDIAL PILON PLATE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem73 | 5184
Product Code | HSB | Break74 | 3375
Regulation No | 888.3 | Device-Device Incompatibility75 | 1735
Premarket Submission No | K990120 | Migration76 | 1581
Establishment Operations | Manufacturer | Fracture77 | 1197
Recalls | -
Device Name | Refobacin Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | LOD | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K171540 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | Refobacin Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | MBB | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K171540 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | Biomet Bone Cement R | TPLC Device Problem | MDR
Device Class | 2 | Loss of or Failure to Bond35 | 3430
Product Code | LOD | Adverse Event Without Identified Device or Use Problem36 | 1140
Regulation No | 888.3 | Loosening of Implant Not Related to Bone-Ingrowth37 | 369
Premarket Submission No | K172408 | Loose or Intermittent Connection38 | 259
Establishment Operations | Manufacturer | Insufficient Information39 | 207
Recalls | -
Device Name | MODIFICATION TO: LIGAFIX INTERFERENCE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K070507 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | 1
Device Name | LIGAFIX INTERFACE SCREW | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem120 | 10188
Product Code | HWC | Break121 | 5202
Regulation No | 888.3 | Migration122 | 1640
Premarket Submission No | K090994 | Device-Device Incompatibility123 | 1508
Establishment Operations | Manufacturer | Fracture124 | 875
Recalls | -
Device Name | A. L. P. S. Clavicle Plating System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K173767 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 2
Device Name | Biomet Proximal Humerus Plating System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K143697 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | 18
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HTN | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | MEDICAL PLATES, LATERAL PLATES, POSTERIOR LATERAL PLATES, OLECRANON PLATES MODEL 131218704-708; 851318704-708, 131218204 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K101421 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | POSTERIOR LATERAL PLATES | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K103408 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | PROXIMAL TIBIA LOCKING PLATE STANDARD AND LARGE, 3.5MM LOW PROFILE AND MULTIDIRECTIONAL SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090877 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | PROXIMAL TIBIA LOCKING PLATE STANDARD AND LARGE, 3.5MM LOW PROFILE AND MULTIDIRECTIONAL SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HWC | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090877 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | LAPIDUS PLATE, 3.5MM INLINE FUSION PLATE, 3.5MM MTP PLATE (LEFT AND RIGHT) | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K131670 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | 1.5MM SMALL LOCKING PLATE, 1.5MM NON-LOCKING AND LOCKING SCREWS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem137 | 6302
Product Code | HRS | Break138 | 2379
Regulation No | 888.3 | Fracture139 | 1051
Premarket Submission No | K090492 | Device-Device Incompatibility140 | 598
Establishment Operations | Manufacturer | Migration141 | 514
Recalls | -
Device Name | PUSHER, SOCKET | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions27 | 3
Product Code | HXO | No Information28 | 1
Regulation No | 888.45
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | CHISEL, SURGICAL, MANUAL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions51 | 51
Product Code | FZO | No Patient Involvement52 | 49
Regulation No | 878.48 | No Consequences Or Impact To Patient53 | 37
Premarket Submission No | - | No Known Impact Or Consequence To Patient54 | 10
Establishment Operations | Manufacturer | Bone Fracture(s)55 | 4
Recalls | -
Device Name | FORCEPS, WIRE HOLDING | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions27 | 2
Product Code | HYA | No Patient Involvement28 | 1
Regulation No | 878.48 | Foreign Body In Patient29 | 1
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | JDI | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | KWZ | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | LPH | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | LZO | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | OQG | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | OQH | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | OQI | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | G7 ACETABULAR SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Difficult to Insert29 | 299
Product Code | PBI | Adverse Event Without Identified Device or Use Problem30 | 133
Regulation No | 888.33 | Fracture31 | 120
Premarket Submission No | K121874 | Insufficient Information32 | 95
Establishment Operations | Manufacturer | Positioning Failure33 | 87
Recalls | 3
Device Name | SCREWDRIVER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 4583
Product Code | HXX | No Consequences Or Impact To Patient129 | 2515
Regulation No | 888.45 | No Patient Involvement130 | 881
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 867
Establishment Operations | Manufacturer | No Code Available132 | 314
Recalls | -
Device Name | CANNULA, SURGICAL, GENERAL & PLASTIC SURGERY | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions92 | 249
Product Code | GEA | No Consequences Or Impact To Patient93 | 181
Regulation No | 878.48 | Not Applicable94 | 27
Premarket Submission No | - | No Known Impact Or Consequence To Patient95 | 20
Establishment Operations | Manufacturer | Abscess96 | 19
Recalls | -
Device Name | DEVICE, PROSTHESIS ALIGNMENT | TPLC Device Problem | MDR
Device Class | 1 | No Consequences Or Impact To Patient24 | 1
Product Code | IQO | No Code Available25 | 1
Regulation No | 890.3
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | RETRACTOR | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions95 | 190
Product Code | GAD | No Consequences Or Impact To Patient96 | 76
Regulation No | 878.48 | No Known Impact Or Consequence To Patient97 | 41
Premarket Submission No | - | No Patient Involvement98 | 33
Establishment Operations | Manufacturer; Specification Developer | Foreign Body In Patient99 | 14
Recalls | -
Device Name | TRAY, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions119 | 396
Product Code | FSM | No Consequences Or Impact To Patient120 | 297
Regulation No | 878.48 | No Known Impact Or Consequence To Patient121 | 38
Premarket Submission No | - | No Patient Involvement122 | 24
Establishment Operations | Manufacturer; Specification Developer | No Code Available123 | 22
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDI | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWA | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWZ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LPH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LZO | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MEH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE MICROPLASTY STEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | OQG | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K110400 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | AGC TOTAL KNEE PROSTHESIS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem83 | 23191
Product Code | JWH | Insufficient Information84 | 7973
Regulation No | 888.35 | Fracture85 | 5898
Premarket Submission No | K833921 | Loss of or Failure to Bond86 | 4499
Establishment Operations | Manufacturer | Appropriate Term/Code Not Available87 | 3220
Recalls | 2
Device Name | PUNCH, SURGICAL | TPLC Device Problem | MDR
Device Class | 1 | No Consequences Or Impact To Patient74 | 173
Product Code | LRY | No Clinical Signs, Symptoms or Conditions75 | 152
Regulation No | 878.48 | Not Applicable76 | 30
Premarket Submission No | - | No Code Available77 | 25
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient78 | 13
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | JDI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWZ | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LPH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LZO | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQG | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 OSSEOTI ACETUBLAR SHELLS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | PBI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K140669 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | JDI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWL | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWY | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWZ | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LPH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LZO | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | MEH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TAPERLOC COMPLETE SIZE 4MM AND XR 123 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQG | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K120030 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | MALLET, SURGICAL, GENERAL & PLASTIC SURGERY | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions28 | 3
Product Code | GFJ | No Consequences Or Impact To Patient29 | 2
Regulation No | 878.48 | No Code Available30 | 1
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | JDI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWZ | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LPH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LZO | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQG | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Finned Acetabular Shell | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | PBI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142746 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | TEMPLATE | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions124 | 5563
Product Code | HWT | No Consequences Or Impact To Patient125 | 1771
Regulation No | 888.48 | Insufficient Information126 | 235
Premarket Submission No | - | No Code Available127 | 139
Establishment Operations | Manufacturer; Specification Developer | No Patient Involvement128 | 115
Recalls | -
Device Name | DRIVER, PROSTHESIS | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions93 | 2572
Product Code | HWR | No Consequences Or Impact To Patient94 | 1629
Regulation No | 888.45 | Not Applicable95 | 116
Premarket Submission No | - | No Code Available96 | 101
Establishment Operations | Manufacturer; Specification Developer | Insufficient Information97 | 86
Recalls | -
Device Name | GAUGE, DEPTH | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions75 | 933
Product Code | HTJ | No Consequences Or Impact To Patient76 | 308
Regulation No | 888.43 | No Patient Involvement77 | 248
Premarket Submission No | - | No Known Impact Or Consequence To Patient78 | 24
Establishment Operations | Manufacturer | Insufficient Information79 | 18
Recalls | -
Device Name | EXTRACTOR | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions88 | 803
Product Code | HWB | No Consequences Or Impact To Patient89 | 660
Regulation No | 888.45 | No Code Available90 | 101
Premarket Submission No | - | No Known Impact Or Consequence To Patient91 | 43
Establishment Operations | Manufacturer | Not Applicable92 | 42
Recalls | -
Device Name | IMPACTOR | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions114 | 5252
Product Code | HWA | No Consequences Or Impact To Patient115 | 3980
Regulation No | 888.45 | Insufficient Information116 | 242
Premarket Submission No | - | No Code Available117 | 199
Establishment Operations | Manufacturer | No Known Impact Or Consequence To Patient118 | 187
Recalls | -
Device Name | INSTRUMENT, SURGICAL, ORTHOPEDIC, AC-POWERED MOTOR AND ACCESSORY/ATTACHMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 8428
Product Code | HWE | No Consequences Or Impact To Patient129 | 6844
Regulation No | 878.48 | No Known Impact Or Consequence To Patient130 | 1869
Premarket Submission No | - | No Patient Involvement131 | 759
Establishment Operations | Manufacturer | Foreign Body In Patient132 | 346
Recalls | -
Device Name | GUIDE, SURGICAL, INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 1714
Product Code | FZX | No Consequences Or Impact To Patient129 | 784
Regulation No | 878.48 | No Patient Involvement130 | 390
Premarket Submission No | - | No Known Impact Or Consequence To Patient131 | 198
Establishment Operations | Manufacturer | No Code Available132 | 148
Recalls | -
Device Name | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | TPLC Device Problem | MDR
Device Class | 1 | Break29 | 10678
Product Code | LXH | Device-Device Incompatibility30 | 3479
Regulation No | 888.45 | Fracture31 | 2599
Premarket Submission No | - | Material Deformation32 | 1333
Establishment Operations | Manufacturer; Specification Developer | Material Twisted/Bent33 | 1264
Recalls | -
Device Name | BROACH | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions100 | 1215
Product Code | HTQ | No Consequences Or Impact To Patient101 | 1129
Regulation No | 888.45 | No Code Available102 | 169
Premarket Submission No | - | Insufficient Information103 | 118
Establishment Operations | Manufacturer | Not Applicable104 | 76
Recalls | -
Device Name | REAMER | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions123 | 10673
Product Code | HTO | No Consequences Or Impact To Patient124 | 4137
Regulation No | 888.45 | No Code Available125 | 459
Premarket Submission No | - | Insufficient Information126 | 396
Establishment Operations | Manufacturer | Foreign Body In Patient127 | 321
Recalls | -
Device Name | VANGUARD 360 REVISION KNEE SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem83 | 23191
Product Code | JWH | Insufficient Information84 | 7973
Regulation No | 888.35 | Fracture85 | 5898
Premarket Submission No | K093293 | Loss of or Failure to Bond86 | 4499
Establishment Operations | Manufacturer | Appropriate Term/Code Not Available87 | 3220
Recalls | 5
Device Name | VANGUARD 360 REVISION KNEE SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem83 | 23191
Product Code | MBV | Insufficient Information84 | 7973
Regulation No | 888.35 | Fracture85 | 5898
Premarket Submission No | K093293 | Loss of or Failure to Bond86 | 4499
Establishment Operations | Manufacturer | Appropriate Term/Code Not Available87 | 3220
Recalls | 5
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDI | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWA | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWZ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LPH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LWJ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LZO | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MAY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K1010869 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MEH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K101086 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | JDI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWZ | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LPH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LZO | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQG | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | G7 Freedom and Offset Liners, Freedom Head, Size 32 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | PBI | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K142882 | Fracture74 | 1609
Establishment Operations | Manufacturer | Naturally Worn75 | 1604
Recalls | -
Device Name | BIT, DRILL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions128 | 3549
Product Code | HTW | No Consequences Or Impact To Patient129 | 2172
Regulation No | 888.45 | Foreign Body In Patient130 | 881
Premarket Submission No | - | No Code Available131 | 342
Establishment Operations | Manufacturer; Specification Developer | No Known Impact Or Consequence To Patient132 | 308
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDI | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWA | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWZ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LPH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LWJ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LZO | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MAY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | TAPERLOC COMPLETE, SIZE 5 AND 6 | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MEH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K103755 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Manufacturer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDI | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | JDL | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWA | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | KWZ | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LPH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | LZO | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MAY | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | ARCOS INTERLOCKING DISTAL STEMS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem28 | 8637
Product Code | MEH | Insufficient Information29 | 3294
Regulation No | 888.33 | Naturally Worn30 | 2385
Premarket Submission No | K100469 | Device Dislodged or Dislocated31 | 2235
Establishment Operations | Specification Developer | Loss of Osseointegration32 | 1196
Recalls | -
Device Name | G7 Dual Mobility System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | KWY | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K150522 | Fracture74 | 1609
Establishment Operations | Specification Developer | Naturally Worn75 | 1604
Recalls | 1
Device Name | G7 Dual Mobility System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LPH | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K150522 | Fracture74 | 1609
Establishment Operations | Specification Developer | Naturally Worn75 | 1604
Recalls | 1
Device Name | G7 Dual Mobility System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | LZO | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K150522 | Fracture74 | 1609
Establishment Operations | Specification Developer | Naturally Worn75 | 1604
Recalls | 1
Device Name | G7 Dual Mobility System | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | OQG | Device Dislodged or Dislocated72 | 5541
Regulation No | 888.33 | Insufficient Information73 | 3795
Premarket Submission No | K150522 | Fracture74 | 1609
Establishment Operations | Specification Developer | Naturally Worn75 | 1604
Recalls | 1
Device Name | OXFORD(TM) MENISCAL UNICOMPARTMENTAL KNEE SYSTEM | TPLC Device Problem | MDR
Device Class | 3 | Adverse Event Without Identified Device or Use Problem71 | 11415
Product Code | NRA | Device Dislodged or Dislocated72 | 5541
Regulation No |  | Insufficient Information73 | 3795
Premarket Submission No | P0100148 | Fracture74 | 1609
Establishment Operations | Manufacturer; Specification Developer | Naturally Worn75 | 1604
Recalls | 22
Device Name | BIOLOX DELTA OPTION CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K082996 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | 44MM E1 ACETABULAR LINER WITH 44MM BIOLOX DELTA OPTION CERAMIC HEAD OR 44MM M A MAGNUM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | JDI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K093549 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | 44MM E1 ACETABULAR LINER WITH 44MM BIOLOX DELTA OPTION CERAMIC HEAD OR 44MM M A MAGNUM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LPH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K093549 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | 44MM E1 ACETABULAR LINER WITH 44MM BIOLOX DELTA OPTION CERAMIC HEAD OR 44MM M A MAGNUM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LWJ | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K093549 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | 44MM E1 ACETABULAR LINER WITH 44MM BIOLOX DELTA OPTION CERAMIC HEAD OR 44MM M A MAGNUM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K093549 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | 44MM E1 ACETABULAR LINER WITH 44MM BIOLOX DELTA OPTION CERAMIC HEAD OR 44MM M A MAGNUM | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | MAY | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K093549 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | ARCOMXL POLYETHYLENE LINERS AND BIOLOX DELTA HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K051411 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | 1
Device Name | 36MM BIOLOX DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K061312 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | SURGICAL TABLE CUSHION | TPLC Device Problem | MDR
Device Class | 1 | Pressure Sores28 | 1
Product Code | LWG | Insufficient Information29 | 1
Regulation No | 878.49 | No Clinical Signs, Symptoms or Conditions30 | 1
Premarket Submission No | -
Establishment Operations | Manufacturer
Recalls | -
Device Name | OSTEOTOME | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions67 | 182
Product Code | HWM | No Consequences Or Impact To Patient68 | 135
Regulation No | 878.48 | No Patient Involvement69 | 11
Premarket Submission No | - | Insufficient Information70 | 10
Establishment Operations | Specification Developer | Not Applicable71 | 5
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | JDI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LPH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LWJ | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | MAY | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQG | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX DELTA CERAMIC OPTION HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K141653 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | JDI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LPH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LWJ | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQG | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | BIOLOX(R) DELTA CERAMIC HEADS | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K131684 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | 2
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | JDI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LPH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LWJ | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | MAY | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQG | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox delta Ceramic Heads, Biolox delta Option Ceramic Heads | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K192683 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Manufacturer; Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | JDI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LPH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LWJ | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | LZO | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | MAY | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQG | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQH | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | Biolox® delta Ceramic Heads, Biolox® delta Option Ceramic Heads MR Labeling | TPLC Device Problem | MDR
Device Class | 2 | Adverse Event Without Identified Device or Use Problem68 | 3573
Product Code | OQI | Device Dislodged or Dislocated69 | 2641
Regulation No | 888.33 | Insufficient Information70 | 2271
Premarket Submission No | K200959 | Appropriate Term/Code Not Available71 | 1209
Establishment Operations | Specification Developer | Fracture72 | 862
Recalls | -
Device Name | INSTRUMENT, CUTTING, ORTHOPEDIC | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions74 | 532
Product Code | HTZ | No Consequences Or Impact To Patient75 | 163
Regulation No | 878.48 | Insufficient Information76 | 25
Premarket Submission No | - | No Patient Involvement77 | 23
Establishment Operations | Specification Developer | No Known Impact Or Consequence To Patient78 | 17
Recalls | -
Device Name | AWL | TPLC Device Problem | MDR
Device Class | 1 | No Clinical Signs, Symptoms or Conditions41 | 17
Product Code | HWJ | No Code Available42 | 13
Regulation No | 888.45 | No Consequences Or Impact To Patient43 | 11
Premarket Submission No | - | Foreign Body In Patient44 | 6
Establishment Operations | Specification Developer | No Known Impact Or Consequence To Patient45 | 3
Recalls | -
Device Name | OXFORD UNICOMPARTMENTAL KNEE FEMORAL COMPONENT | TPLC Device Problem | MDR
Device Class | 2 | Break30 | 1431
Product Code | HRY | Fracture31 | 411
Regulation No | 888.35 | Adverse Event Without Identified Device or Use Problem32 | 392
Premarket Submission No | K011138 | Insufficient Information33 | 385
Establishment Operations | Specification Developer | Loosening of Implant Not Related to Bone-Ingrowth34 | 200
Recalls | 1

## Info
Establishment Name | Error | Timestamp
BIOMET UK LTD | Warning Letters Not Found | 2023-08-07 16:34