## Entity Summary
Field | Extracted Data | Remarks
Establishment Name | ENCISION, INC.
Registration No | 1722040
Registration Year | 2023
Total No Of Devices | 8
Class 3 Devices | -
No of Inspections | 15
No Of Recalls | 26
Warning Letters | -
1.Response Letter | N/A
2.Closeout Letter | N/A
Adverse Events | 2
Is This an adverse Report? | -
Establishment | ENCISION, INC. 6797 WINCHESTER CIR. BOULDER, CO 80301 Registration Number: 1722040 FEI Number*: 1000117361 Status: Active Initial Distributor/Importer: Yes *Note Firm may have additional establishment types. Please review listings for further information. Date of Registration Status: 2023
Owner/Operator | ENCISION, INC. 6797 WINCHESTER CIR. Boulder, CO US 80301 Owner/Operator Number: 9003185
Official Correspondent | danielle kaessner 6797 WINCHESTER CIR. Boulder, CO 80301 Phone: 303-4442600-125
510(k) Device Name | Applicant | 510(k)Number | Decision Date
aem disposable suction irrigation electrodes | ENCISION, INC. | K131949 | 08/26/2013
em3 aem monitor aem connectors | ENCISION, INC. | K122383 | 01/10/2013
aem disposable suction irrigation handset model es1700; aem disposable suction irrigation electrodes model es077x series | ENCISION, INC. | K122580 | 11/09/2012
aem monitoring system, models em2, emr, es4007, es4107, es9005 and es9015 | ENCISION, INC. | K093622 | 08/20/2010
suction irrigation electrodes and adapters | ENCISION, INC. | K100711 | 06/09/2010
aem disposable electrodes,model eso300 series | ENCISION, INC. | K091074 | 08/17/2009
disposable handle assembles, models dh0100 and dh0500 series | ENCISION, INC. | K090979 | 05/28/2009
disposable aem fixed tip electrodes and handpieces | ENCISION, INC. | K072789 | 10/19/2007

## Recalls
Recall Link | Trade Name | Recall Class | Posted internet Date | Termination Date | FDA Determine Cause | Quantity in Commerce | Manufacturer Reason For Recall | Action | Distribution
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=108356 | Device Recall AEM Disposable Electrodes | Class 2 | July 06, 2012 | August 07, 2012 | Device Design | 48 boxes | Encision is performing a field correction to their AEM Disposable Electrodes due to two consumer complaints reporting unanticipated alarms from the AEM monitor while using the device. This issue could result delays or complications during surgical procedures. | Encision sent an Urgent Medical Device Recall letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to lubricate the contact on the instrument cord with a conductive gel (to be supplied by Encision) prior to making the connection of the cord to the instrument. Instructions for applying the gel were attached. This corrective action was temporary; allowing customers to use their inventory of devices and for new product in the short term, until a design correction can resolve the problem. Customers would be contacted by their Encision Sales Representative to review this information for their specific situation and train their staff on its correction. For questions customers should call 303-339-6917 or Customer Service at 800-998-0986. For questions regarding this recall call 303-339-6917. | Worldwide Distribution - USA including (CA, TX, GA, NY, NC, OH), and New Zealand.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=103727 | Device Recall Encision AEM Suction Irrigation Electrode, LDiamond | Class 2 | August 30, 2011 | September 04, 2012 | Device Design | 169 units | Electrode tips have been reported breaking off during use or cleaning. | Encision Inc. sent out an Urgent Product Recall letter dated August 24, 2011 to all affected customers. The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory. For information on this recall call Encision at (303) 339-6917. | Worldwide Distribution: Nationwide distribution, and the country of New Zealand.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=103728 | Device Recall Encision AEM Suction Irrigation Electrode, Spatula | Class 2 | August 30, 2011 | September 04, 2012 | Device Design | 604 units | Electrode tips have been reported breaking off during use or cleaning. | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers. The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory. For information on this recall call Encision at (303) 339-6917. | Worldwide Distribution: Nationwide distribution, and the country of New Zealand.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=103729 | Device Recall Encision AEM Suction Irrigation Electrode, LHook, | Class 2 | August 30, 2011 | September 04, 2012 | Device Design | 212 units | Electrode tips have been reported breaking off during use or cleaning. | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers. The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory. For information on this recall call Encision at (303) 339-6917. | Worldwide Distribution: Nationwide distribution, and the country of New Zealand.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=103730 | Device Recall Encision AEM Suction Irrigation Electrode, JHook | Class 2 | August 30, 2011 | September 04, 2012 | Device Design | 83 units | Electrode tips have been reported breaking off during use or cleaning. | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers. The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory. For information on this recall call Encision at (303) 339-6917. | Worldwide Distribution: Nationwide distribution, and the country of New Zealand.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99522 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 44 Ten-Count Boxes | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99523 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 95 Ten-Count Boxes | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99505 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 574 Ten-Count Boxes, 24 Kits | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99517 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 952 Ten-Count Boxes | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99518 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 129 Ten-Count Boxes | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=99521 | Device Recall AEM Disposable Scissors | Class 2 | May 31, 2011 | August 09, 2011 | Package design/selection | 233 Ten Count Boxes, 1 Kit | Sterility may be compromised due to holes in packaging. | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection. Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90486 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 10 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90484 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 330 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90483 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 40 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90482 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 240 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90481 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 130 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90480 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 20 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90477 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 260 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=90485 | Device Recall AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | Class 2 | May 06, 2010 | May 21, 2010 | Nonconforming Material/Component | 110 units | Electrosurgical electrode tip insulation may crack and break off during use. | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=33733 | Device Recall Fixed Tip Electrosurgical Electrode. | Class 2 | September 22, 2004 | November 09, 2004 | Other | 335 units | Electrosurgical tip may break or come off during use. | Customers were notified beginning 6/15/2004 by phone or visit. | AL, CT, GA, IA, MN, MS, NC, OK, OR, TN, WA. Foreign distribution to Canada and Australia.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=34118 | Device Recall Disposable AEM Cord | Class 2 | July 22, 2004 | December 14, 2004 | Other | 73 units | Firm found pinholes in packaging which would compromise sterility. | Affected customers were notified by fax or mail on 7/15/2004. | AL, CA, GA, KS, KY, MI, MN, MS, NC, NH, NY, OH, OK, SD, TN, WA, WY. Foreign Distribution to Australia and Japan. Includes military facilities in GA, MS, and Japan.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=27981 | Device Recall Active Electrode Monitoring System | Class 2 | July 24, 2003 | November 06, 2003 | Other | 2 units | Use of NC versions of AEM monitor presents burn hazard when used with single pad and capacitively coupled patient return electrodes. | Accounts were visited by sales representatives beginning 6/23/2003. | CA, IL, KS, OH, OK, PA, TN. No government or military accounts. Foreign distribution to Canada and Italy.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=27980 | Device Recall Active Electrode Monitoring System | Class 2 | July 24, 2003 | November 06, 2003 | Other | 4 units | Use of NC versions of AEM monitor presents burn hazard when used with single pad and capacitively coupled patient return electrodes. | Accounts were visited by sales representatives beginning 6/23/2003. | CA, IL, KS, OH, OK, PA, TN. No government or military accounts. Foreign distribution to Canada and Italy.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=27979 | Device Recall Active Electrode Monitoring System | Class 2 | July 24, 2003 | November 06, 2003 | Other | 21 units | Use of NC versions of AEM monitor presents burn hazard when used with single pad and capacitively coupled patient return electrodes. | Accounts were visited by sales representatives beginning 6/23/2003. | CA, IL, KS, OH, OK, PA, TN. No government or military accounts. Foreign distribution to Canada and Italy.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=25675 | Device Recall Monopolar Laparoscopic Electrode | Class 2 | January 30, 2003 | April 14, 2003 | Other | 58 devices. | Reports of needle fixed-tip electrodes becoming loose or breaking off. | All consignees will be notified by phone or visit beginning 01/15/2003. | Domestic distribution to LA, MA, MN, OR, UT and WA.

## MAUDE
Report Number | Event Date | FDA Received Date | Report Lag | Event Type | Report Source | Manufacturer-D | Manufacturer-G | Brand Name | Type Of Device | Manufacturer Review | Event Text | Manufacturer Narrative | Was Device Available For Evaluation ? | Is This an Adverse Event Report ? | MAUDE Link
MW5105013 | 10/25/2021 | 10/28/2021 | 3 | malfunction | Voluntary | ENCISION, INC. | - | AEM MONITOR | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | - | Aem cable wire, proximal to the machine was defective (burnt intra-op during laparoscopic cholecystectomy, witnessed by rn. ) surgeon was informed at once and cable was disconnected from the machine. Bovie pad was changed to a new pad. Skin was clear and no abnormalities. Fda safety report id # (b)(4). | - | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=12719381&pc=GEI&device_sequence_no=1
1722040-2015-00001 | 09/29/2014 | 04/02/2015 | 185 | malfunction | Manufacturer | ENCISION, INC. boulder CO | - | AEM DISPOSABLE CURVED SCISSORS, 3/4 | ELECTROSURGICAL ELECTRODE | - | Aem disposable curved scissors (es0102 lot udf): broke tip off of scissors when using device. Physician retrieved broken scissors head. X-ray taken confirmed that there was zero retained metal. | Only the insert rod was returned. The scissor housing and all components were missing. Evaluation of the returned component, the work orders used to manufacture this product, complaint history for this product family, and testing other scissor inserts to failure all indicate that fracturing of the rod resulting in disengagement of the scissor housing will not occur if the product is used as indicated in the instructions for use. It is possible to fracture the rod with forces exceeding those expected when cutting soft tissue in surgery but the scissor housing is threaded into the handle retaining all broken components. Encision has not received any other reports of a fracture of the scissor rod resulting in components detaching. No trend has been identified. No corrective action required. | Yes | No | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/detail.cfm?mdrfoi__id=4666729&pc=GEI&device_sequence_no=1

## Inspections
City | State | Zip Code | Inspection End Date | Project Area | Classification
Boulder | Colorado | 80301 | 2019-09-20 | Compliance: Devices | Voluntary Action Indicated (VAI)
Boulder | Colorado | 80301 | 2019-09-20 | Postmarket Assurance: Devices | No Action Indicated (NAI)
Boulder | Colorado | 80301 | 2015-10-19 | Compliance: Devices | Voluntary Action Indicated (VAI)
Boulder | Colorado | 80301 | 2015-10-19 | Postmarket Assurance: Devices | No Action Indicated (NAI)
Boulder | Colorado | 80301 | 2012-12-07 | Compliance: Devices | Voluntary Action Indicated (VAI)
Boulder | Colorado | 80301 | 2012-12-07 | Postmarket Assurance: Devices | No Action Indicated (NAI)
Boulder | Colorado | 80301 | 2009-12-02 | Compliance: Devices | Voluntary Action Indicated (VAI)
Boulder | Colorado | 80301 | 2009-12-02 | Postmarket Assurance: Devices | Voluntary Action Indicated (VAI)

## Devices
Device Summary | Device Detail | TPLC Information | # Reported Events
Device Name | AEM DISPOSABLE ELECTRODES,MODEL ESO300 SERIES | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K091074 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer; Specification Developer | Detachment of Device or Device Component190 | 1833
Recalls | 9
Device Name | SUCTION IRRIGATION ELECTRODES AND ADAPTERS | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K100711 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | 5
Device Name | ENDOSHIELD BURN PROTECTION SYSTEM | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K140006 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | -
Device Name | MONOPOLAR LAPAROSCOPIC ELECTRODES W/ELETROSHIELD | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K912780 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | 8
Device Name | EM3 AEM MONITOR AEM CONNECTORS | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K122383 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer; Specification Developer | Detachment of Device or Device Component190 | 1833
Recalls | -
Device Name | AEM MONITORING SYSTEM, MODELS EM2, EMR, ES4007, ES4107, ES9005 AND ES9015 | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K093622 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | -
Device Name | Encision AEM Monopolar Laparoscopic Instruments and Accessories | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K191612 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | -
Device Name | Encision AEM enTouch 2X Scissor | TPLC Device Problem | MDR
Device Class | 2 | Break186 | 6041
Product Code | GEI | Adverse Event Without Identified Device or Use Problem187 | 3961
Regulation No | 878.44 | Energy Output Problem188 | 2516
Premarket Submission No | K201018 | Difficult to Open or Close189 | 1860
Establishment Operations | Manufacturer | Detachment of Device or Device Component190 | 1833
Recalls | -

## Info
Establishment Name | Error | Timestamp
ENCISION, INC. | Warning Letters Not Found | 2023-09-01 19:40